## User Inputs
|  |  | Outputs |  |  |  | Saturation Limits |  | Standard Camera | None
User Inputs |  |  | Power after filter | 9.9661e-05 | mW |  | Estimated Normalized ADC Peak % |  | WinCamD-IR-BB | PPBS-FS
Wavelength (nm) | 675 |  |  | 0.099661 | μW |  | 6.345e-11% |  | BeamR2/BeamMap | PPBS-CF
Laser Output Power (mW) | 2 | [7 empty] | TaperCamD-LCM | PPBS-BF
First Attenuation Optic | ND-4 |  | Sensor Irradiance | 0.063446 | W/m^2 |  | Estimated CW Exposure Time (ms) |  |  | PPBS-ZS
Second Attenuation Optic | None |  |  | 0.0063446 | mW/cm^2 |  | 2.5755 |  |  | CBS
Third Attenuation Optic | None | [8 empty] | ND-0.5
Gaussian Waist Diameter (um) | 2000 |  | Total Attenuation |  |  |  | Estimated Single Pulse ADC % |  |  | ND-1
Effective Beam Diameter (um) | 1414.2136 |  | Percent: | 0.009966% |  |  | 38.83% |  |  | ND-2
Pulsed? | Yes |  | Ratio: | 10033.9804 | :1 |  |  |  |  | ND-3
Pulse Repetition Rate (kHz) | 1 |  | OD value: | 4.0015 | [5 empty] | ND-4
Pulse Energy (mJ) | 0.001 | [8 empty] | ND-5
Pulse Width (ns) | 10 |  | CW Measurement | [6 empty] | ND-0.5-UV
Average Power (mW) | 1 | z (mm) | Element | Irradiance (mW/cm^2) | Power (W) | Damage Threshold (mW/cm^2) | Power Limit (W) | Diameter (um) | Power Absorbed (W) | ND-1-UV
|  | -19.1 | ND-4 | 63.6598 | 0.001 | 571154.5077 | 8.9718 | 1414.2255 | 0.0009999 | ND-2-UV
Sensor | Standard Camera | -12.8 | None | 0.0063445 | 9.9661e-08 | N/A | N/A | 1414.2189 | 0 | ND-3-UV
Instructions: Change/select only user inputs above Estimated CW Exposure time assumes WinCamD-LCM Pulsed lasers must satisy both Pulse and CW damage thresholds **All numbers are estimates, results vary depending on beam** **Damage thresholds are typical and can vary** v2020.02.19 - Compatible with Excel 2007 or later Questions? Contact support@dataray.com |  | -12.8 | None | 0.0063445 | 9.9661e-08 | N/A | N/A | 1414.2189 | 0 | ND-4-UV
|  | 0 | S-WCD-LCM Sensor | 0.0063446 | 9.9661e-08 | 16154.8251 | 0.25376 | 1414.2136 | 9.9661e-08 | ND-0.3-IR
ND-0.6-IR
|  |  | Pulse Measurement | [6 empty] | ND-1-IR
|  | z (mm) | Element | Irradiance (mJ/cm^2) | Pulse Energy (mJ) | Damage Threshold (mJ/cm^2) | Energy Limit (mJ) | Diameter (um) | Energy Absorbed (mJ) | ND-2-IR
|  | -19.1 | ND-4 | 0.06366 | 0.001 | 38255.6687 | 600.9288 | 1414.2255 | 0.0009999 | ND-3-IR
|  | -12.8 | None | 6.3445e-06 | 9.9661e-08 | N/A | N/A | 1414.2189 | 0
|  | -12.8 | None | 6.3445e-06 | 9.9661e-08 | N/A | N/A | 1414.2189 | 0
|  | 0 | S-WCD-LCM Sensor | 6.3446e-06 | 9.9661e-08 | 63.4398 | 0.99651 | 1414.2136 | 9.9661e-08
Rayleigh Range (mm) | Waist Area (mm^2)
4654.2113 | 1.5708
Yes
No
Note on B9: Loren Jones: Equivalent irradiance diameter of a flat circular beam assuming Gaussian beam is 2x the irradiance of an identical diameter flat beam
Note on J14: Effective Diameter (Diameter of equivalent irradiance flat-top)
Note on K14: Loren Jones: Or reflected
Note on D18: Default is 0 to have beam waist at the measurement plane. Can change for far-field measurements.

## Pulsed (hidden)
Outputs
|  |  | Pulse Energy after filter | 9.9661e-08 | mJ |  | QE | 47.36%
Wavelength (nm) | 675 |  |  | 9.9661e-05 | μJ |  | Energy of photon (J/photon) | 2.9449e-19
Laser Output Power (mW) | 1 | [5 empty] | Pixel Saturation (electrons/px) | 8000
First Attenuation Optic | ND-4 |  | Sensor Irradiance | 6.3446e-05 | J/m^2 |  | Pixel Saturation (photons/px) | 16893.0239
Second Attenuation Optic | None |  |  | 6.3446e-06 | mJ/cm^2 |  | Pixel Saturation (J/px) | 4.9749e-15
Third Attenuation Optic | None | [5 empty] | Pixel Size (microns) | 5.5
Spot Size Diameter (um) | 1414.2136 |  | Total Attenuation |  |  |  | Pixel Density (px/mm^2) | 32847
|  |  | Percent: | 0.009966% |  |  | Saturation (uJ/cm^2) | 0.016341
Pulsed? | Yes |  | Ratio: | 10033.9804 | :1 |  | 40 us Saturation (uW/cm^2) | 408.5231
Pulse Repetition Rate (kHz) | 1 |  | OD value: | 4.0015 |  |  | mW/cm^2 | 0.40852
Pulse Energy (mJ) | 0.001
Pulse Width (ns) | 10 |  | Element | Irradiance (mW/cm^2) | Power (W) | Damage Threshold (mW/cm^2) | Power Limit (W)
|  | 1 | ND-4 | 63.6598 | 0.001 | 0 | 8.9717
Average Power (mW) | 1 | 2 | None | 0.0063445 | 9.9661e-08 | N/A | N/A
|  | 3 | None | 0.0063445 | 9.9661e-08 | N/A | N/A
Planck's Constant | 6.6261e-34 | 4 | Sensor | 0.0063446 | 9.9661e-08 | N/A | #N/A
C (m/s) | 300000000
Frequency | 4.4444e+14 |  | Element | Irradiance (mJ/cm^2) | Pulse Energy (mJ) | Damage Threshold (mJ/cm^2) | Energy Limit (mJ) | Saturation Energy (uJ/cm^2)
|  | 1 | ND-4 | 0.06366 | 0.001 | 38255.6687 | 600.9186
|  | 2 | None | 6.3445e-06 | 9.9661e-08 | N/A | N/A
|  | 3 | None | 6.3445e-06 | 9.9661e-08 | N/A | N/A
|  | 4 | Sensor | 6.3446e-06 | 9.9661e-08 | N/A | #N/A | 0.016341
Beam Area
|  |  | 1.5708 | mm^2 |  | Estimated Single Pulse ADC % | Estimated CW Exposure Time (ms)
[6 empty] | 38.83% | 2.5755

## Surface Transmittance (hidden)
s_pol | p_pol | # | Material | Index | Angle (degrees) | Angle (radians)
0.5 | 0.5 | n1 | air | 1 | 3 | 0.05236
|  | n2 | ng9, ng4 | 1.51 | 1.9862 | 0.034667
|  | n3 | air | 1 | 3 | 0.05236
|  |  | Surface 1 | Surface 2
|  | r_s | -0.20356 | 0.20356
|  | t_s | 0.79644 | 1.2036
|  | r_p | 0.20282 | -0.20282
|  | t_p | 0.79657 | 1.2037
|  | T_s | 0.95856 | 0.95856
|  | T_p | 0.95886 | 0.95886
Transmittance
|  | T | 0.95871 | 0.95871 | 0.91913
Note on A1: Loren Jones: Polarization. Both set at 0.5 for unpolarized light
Note on E3: Loren Jones: Assumed constant with wavelength, would need more information from Schott to calculate per wavelengh
Note on F14: Loren Jones: After surfaces, using Fresnel equations

## PPBS Reflectance (hidden)
s_pol | p_pol | # | Material | Index | Angle (degrees) | Angle (radians)
0.5 | 0.5 | n1 | air | 1 | 45 | 0.7854
|  | n2 | UVFS | FALSE | #DIV/0! | #DIV/0!
|  | n3 | air | 1 | #DIV/0! | #DIV/0!
[7 empty] | wavelength= | 0.675 | um
|  |  | Surface 1 | Surface 2
|  | r_s | #DIV/0! | #DIV/0! |  |  | Material | Index of refraction
|  | t_s | #DIV/0! | #DIV/0! |  |  | UVFS | 1.4559
|  | r_p | #DIV/0! | #DIV/0! |  |  | CaF2 | 1.4321
|  | t_p | #DIV/0! | #DIV/0! |  |  | BaF2 | 1.4724
[7 empty] | ZnSe | 2.5676
|  | T_s | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
|  | T_p | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
Worst Case residual beams for Beam Dump(s)
[6 empty] | Reflectance | First Residual % of input | #DIV/0!
[6 empty] | #DIV/0! | Second Residual % of input | #DIV/0!
Note on A1: Loren Jones: Polarization. Both set at 0.5 for unpolarized light

## Internal Transmittance (hidden)
| NG9 | NG4 | NG1 | [11 empty] | 0.00010843
Wavelength | t_i per 1mm | t_i per 1mm | t_i per 1 mm |  |  |  | Filter # | Ref column | PPBS? | UV? | IR? | CBS? | Transmit per 1mm | Ref Thickness | Actual Thickness | t_i per actual
200 | 1e-05 | 1e-05 | 1e-05 | Power (mW) | 1 |  | 1st | 2 | 0 | 0 | 0 | 0 | 0.0658 | 1 | 3.355 | 0.00010843
201 | 1e-05 | 1e-05 | 1e-05 | Wavelength (nm) | 675 |  | 2nd | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1
202 | 1e-05 | 1e-05 | 1e-05 | Total Transmittance | 9.9661e-05 |  | 3rd | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1
203 | 1e-05 | 1e-05 | 1e-05 | Power after filter | 9.9661e-05
204 | 1e-05 | 1e-05 | 1e-05
205 | 1e-05 | 1e-05 | 1e-05 |  | W/cm^2 | W/(spot size) | mJ/cm^2 | mJ/(spot size) | z |  | LDT_LP | wl spec | pulse width spec (ns) |  | Thickness
206 | 1e-05 | 1e-05 | 1e-05 | None | N/A | N/A | N/A | N/A | 0 | [5 empty] | 0 | None | 2
207 | 1e-05 | 1e-05 | 1e-05 | PPBS |  |  |  |  | 25 | [6 empty] | PPBS-FS | 0 | 4
208 | 1e-05 | 1e-05 | 1e-05 | PPBS-FS | 398245.683 | 6255.6286 | 51798.4203 | 813.6477 | 25 |  | 30000 | 532 | 5.4 |  |  | PPBS-CF | 0
209 | 1e-05 | 1e-05 | 1e-05 | PPBS-BF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 25 |  | FALSE | FALSE | FALSE |  |  | PPBS-BF | 0
210 | 1e-05 | 1e-05 | 1e-05 | PPBS-CF | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 25 |  | FALSE | FALSE | FALSE |  |  | PPBS-ZS | 0
211 | 1e-05 | 1e-05 | 1e-05 | PPBS-ZS | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 25 |  | FALSE | FALSE | FALSE |  | 1.661 | CBS | 0
212 | 1e-05 | 1e-05 | 1e-05 | CBS | 38063.9098 | 597.9065 | 3806.391 | 59.7906 | 15 |  | 3000 | 532 | 10 |  | 0.975 | ND-0.5 | 0
213 | 1e-05 | 1e-05 | 1e-05 | ND-0.5 |  | 8.9717 | 38255.6687 | 600.9186 | 6.3 |  | 20000 | 532 | 4.4 |  | 0.815 | ND-1 | 0
214 | 1e-05 | 1e-05 | 1e-05 | ND-1 |  | 8.9717 | 38255.6687 | 600.9186 | 6.3 |  | 20000 | 532 | 4.4 | Yes | 1.661 | ND-2 | 0
215 | 1e-05 | 1e-05 | 1e-05 | ND-2 |  | 8.9717 | 38255.6687 | 600.9186 | 6.3 |  | 20000 | 532 | 4.4 | No | 2.508 | ND-3 | 0
216 | 1e-05 | 1e-05 | 1e-05 | ND-3 |  | 8.9717 | 38255.6687 | 600.9186 | 6.3 |  | 20000 | 532 | 4.4 |  | 3.355 | ND-4 | 1
217 | 1e-05 | 1e-05 | 1e-05 | ND-4 |  | 8.9717 | 38255.6687 | 600.9186 | 6.3 |  | 20000 | 532 | 4.4 |  | 4.202 | ND-5 | 0
218 | 1e-05 | 1e-05 | 1e-05 | ND-5 |  | 8.9717 | 38255.6687 | 600.9186 | 6.3 |  | 20000 | 532 | 4.4 |  |  | ND-0.5-UV | 0
219 | 1e-05 | 1e-05 | 1e-05 | ND-0.5-UV | 1064.7887 | 16.7257 | 147.6596 | 2.3194 | 6.3 |  | 56 | 355 | 5.2 |  |  | ND-1-UV | 0
220 | 1e-05 | 1e-05 | 1e-05 | ND-1-UV | 1064.7887 | 16.7257 | 147.6596 | 2.3194 | 6.3 |  | 56 | 355 | 5.2 |  |  | ND-2-UV | 0
221 | 1e-05 | 1e-05 | 1e-05 | ND-2-UV | 1064.7887 | 16.7257 | 147.6596 | 2.3194 | 6.3 |  | 56 | 355 | 5.2 |  |  | ND-3-UV | 0
222 | 1e-05 | 1e-05 | 1e-05 | ND-3-UV | 1064.7887 | 16.7257 | 147.6596 | 2.3194 | 6.3 |  | 56 | 355 | 5.2 |  |  | ND-4-UV | 0
223 | 1e-05 | 1e-05 | 1e-05 | ND-4-UV | 1064.7887 | 16.7257 | 147.6596 | 2.3194 | 6.3 |  | 56 | 355 | 5.2 |  |  | ND-0.3-IR | 0
224 | 1e-05 | 1e-05 | 1e-05 | S-WCD-LCM Sensor | 16.1548 | 0.25376 | 63.4398 | 0.99651 | 0 |  | 100 | 1064 | 10 |  |  | ND-0.6-IR | 0
225 | 1e-05 | 1e-05 | 1e-05 | BeamR2/BeamMap | 63.662 | 1 | 221.4894 | 3.4791 | 0 |  | 84 | 355 | 5.2 |  |  | ND-1-IR | 0
226 | 1e-05 | 1e-05 | 1e-05 | ND-0.3-IR | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 7 |  | FALSE | FALSE | FALSE |  |  | ND-2-IR | 0
227 | 1e-05 | 1e-05 | 1e-05 | ND-1-IR | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 7 |  | FALSE | FALSE | FALSE |  |  | ND-3-IR | 0
228 | 1e-05 | 1e-05 | 1e-05 | ND-2-IR | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 7 |  | FALSE | FALSE | FALSE
229 | 1e-05 | 1e-05 | 1e-05 | ND-3-IR | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 7 |  | FALSE | FALSE | FALSE
230 | 1e-05 | 1e-05 | 1e-05 | S-WCD-IR-BB Sensor | 0.2 | 0.0031416 | 6 | 0.094248 | 0
231 | 1e-05 | 1e-05 | 1e-05
232 | 1e-05 | 1e-05 | 1e-05 |  | ND | 8.9717
233 | 1e-05 | 1e-05 | 1e-05 |  | ND W/mm | 6.344 | 5000 | mJ/cm^2
234 | 1e-05 | 1e-05 | 1e-05 |  |  |  | 532 | nm
235 | 1e-05 | 1e-05 | 1e-05 |  | Slit-2.5 | 7.1774 | 10 | ns
236 | 1e-05 | 1e-05 | 1e-05 |  | W/mm | 5.0752
237 | 1e-05 | 1e-05 | 1e-05
238 | 1e-05 | 1e-05 | 1e-05 |  |  |  | 10000 | mJ/cm^2
239 | 1e-05 | 1e-05 | 1e-05 |  |  |  | 532 | nm
240 | 1e-05 | 1e-05 | 1e-05 |  |  |  | 10 | ns
241 | 1e-05 | 1e-05 | 1e-05
242 | 1e-05 | 1e-05 | 1e-05
243 | 1e-05 | 1e-05 | 1e-05 |  |  |  |  | 1
244 | 1e-05 | 1e-05 | 1e-05 |  | LCM Sensor | 0.2
245 | 1e-05 | 1e-05 | 1e-05 |  | W/mm | 0.1
246 | 1e-05 | 1e-05 | 1e-05
247 | 1e-05 | 1e-05 | 1e-05
248 | 1e-05 | 1e-05 | 1e-05
249 | 1e-05 | 1e-05 | 1e-05
250 | 1e-05 | 1e-05 | 1e-05
251 | 1e-05 | 1e-05 | 1e-05
252 | 1e-05 | 1e-05 | 1e-05
253 | 1e-05 | 1e-05 | 1e-05
254 | 1e-05 | 1e-05 | 1e-05
255 | 1e-05 | 1e-05 | 1e-05
256 | 1e-05 | 1e-05 | 1e-05
257 | 1e-05 | 1e-05 | 1e-05
258 | 1e-05 | 1e-05 | 1e-05
259 | 1e-05 | 1e-05 | 1e-05
260 | 1e-05 | 1e-05 | 1e-05
261 | 1e-05 | 1e-05 | 1e-05
262 | 1e-05 | 1e-05 | 1e-05
263 | 1e-05 | 1e-05 | 1e-05
264 | 1e-05 | 1e-05 | 1e-05
265 | 1e-05 | 1e-05 | 1e-05
266 | 1e-05 | 1e-05 | 1e-05
267 | 1e-05 | 1e-05 | 1e-05
268 | 1e-05 | 1e-05 | 1e-05
269 | 1e-05 | 1e-05 | 1e-05
270 | 1e-05 | 1e-05 | 1e-05
271 | 1e-05 | 1e-05 | 1e-05
272 | 1e-05 | 1e-05 | 1e-05
273 | 1e-05 | 1e-05 | 1e-05
274 | 1e-05 | 1e-05 | 1e-05
275 | 1e-05 | 1e-05 | 1e-05
276 | 1e-05 | 1e-05 | 1e-05
277 | 1e-05 | 1e-05 | 1e-05
278 | 1e-05 | 1e-05 | 1e-05
279 | 1e-05 | 1e-05 | 1e-05
280 | 1e-05 | 1e-05 | 1e-05
281 | 1e-05 | 1e-05 | 1e-05
282 | 1e-05 | 1e-05 | 1e-05
283 | 1e-05 | 1e-05 | 1e-05
284 | 1e-05 | 1e-05 | 1e-05
285 | 1e-05 | 1e-05 | 1e-05
286 | 1e-05 | 1e-05 | 1e-05
287 | 1e-05 | 1e-05 | 1e-05
288 | 1e-05 | 1e-05 | 1e-05
289 | 1e-05 | 1e-05 | 1e-05
290 | 1e-05 | 1e-05 | 1e-05
291 | 1e-05 | 1e-05 | 1e-05
292 | 1e-05 | 1e-05 | 1e-05
293 | 1e-05 | 1e-05 | 1e-05
294 | 1e-05 | 1e-05 | 1e-05
295 | 1e-05 | 1e-05 | 1e-05
296 | 1e-05 | 1e-05 | 1e-05
297 | 1e-05 | 1e-05 | 1e-05
298 | 1e-05 | 1e-05 | 1e-05
299 | 1e-05 | 1e-05 | 1e-05
300 | 1e-05 | 1e-05 | 1e-05
301 | 1e-05 | 1e-05 | 1e-05
302 | 1e-05 | 1e-05 | 1e-05
303 | 1e-05 | 1e-05 | 1e-05
304 | 1e-05 | 1e-05 | 1e-05
305 | 1e-05 | 1e-05 | 1e-05
306 | 1e-05 | 1e-05 | 1e-05
307 | 1e-05 | 1e-05 | 1e-05
308 | 1e-05 | 1e-05 | 1e-05
309 | 1e-05 | 1e-05 | 1e-05
310 | 1e-05 | 1e-05 | 1e-05
311 | 1e-05 | 1e-05 | 1e-05
312 | 1e-05 | 1e-05 | 1e-05
313 | 1e-05 | 1e-05 | 1e-05
314 | 1e-05 | 1e-05 | 1e-05
315 | 1e-05 | 1e-05 | 1e-05
316 | 1e-05 | 1e-05 | 1e-05
317 | 1e-05 | 1e-05 | 1e-05
318 | 1e-05 | 1e-05 | 1e-05
319 | 1e-05 | 1e-05 | 1e-05
320 | 1e-05 | 1e-05 | 1e-05
321 | 1e-05 | 1e-05 | 1e-05
322 | 1e-05 | 1e-05 | 1e-05
323 | 1e-05 | 1e-05 | 1e-05
324 | 1e-05 | 1e-05 | 1e-05
325 | 1e-05 | 1.4757e-05 | 1e-05
326 | 1e-05 | 2.5003e-05 | 1e-05
327 | 1e-05 | 4.5394e-05 | 1e-05
328 | 1e-05 | 7.7625e-05 | 1e-05
329 | 1e-05 | 0.00013002 | 1e-05
330 | 1e-05 | 0.00019907 | 1e-05
331 | 1e-05 | 0.00030061 | 1e-05
332 | 1e-05 | 0.00045604 | 1e-05
333 | 1e-05 | 0.00069663 | 1e-05
334 | 1e-05 | 0.0010666 | 1e-05
335 | 1e-05 | 0.0015311 | 1e-05
336 | 1e-05 | 0.0022542 | 1e-05
337 | 1e-05 | 0.0029512 | 1e-05
338 | 1e-05 | 0.0038194 | 1e-05
339 | 1e-05 | 0.0048195 | 1e-05
340 | 1e-05 | 0.0060395 | 1e-05
341 | 1e-05 | 0.0075162 | 1e-05
342 | 1e-05 | 0.0091622 | 1e-05
343 | 1e-05 | 0.01122 | 1e-05
344 | 1e-05 | 0.013836 | 1e-05
345 | 1e-05 | 0.016711 | 1e-05
346 | 1.5031e-05 | 0.019634 | 1e-05
347 | 2.2751e-05 | 0.022751 | 1e-05
348 | 3.2359e-05 | 0.026062 | 1e-05
349 | 4.5604e-05 | 0.029923 | 1e-05
350 | 6.4121e-05 | 0.0339 | 1e-05
351 | 8.8716e-05 | 0.0382 | 1e-05
352 | 0.00011912 | 0.0433 | 1e-05
353 | 0.00016255 | 0.0481 | 1e-05
354 | 0.00021528 | 0.0534 | 1e-05
355 | 0.00027669 | 0.0585 | 1e-05
356 | 0.00035156 | 0.0645 | 1e-05
357 | 0.00045186 | 0.07 | 1e-05
358 | 0.00055335 | 0.0763 | 1e-05
359 | 0.00069024 | 0.0826 | 1e-05
360 | 0.00084723 | 0.0888 | 1e-05
361 | 0.0010233 | 0.0947 | 1e-05
362 | 0.0012246 | 0.1012 | 1e-05
363 | 0.0014488 | 0.1074 | 1e-05
364 | 0.0016749 | 0.1134 | 1e-05
365 | 0.0019543 | 0.119 | 1e-05
366 | 0.0022182 | 0.1238 | 1e-05
367 | 0.002541 | 0.1286 | 1e-05
368 | 0.0028249 | 0.1329 | 1e-05
369 | 0.0031189 | 0.1366 | 1e-05
370 | 0.0034041 | 0.1391 | 1e-05
371 | 0.0035892 | 0.1397 | 1e-05
372 | 0.0037325 | 0.1391 | 1e-05
373 | 0.0038371 | 0.1374 | 1e-05
374 | 0.0039264 | 0.1363 | 1e-05
375 | 0.0040738 | 0.136 | 1e-05
376 | 0.0042364 | 0.136 | 1e-05
377 | 0.0044463 | 0.1366 | 1e-05
378 | 0.0046452 | 0.1374 | 1e-05
379 | 0.0049317 | 0.1397 | 1e-05
380 | 0.0052845 | 0.1439 | 1e-05
381 | 0.0057677 | 0.1498 | 1e-05
382 | 0.0063826 | 0.1578 | 1e-05
383 | 0.007145 | 0.1665 | 1e-05
384 | 0.0079799 | 0.1759 | 1e-05
385 | 0.0089125 | 0.1849 | 1e-05
386 | 0.01005 | 0.1928 | 1e-05
387 | 0.01129 | 0.2013 | 1e-05
388 | 0.01236 | 0.2098 | 1e-05
389 | 0.01361 | 0.2177 | 1e-05
390 | 0.01464 | 0.225 | 1e-05
391 | 0.01554 | 0.2313 | 1e-05
392 | 0.01653 | 0.2372 | 1e-05
393 | 0.01743 | 0.2429 | 1e-05
394 | 0.01825 | 0.2479 | 1e-05
395 | 0.0191 | 0.253 | 1e-05
396 | 0.01992 | 0.2567 | 1e-05
397 | 0.02065 | 0.2607 | 1e-05
398 | 0.02147 | 0.2635 | 1e-05
399 | 0.02224 | 0.266 | 1e-05
400 | 0.02301 | 0.2689 | 1e-05
401 | 0.0237 | 0.271 | 1e-05
402 | 0.0243 | 0.27241 | 1e-05
403 | 0.0249 | 0.27346 | 1e-05
404 | 0.02546 | 0.2744 | 1e-05
405 | 0.02606 | 0.27508 | 1e-05
406 | 0.02653 | 0.27539 | 1e-05
407 | 0.02692 | 0.27534 | 1e-05
408 | 0.02735 | 0.27513 | 1e-05
409 | 0.02765 | 0.27497 | 1e-05
410 | 0.02795 | 0.27503 | 1e-05
411 | 0.02821 | 0.27518 | 1e-05
412 | 0.02846 | 0.2755 | 1e-05
413 | 0.02868 | 0.27592 | 1e-05
414 | 0.02894 | 0.27644 | 1e-05
415 | 0.02915 | 0.27691 | 1e-05
416 | 0.02945 | 0.27754 | 1e-05
417 | 0.02979 | 0.27832 | 1e-05
418 | 0.03022 | 0.27911 | 1e-05
419 | 0.03065 | 0.2799 | 1e-05
420 | 0.03108 | 0.28094 | 1e-05
421 | 0.03151 | 0.28209 | 1e-05
422 | 0.03198 | 0.2833 | 1e-05
423 | 0.03237 | 0.28461 | 1e-05
424 | 0.03271 | 0.28581 | 1e-05
425 | 0.03301 | 0.28707 | 1e-05
426 | 0.03336 | 0.28806 | 1e-05
427 | 0.03366 | 0.28895 | 1e-05
428 | 0.03396 | 0.28974 | 1e-05
429 | 0.03417 | 0.29047 | 1e-05
430 | 0.03435 | 0.29099 | 1e-05
431 | 0.03452 | 0.29126 | 1e-05
432 | 0.03465 | 0.29141 | 1e-05
433 | 0.03473 | 0.29131 | 1e-05
434 | 0.03486 | 0.29115 | 1e-05
435 | 0.03503 | 0.29099 | 1e-05
436 | 0.0352 | 0.29058 | 1e-05
437 | 0.03525 | 0.29 | 1e-05
438 | 0.03538 | 0.28937 | 1e-05
439 | 0.03546 | 0.2889 | 1e-05
440 | 0.03555 | 0.28901 | 1e-05
441 | 0.03563 | 0.28942 | 1e-05
442 | 0.03581 | 0.29031 | 1e-05
443 | 0.03611 | 0.29147 | 1e-05
444 | 0.03632 | 0.29267 | 1e-05
445 | 0.03666 | 0.29398 | 1e-05
446 | 0.03696 | 0.29534 | 1e-05
447 | 0.03735 | 0.29686 | 1e-05
448 | 0.03787 | 0.29864 | 1e-05
449 | 0.03821 | 0.30031 | 1e-05
450 | 0.038717 | 0.30199 | 1e-05
451 | 0.039204 | 0.30362 | 1e-05
452 | 0.039699 | 0.30521 | 1e-05
453 | 0.040225 | 0.30693 | 1e-05
454 | 0.04072 | 0.30848 | 1e-05
455 | 0.041223 | 0.31 | 1e-05
456 | 0.041639 | 0.31146 | 1e-05
457 | 0.042031 | 0.31271 | 1.0046e-05
458 | 0.042408 | 0.31401 | 1.0471e-05
459 | 0.042777 | 0.31506 | 1.0889e-05
460 | 0.043084 | 0.31598 | 1.1194e-05
461 | 0.043288 | 0.31652 | 1.1535e-05
462 | 0.043484 | 0.31686 | 1.1885e-05
463 | 0.043626 | 0.31698 | 1.218e-05
464 | 0.043736 | 0.31698 | 1.2474e-05
465 | 0.043814 | 0.31698 | 1.2708e-05
466 | 0.043814 | 0.31677 | 1.2953e-05
467 | 0.043783 | 0.31648 | 1.3154e-05
468 | 0.043712 | 0.31602 | 1.3305e-05
469 | 0.043634 | 0.31548 | 1.3468e-05
470 | 0.043555 | 0.31497 | 1.365e-05
471 | 0.043469 | 0.31443 | 1.3795e-05
472 | 0.043382 | 0.31388 | 1.3952e-05
473 | 0.043288 | 0.3133 | 1.4103e-05
474 | 0.043186 | 0.31263 | 1.4223e-05
475 | 0.04306 | 0.312 | 1.4341e-05
476 | 0.042927 | 0.31146 | 1.4398e-05
477 | 0.042777 | 0.31083 | 1.4392e-05
478 | 0.042636 | 0.3102 | 1.4385e-05
479 | 0.042479 | 0.30965 | 1.436e-05
480 | 0.042314 | 0.30894 | 1.4341e-05
481 | 0.042149 | 0.30815 | 1.4322e-05
482 | 0.041953 | 0.30739 | 1.4341e-05
483 | 0.04178 | 0.30655 | 1.4354e-05
484 | 0.041599 | 0.30572 | 1.4348e-05
485 | 0.04145 | 0.30501 | 1.4341e-05
486 | 0.041277 | 0.30429 | 1.4322e-05
487 | 0.04112 | 0.30366 | 1.4285e-05
488 | 0.040979 | 0.30304 | 1.4223e-05
489 | 0.04083 | 0.30249 | 1.4158e-05
490 | 0.040688 | 0.30203 | 1.4093e-05
491 | 0.040555 | 0.30145 | 1.4008e-05
492 | 0.040429 | 0.30103 | 1.3914e-05
493 | 0.040304 | 0.30069 | 1.3804e-05
494 | 0.040178 | 0.30036 | 1.374e-05
495 | 0.040076 | 0.29998 | 1.3631e-05
496 | 0.039958 | 0.2996 | 1.355e-05
497 | 0.039856 | 0.29914 | 1.3512e-05
498 | 0.039754 | 0.29872 | 1.3468e-05
499 | 0.03966 | 0.29839 | 1.343e-05
500 | 0.039589 | 0.29801 | 1.3399e-05
501 | 0.039493 | 0.29769 | 1.3386e-05
502 | 0.039408 | 0.2975 | 1.3399e-05
503 | 0.039328 | 0.29731 | 1.3399e-05
504 | 0.039262 | 0.29715 | 1.3399e-05
505 | 0.039206 | 0.29699 | 1.3405e-05
506 | 0.039163 | 0.29677 | 1.3355e-05
507 | 0.039121 | 0.29665 | 1.328e-05
508 | 0.039102 | 0.29649 | 1.3229e-05
509 | 0.039088 | 0.2963 | 1.3183e-05
510 | 0.039079 | 0.29621 | 1.3173e-05
511 | 0.039088 | 0.29615 | 1.3185e-05
512 | 0.039102 | 0.29608 | 1.3217e-05
513 | 0.039135 | 0.29608 | 1.3273e-05
514 | 0.039168 | 0.29612 | 1.3349e-05
515 | 0.039201 | 0.29621 | 1.3397e-05
516 | 0.039225 | 0.29634 | 1.3462e-05
517 | 0.039239 | 0.29646 | 1.349e-05
518 | 0.039262 | 0.29662 | 1.3521e-05
519 | 0.039295 | 0.29687 | 1.3583e-05
520 | 0.039333 | 0.29709 | 1.365e-05
521 | 0.03938 | 0.29731 | 1.3738e-05
522 | 0.039455 | 0.29756 | 1.382e-05
523 | 0.039536 | 0.29794 | 1.3902e-05
524 | 0.039611 | 0.29819 | 1.3996e-05
525 | 0.039701 | 0.29857 | 1.4096e-05
526 | 0.03979 | 0.29891 | 1.4203e-05
527 | 0.039884 | 0.29938 | 1.4289e-05
528 | 0.039993 | 0.29982 | 1.4398e-05
529 | 0.040101 | 0.30029 | 1.4492e-05
530 | 0.040209 | 0.3007 | 1.4593e-05
531 | 0.040299 | 0.30117 | 1.4689e-05
532 | 0.040403 | 0.30161 | 1.4794e-05
533 | 0.040497 | 0.30212 | 1.4901e-05
534 | 0.040596 | 0.30252 | 1.4997e-05
535 | 0.04069 | 0.30299 | 1.5101e-05
536 | 0.040803 | 0.30334 | 1.5205e-05
537 | 0.040902 | 0.30369 | 1.5311e-05
538 | 0.04101 | 0.30403 | 1.5382e-05
539 | 0.041119 | 0.30431 | 1.5488e-05
540 | 0.041199 | 0.30463 | 1.5585e-05
541 | 0.04126 | 0.30491 | 1.5654e-05
542 | 0.041321 | 0.30516 | 1.574e-05
543 | 0.041364 | 0.30541 | 1.5786e-05
544 | 0.041406 | 0.3056 | 1.5812e-05
545 | 0.041439 | 0.30579 | 1.5818e-05
546 | 0.041482 | 0.30595 | 1.5799e-05
547 | 0.041515 | 0.30601 | 1.574e-05
548 | 0.041538 | 0.30604 | 1.5704e-05
549 | 0.041566 | 0.30607 | 1.5642e-05
550 | 0.041576 | 0.30601 | 1.5585e-05
551 | 0.041562 | 0.30598 | 1.5529e-05
552 | 0.041543 | 0.30585 | 1.5488e-05
553 | 0.041515 | 0.30576 | 1.5417e-05
554 | 0.041486 | 0.3056 | 1.5359e-05
555 | 0.041444 | 0.30548 | 1.529e-05
556 | 0.041406 | 0.30523 | 1.5208e-05
557 | 0.041369 | 0.30497 | 1.5136e-05
558 | 0.041326 | 0.30469 | 1.5045e-05
559 | 0.041274 | 0.30438 | 1.4957e-05
560 | 0.041199 | 0.30397 | 1.4819e-05
561 | 0.041095 | 0.30353 | 1.4699e-05
562 | 0.040973 | 0.30303 | 1.453e-05
563 | 0.040836 | 0.30246 | 1.4392e-05
564 | 0.040709 | 0.30196 | 1.4228e-05
565 | 0.040572 | 0.30146 | 1.4077e-05
566 | 0.040445 | 0.3008 | 1.3939e-05
567 | 0.040318 | 0.3002 | 1.3795e-05
568 | 0.040181 | 0.29957 | 1.3657e-05
569 | 0.040068 | 0.29894 | 1.3537e-05
570 | 0.03995 | 0.29828 | 1.338e-05
571 | 0.039847 | 0.29759 | 1.3243e-05
572 | 0.039743 | 0.2969 | 1.3166e-05
573 | 0.039639 | 0.29627 | 1.3079e-05
574 | 0.03954 | 0.29558 | 1.3003e-05
575 | 0.039451 | 0.29498 | 1.2921e-05
576 | 0.039366 | 0.29432 | 1.2859e-05
577 | 0.039281 | 0.29373 | 1.2815e-05
578 | 0.039206 | 0.29313 | 1.2764e-05
579 | 0.03914 | 0.2925 | 1.2727e-05
580 | 0.039083 | 0.29197 | 1.272e-05
581 | 0.039036 | 0.2915 | 1.2733e-05
582 | 0.039003 | 0.29106 | 1.2758e-05
583 | 0.038975 | 0.29068 | 1.2802e-05
584 | 0.038956 | 0.2903 | 1.2859e-05
585 | 0.038956 | 0.28996 | 1.2896e-05
586 | 0.038965 | 0.28968 | 1.294e-05
587 | 0.03897 | 0.28946 | 1.2997e-05
588 | 0.038994 | 0.2892 | 1.3035e-05
589 | 0.039027 | 0.28905 | 1.3072e-05
590 | 0.039083 | 0.28898 | 1.3148e-05
591 | 0.039149 | 0.28886 | 1.3254e-05
592 | 0.03925 | 0.28883 | 1.3455e-05
593 | 0.03939 | 0.2888 | 1.3669e-05
594 | 0.03955 | 0.2888 | 1.3883e-05
595 | 0.03971 | 0.28898 | 1.4128e-05
596 | 0.0399 | 0.28927 | 1.4354e-05
597 | 0.04008 | 0.28961 | 1.458e-05
598 | 0.04027 | 0.29 | 1.4794e-05
599 | 0.04047 | 0.29046 | 1.5058e-05
600 | 0.04067 | 0.29099 | 1.5328e-05
601 | 0.04087 | 0.29153 | 1.56e-05
602 | 0.04106 | 0.29214 | 1.5704e-05
603 | 0.04127 | 0.2927 | 1.5959e-05
604 | 0.04148 | 0.29326 | 1.6293e-05
605 | 0.04169 | 0.29376 | 1.6596e-05
606 | 0.0419 | 0.29434 | 1.6904e-05
607 | 0.04212 | 0.29487 | 1.7298e-05
608 | 0.04234 | 0.29546 | 1.7701e-05
609 | 0.04253 | 0.29601 | 1.803e-05
610 | 0.04276 | 0.29657 | 1.8408e-05
611 | 0.04297 | 0.29719 | 1.8836e-05
612 | 0.04318 | 0.29772 | 1.9143e-05
613 | 0.0434 | 0.29836 | 1.9543e-05
614 | 0.04363 | 0.29898 | 1.9907e-05
615 | 0.04387 | 0.29953 | 2.0324e-05
616 | 0.04409 | 0.3 | 2.0654e-05
617 | 0.04432 | 0.30047 | 2.0989e-05
618 | 0.04455 | 0.30094 | 2.138e-05
619 | 0.04475 | 0.30138 | 2.1727e-05
620 | 0.045 | 0.30182 | 2.208e-05
621 | 0.0452 | 0.30235 | 2.2646e-05
622 | 0.04538 | 0.30279 | 2.3014e-05
623 | 0.04559 | 0.3032 | 2.3823e-05
624 | 0.04579 | 0.30361 | 2.421e-05
625 | 0.04598 | 0.30399 | 2.4831e-05
626 | 0.04618 | 0.30425 | 2.5177e-05
627 | 0.04639 | 0.30443 | 2.5645e-05
628 | 0.04661 | 0.30461 | 2.6062e-05
629 | 0.04679 | 0.30478 | 2.6424e-05
630 | 0.047 | 0.30496 | 2.6915e-05
631 | 0.04718 | 0.30519 | 2.7353e-05
632 | 0.0474 | 0.30537 | 2.7861e-05
633 | 0.04758 | 0.30557 | 2.8445e-05
634 | 0.0478 | 0.30578 | 2.8973e-05
635 | 0.048 | 0.30598 | 2.9512e-05
636 | 0.04818 | 0.30616 | 3.013e-05
637 | 0.0484 | 0.30639 | 3.0549e-05
638 | 0.04859 | 0.3066 | 3.1117e-05
639 | 0.04878 | 0.30677 | 3.1842e-05
640 | 0.049 | 0.30698 | 3.2509e-05
641 | 0.04918 | 0.30719 | 3.3266e-05
642 | 0.04937 | 0.30739 | 3.3729e-05
643 | 0.04954 | 0.3076 | 3.4514e-05
644 | 0.04974 | 0.3078 | 3.5075e-05
645 | 0.04999 | 0.30801 | 3.5975e-05
646 | 0.05031 | 0.30815 | 3.6813e-05
647 | 0.05063 | 0.30833 | 3.7584e-05
648 | 0.05092 | 0.30845 | 3.8726e-05
649 | 0.05123 | 0.30865 | 3.9811e-05
650 | 0.05159 | 0.309 | 4.0738e-05
651 | 0.05189 | 0.30942 | 4.1976e-05
652 | 0.05219 | 0.31 | 4.3451e-05
653 | 0.05258 | 0.31055 | 4.4978e-05
654 | 0.05292 | 0.31121 | 4.6345e-05
655 | 0.05326 | 0.31197 | 4.7643e-05
656 | 0.05361 | 0.31272 | 4.9091e-05
657 | 0.05399 | 0.31338 | 5.0582e-05
658 | 0.05442 | 0.31432 | 5.2119e-05
659 | 0.05485 | 0.31508 | 5.3456e-05
660 | 0.05537 | 0.31593 | 5.559e-05
661 | 0.0558 | 0.31687 | 5.7412e-05
662 | 0.0564 | 0.31781 | 5.9704e-05
663 | 0.057 | 0.31875 | 6.223e-05
664 | 0.05764 | 0.31979 | 6.4565e-05
665 | 0.05833 | 0.32092 | 6.7608e-05
666 | 0.05893 | 0.32215 | 7.0469e-05
667 | 0.05953 | 0.32347 | 7.2778e-05
668 | 0.06017 | 0.32497 | 7.6384e-05
669 | 0.06082 | 0.32629 | 7.9433e-05
670 | 0.06155 | 0.3279 | 8.3176e-05
671 | 0.06228 | 0.32959 | 8.6696e-05
672 | 0.06305 | 0.33129 | 9.0157e-05
673 | 0.06395 | 0.33308 | 9.4842e-05
674 | 0.0649 | 0.33496 | 9.8628e-05
675 | 0.0658 | 0.33694 | 0.00010375
676 | 0.06666 | 0.33902 | 0.00010889
677 | 0.06743 | 0.34118 | 0.00011429
678 | 0.06842 | 0.34335 | 0.00011967
679 | 0.06936 | 0.3458 | 0.00012589
680 | 0.07039 | 0.34797 | 0.00013213
681 | 0.07134 | 0.34985 | 0.00013868
682 | 0.07237 | 0.35174 | 0.00014655
683 | 0.07348 | 0.35381 | 0.00015382
684 | 0.0746 | 0.35579 | 0.00016144
685 | 0.07567 | 0.35796 | 0.00016943
686 | 0.07675 | 0.36013 | 0.00017742
687 | 0.07769 | 0.3622 | 0.00018578
688 | 0.07872 | 0.36446 | 0.00019364
689 | 0.07979 | 0.36653 | 0.00020184
690 | 0.081 | 0.36851 | 0.00021038
691 | 0.08194 | 0.37059 | 0.00022131
692 | 0.08297 | 0.37257 | 0.00023121
693 | 0.08396 | 0.37436 | 0.00024322
694 | 0.08499 | 0.37634 | 0.00025351
695 | 0.08602 | 0.37784 | 0.00026607
696 | 0.08696 | 0.37973 | 0.00027606
697 | 0.08804 | 0.38171 | 0.00028774
698 | 0.08898 | 0.3835 | 0.0003013
699 | 0.09 | 0.38529 | 0.00031405
700 | 0.09104 | 0.38698 | 0.00032434
701 | 0.09202 | 0.38854 | 0.00033497
702 | 0.09301 | 0.38988 | 0.00035075
703 | 0.09399 | 0.39131 | 0.00036392
704 | 0.09509 | 0.39261 | 0.00037757
705 | 0.09601 | 0.39395 | 0.00039264
706 | 0.09682 | 0.39525 | 0.00040458
707 | 0.09762 | 0.39646 | 0.00041687
708 | 0.09846 | 0.39763 | 0.00043351
709 | 0.09924 | 0.39876 | 0.00044668
710 | 0.1 | 0.39998 | 0.00046238
711 | 0.10081 | 0.40107 | 0.00047643
712 | 0.10158 | 0.40212 | 0.00049091
713 | 0.10242 | 0.40308 | 0.00050582
714 | 0.10319 | 0.40396 | 0.00052119
715 | 0.10407 | 0.40475 | 0.00053703
716 | 0.10466 | 0.40559 | 0.00055463
717 | 0.10536 | 0.40626 | 0.00056885
718 | 0.10605 | 0.40685 | 0.00058614
719 | 0.10675 | 0.40739 | 0.00060256
720 | 0.10737 | 0.40798 | 0.00061518
721 | 0.10799 | 0.4084 | 0.00063387
722 | 0.10862 | 0.40882 | 0.00064863
723 | 0.10931 | 0.40919 | 0.00066681
724 | 0.10994 | 0.40965 | 0.00068549
725 | 0.11056 | 0.41 | 0.00069984
726 | 0.11104 | 0.41037 | 0.0007145
727 | 0.11162 | 0.41079 | 0.0007379
728 | 0.11217 | 0.41108 | 0.00075509
729 | 0.11276 | 0.41133 | 0.00077804
730 | 0.11338 | 0.41166 | 0.00079983
731 | 0.1139 | 0.41196 | 0.00081846
732 | 0.11448 | 0.41229 | 0.00083368
733 | 0.1151 | 0.41254 | 0.00085114
734 | 0.11565 | 0.41284 | 0.00086896
735 | 0.11628 | 0.41301 | 0.00089125
736 | 0.11679 | 0.41317 | 0.00090991
737 | 0.1173 | 0.41334 | 0.00092683
738 | 0.11767 | 0.41338 | 0.00094624
739 | 0.11815 | 0.41338 | 0.00096605
740 | 0.11862 | 0.41342 | 0.00098628
741 | 0.11906 | 0.41351 | 0.0010046
742 | 0.11947 | 0.41355 | 0.0010233
743 | 0.11994 | 0.41363 | 0.0010544
744 | 0.12035 | 0.41368 | 0.0010715
745 | 0.12082 | 0.41372 | 0.0011041
746 | 0.12115 | 0.41372 | 0.0011272
747 | 0.12159 | 0.41359 | 0.0011429
748 | 0.12196 | 0.41355 | 0.0011641
749 | 0.12232 | 0.41351 | 0.0011803
750 | 0.1228 | 0.41351 | 0.0012078
751 | 0.12322 | 0.41336 | 0.0012331
752 | 0.12365 | 0.41326 | 0.0012589
753 | 0.12412 | 0.41313 | 0.0012882
754 | 0.12456 | 0.413 | 0.0013183
755 | 0.12499 | 0.41274 | 0.0013459
756 | 0.12534 | 0.41258 | 0.0013709
757 | 0.12573 | 0.41232 | 0.0013996
758 | 0.12608 | 0.41203 | 0.0014223
759 | 0.1264 | 0.41174 | 0.0014421
760 | 0.12675 | 0.41154 | 0.0014689
761 | 0.12705 | 0.41141 | 0.0014894
762 | 0.12739 | 0.41131 | 0.0015171
763 | 0.12772 | 0.41122 | 0.0015488
764 | 0.12804 | 0.41112 | 0.0015849
765 | 0.12834 | 0.41099 | 0.0016069
766 | 0.12861 | 0.41073 | 0.0016406
767 | 0.12888 | 0.41034 | 0.0016672
768 | 0.12913 | 0.40988 | 0.0016904
769 | 0.12938 | 0.4094 | 0.0017179
770 | 0.12965 | 0.40901 | 0.0017378
771 | 0.12988 | 0.40859 | 0.0017701
772 | 0.13012 | 0.40816 | 0.0017989
773 | 0.13037 | 0.40774 | 0.0018281
774 | 0.13064 | 0.40739 | 0.0018535
775 | 0.13089 | 0.407 | 0.001888
776 | 0.13116 | 0.40661 | 0.0019187
777 | 0.13141 | 0.40622 | 0.0019543
778 | 0.13159 | 0.40579 | 0.0019907
779 | 0.13178 | 0.40534 | 0.0020091
780 | 0.13198 | 0.40498 | 0.0020464
781 | 0.13218 | 0.40459 | 0.0020797
782 | 0.13238 | 0.4042 | 0.0021184
783 | 0.1326 | 0.40381 | 0.0021478
784 | 0.13278 | 0.40343 | 0.0021777
785 | 0.133 | 0.40297 | 0.0021979
786 | 0.13319 | 0.40258 | 0.0022336
787 | 0.13339 | 0.40219 | 0.0022699
788 | 0.13359 | 0.4018 | 0.0022961
789 | 0.13379 | 0.40138 | 0.0023281
790 | 0.13401 | 0.40096 | 0.0023605
791 | 0.13419 | 0.40057 | 0.0023988
792 | 0.13439 | 0.40021 | 0.0024322
793 | 0.13461 | 0.39979 | 0.0024717
794 | 0.13478 | 0.3994 | 0.0025003
795 | 0.13499 | 0.39898 | 0.0025351
796 | 0.13516 | 0.39862 | 0.0025645
797 | 0.13536 | 0.3982 | 0.002607
798 | 0.13556 | 0.39781 | 0.002645
799 | 0.13572 | 0.39739 | 0.0026853
800 | 0.13589 | 0.39697 | 0.0027102
801 | 0.13601 | 0.39661 | 0.002741
802 | 0.13615 | 0.39619 | 0.002783
803 | 0.13626 | 0.39569 | 0.002817
804 | 0.13636 | 0.3951 | 0.002846
805 | 0.13646 | 0.39455 | 0.002892
806 | 0.13656 | 0.39409 | 0.002926
807 | 0.13664 | 0.39363 | 0.002959
808 | 0.13673 | 0.39313 | 0.00298
809 | 0.13683 | 0.39263 | 0.003014
810 | 0.13693 | 0.39196 | 0.003039
811 | 0.13702 | 0.39137 | 0.003085
812 | 0.13712 | 0.39074 | 0.003114
813 | 0.1372 | 0.39016 | 0.003148
814 | 0.1373 | 0.38957 | 0.003181
815 | 0.13738 | 0.38894 | 0.003219
816 | 0.13745 | 0.38827 | 0.003252
817 | 0.13749 | 0.38773 | 0.003282
818 | 0.13754 | 0.3871 | 0.003319
819 | 0.13758 | 0.38651 | 0.003361
820 | 0.13762 | 0.38601 | 0.003391
821 | 0.13764 | 0.38551 | 0.003432
822 | 0.13765 | 0.38509 | 0.003462
823 | 0.13767 | 0.38475 | 0.003495
824 | 0.13769 | 0.38425 | 0.003533
825 | 0.13773 | 0.38371 | 0.003566
826 | 0.1378 | 0.38316 | 0.003604
827 | 0.13788 | 0.38253 | 0.003638
828 | 0.13798 | 0.38191 | 0.003667
829 | 0.13806 | 0.38132 | 0.003696
830 | 0.13807 | 0.38073 | 0.00373
831 | 0.13802 | 0.38015 | 0.003768
832 | 0.13791 | 0.37956 | 0.003797
833 | 0.1378 | 0.37897 | 0.003826
834 | 0.13768 | 0.37839 | 0.003855
835 | 0.13761 | 0.3778 | 0.003885
836 | 0.1376 | 0.3773 | 0.003913
837 | 0.13762 | 0.3768 | 0.003957
838 | 0.13766 | 0.37621 | 0.003995
839 | 0.1377 | 0.37558 | 0.004039
840 | 0.13772 | 0.37495 | 0.004067
841 | 0.13775 | 0.37432 | 0.004108
842 | 0.13778 | 0.37378 | 0.004139
843 | 0.1378 | 0.37319 | 0.004168
844 | 0.13779 | 0.37265 | 0.004199
845 | 0.13772 | 0.37194 | 0.004227
846 | 0.13759 | 0.37118 | 0.004252
847 | 0.13742 | 0.3703 | 0.004271
848 | 0.13722 | 0.36934 | 0.00429
849 | 0.13699 | 0.36854 | 0.004303
850 | 0.13681 | 0.36796 | 0.004328
851 | 0.13661 | 0.36761 | 0.004359
852 | 0.13641 | 0.36746 | 0.004387
853 | 0.13621 | 0.36746 | 0.004431
854 | 0.13605 | 0.3673 | 0.004463
855 | 0.13588 | 0.36695 | 0.004497
856 | 0.13572 | 0.36636 | 0.004523
857 | 0.13555 | 0.36557 | 0.004551
858 | 0.13537 | 0.36475 | 0.004585
859 | 0.1352 | 0.36384 | 0.004607
860 | 0.135 | 0.36299 | 0.004614
861 | 0.13479 | 0.36215 | 0.004601
862 | 0.1346 | 0.36117 | 0.004579
863 | 0.13439 | 0.36036 | 0.004532
864 | 0.13419 | 0.35957 | 0.004513
865 | 0.13401 | 0.35897 | 0.004513
866 | 0.13378 | 0.3586 | 0.004535
867 | 0.13356 | 0.35838 | 0.004579
868 | 0.13332 | 0.35828 | 0.004636
869 | 0.13313 | 0.35819 | 0.004695
870 | 0.13299 | 0.358 | 0.004761
871 | 0.133 | 0.35765 | 0.004823
872 | 0.13306 | 0.35728 | 0.004898
873 | 0.13315 | 0.3569 | 0.004972
874 | 0.13326 | 0.35646 | 0.005045
875 | 0.13324 | 0.35596 | 0.005097
876 | 0.13309 | 0.35542 | 0.005125
877 | 0.13285 | 0.35483 | 0.005137
878 | 0.13266 | 0.35414 | 0.005141
879 | 0.13255 | 0.35354 | 0.005143
880 | 0.1325 | 0.35297 | 0.00515
881 | 0.13258 | 0.35253 | 0.005139
882 | 0.1327 | 0.35219 | 0.005131
883 | 0.13287 | 0.35187 | 0.005123
884 | 0.13302 | 0.35147 | 0.005108
885 | 0.133 | 0.35096 | 0.00511
886 | 0.13285 | 0.35027 | 0.005123
887 | 0.13262 | 0.34949 | 0.005139
888 | 0.13236 | 0.34858 | 0.005164
889 | 0.13213 | 0.34776 | 0.005192
890 | 0.13191 | 0.34697 | 0.005234
891 | 0.13174 | 0.34631 | 0.005294
892 | 0.13161 | 0.34569 | 0.005365
893 | 0.13151 | 0.34509 | 0.005449
894 | 0.13148 | 0.34452 | 0.005525
895 | 0.13151 | 0.34396 | 0.005575
896 | 0.13157 | 0.34336 | 0.005583
897 | 0.13165 | 0.3427 | 0.005575
898 | 0.13177 | 0.34209 | 0.005554
899 | 0.13191 | 0.34152 | 0.005541
900 | 0.13203 | 0.34104 | 0.005543
901 | 0.13213 | 0.34052 | 0.005588
902 | 0.13219 | 0.34012 | 0.00565
903 | 0.13221 | 0.33973 | 0.005716
904 | 0.13213 | 0.33936 | 0.005775
905 | 0.132 | 0.33899 | 0.005804
906 | 0.13185 | 0.3386 | 0.00579
907 | 0.13167 | 0.3382 | 0.005748
908 | 0.13145 | 0.33781 | 0.005704
909 | 0.1312 | 0.33744 | 0.005665
910 | 0.131 | 0.33699 | 0.005665
911 | 0.13079 | 0.33665 | 0.005708
912 | 0.13058 | 0.33634 | 0.005769
913 | 0.1304 | 0.33602 | 0.005848
914 | 0.13019 | 0.33563 | 0.005924
915 | 0.12999 | 0.33503 | 0.005963
916 | 0.12977 | 0.3341 | 0.005961
917 | 0.12955 | 0.333 | 0.005928
918 | 0.12933 | 0.33184 | 0.005896
919 | 0.12912 | 0.33074 | 0.005863
920 | 0.129 | 0.32997 | 0.005865
921 | 0.12894 | 0.32961 | 0.005899
922 | 0.12897 | 0.32946 | 0.005955
923 | 0.12901 | 0.32944 | 0.00603
924 | 0.12906 | 0.32932 | 0.006097
925 | 0.12898 | 0.32901 | 0.006133
926 | 0.12886 | 0.32836 | 0.006127
927 | 0.12862 | 0.32749 | 0.006104
928 | 0.12838 | 0.32661 | 0.006075
929 | 0.12815 | 0.32565 | 0.006046
930 | 0.12798 | 0.325 | 0.006039
931 | 0.1279 | 0.32438 | 0.006053
932 | 0.1279 | 0.32398 | 0.006087
933 | 0.12793 | 0.32361 | 0.006133
934 | 0.12797 | 0.32327 | 0.006173
935 | 0.12799 | 0.32299 | 0.006223
936 | 0.12802 | 0.32274 | 0.006267
937 | 0.12804 | 0.32265 | 0.006317
938 | 0.12806 | 0.32251 | 0.00637
939 | 0.12806 | 0.32234 | 0.006413
940 | 0.12799 | 0.32197 | 0.006463
941 | 0.12784 | 0.32132 | 0.006501
942 | 0.12761 | 0.32045 | 0.006537
943 | 0.12737 | 0.31954 | 0.006574
944 | 0.12713 | 0.31866 | 0.006604
945 | 0.127 | 0.31796 | 0.006634
946 | 0.12694 | 0.31742 | 0.00665
947 | 0.12696 | 0.31703 | 0.006663
948 | 0.12702 | 0.31671 | 0.006665
949 | 0.12704 | 0.31635 | 0.006675
950 | 0.12697 | 0.31601 | 0.006682
951 | 0.12672 | 0.31559 | 0.006684
952 | 0.12648 | 0.31521 | 0.006698
953 | 0.12625 | 0.31477 | 0.006698
954 | 0.126 | 0.31441 | 0.006706
955 | 0.12576 | 0.31398 | 0.006742
956 | 0.12553 | 0.31362 | 0.006786
957 | 0.12527 | 0.31321 | 0.006874
958 | 0.12508 | 0.3128 | 0.006948
959 | 0.12484 | 0.31239 | 0.007028
960 | 0.12463 | 0.312 | 0.007102
961 | 0.12441 | 0.31159 | 0.007138
962 | 0.12422 | 0.31116 | 0.007168
963 | 0.12401 | 0.31077 | 0.007182
964 | 0.12385 | 0.31039 | 0.007212
965 | 0.12366 | 0.31 | 0.007241
966 | 0.12348 | 0.30959 | 0.007299
967 | 0.12332 | 0.30918 | 0.007402
968 | 0.12317 | 0.30877 | 0.007483
969 | 0.12304 | 0.30839 | 0.007563
970 | 0.12291 | 0.308 | 0.007607
971 | 0.12284 | 0.30759 | 0.007593
972 | 0.12277 | 0.30721 | 0.007556
973 | 0.12273 | 0.3068 | 0.007497
974 | 0.12272 | 0.30639 | 0.007446
975 | 0.12274 | 0.30601 | 0.007424
976 | 0.12279 | 0.30557 | 0.007424
977 | 0.12284 | 0.30519 | 0.007461
978 | 0.12293 | 0.3048 | 0.007497
979 | 0.123 | 0.30439 | 0.007541
980 | 0.1231 | 0.30398 | 0.007578
981 | 0.12316 | 0.30355 | 0.007615
982 | 0.12323 | 0.30314 | 0.007681
983 | 0.12328 | 0.30268 | 0.007717
984 | 0.12332 | 0.30234 | 0.007783
985 | 0.12332 | 0.30198 | 0.007835
986 | 0.12331 | 0.30174 | 0.007915
987 | 0.12327 | 0.30157 | 0.007996
988 | 0.12323 | 0.30143 | 0.008062
989 | 0.12318 | 0.30121 | 0.008135
990 | 0.12312 | 0.30102 | 0.008216
991 | 0.12305 | 0.30068 | 0.008296
992 | 0.12298 | 0.30023 | 0.008377
993 | 0.12292 | 0.29979 | 0.008465
994 | 0.12286 | 0.29936 | 0.008546
995 | 0.12278 | 0.29902 | 0.008604
996 | 0.12272 | 0.29873 | 0.008619
997 | 0.12266 | 0.29852 | 0.008582
998 | 0.12261 | 0.29832 | 0.008582
999 | 0.12258 | 0.29818 | 0.008575
1000 | 0.12255 | 0.29799 | 0.00856
1001 | 0.12255 | 0.2978 | 0.008578
1002 | 0.12255 | 0.29761 | 0.008606
1003 | 0.12259 | 0.29742 | 0.00865
1004 | 0.12264 | 0.29719 | 0.008701
1005 | 0.12272 | 0.29698 | 0.008754
1006 | 0.1228 | 0.2968 | 0.008807
1007 | 0.12287 | 0.29659 | 0.008867
1008 | 0.12297 | 0.29639 | 0.008933
1009 | 0.12305 | 0.2962 | 0.008996
1010 | 0.12313 | 0.29599 | 0.009052
1011 | 0.1232 | 0.29581 | 0.009084
1012 | 0.12327 | 0.29563 | 0.009096
1013 | 0.12332 | 0.29546 | 0.009118
1014 | 0.12336 | 0.29524 | 0.009147
1015 | 0.12338 | 0.295 | 0.009184
1016 | 0.12339 | 0.29468 | 0.00925
1017 | 0.12336 | 0.29428 | 0.009351
1018 | 0.12331 | 0.29383 | 0.009439
1019 | 0.12325 | 0.29339 | 0.009527
1020 | 0.12317 | 0.29298 | 0.00958
1021 | 0.12308 | 0.29255 | 0.00958
1022 | 0.12298 | 0.29207 | 0.009536
1023 | 0.12286 | 0.29163 | 0.009483
1024 | 0.12275 | 0.29125 | 0.009442
1025 | 0.12264 | 0.29098 | 0.009432
1026 | 0.12253 | 0.29087 | 0.009483
1027 | 0.1224 | 0.29091 | 0.009564
1028 | 0.12228 | 0.29097 | 0.009659
1029 | 0.12217 | 0.29105 | 0.009756
1030 | 0.12206 | 0.29102 | 0.009838
1031 | 0.12197 | 0.29087 | 0.00991
1032 | 0.12189 | 0.29067 | 0.009966
1033 | 0.1218 | 0.29042 | 0.010007
1034 | 0.12173 | 0.29018 | 0.010067
1035 | 0.12166 | 0.28998 | 0.010142
1036 | 0.1216 | 0.28977 | 0.01023
1037 | 0.12154 | 0.28959 | 0.01034
1038 | 0.12148 | 0.2894 | 0.01045
1039 | 0.12144 | 0.28919 | 0.010541
1040 | 0.12139 | 0.28901 | 0.010635
1041 | 0.12134 | 0.28879 | 0.010697
1042 | 0.1213 | 0.28857 | 0.010752
1043 | 0.12126 | 0.28831 | 0.010794
1044 | 0.12122 | 0.28814 | 0.010836
1045 | 0.12119 | 0.288 | 0.010882
1046 | 0.12116 | 0.28794 | 0.010911
1047 | 0.12112 | 0.28796 | 0.010928
1048 | 0.1211 | 0.28803 | 0.010945
1049 | 0.12108 | 0.28806 | 0.010965
1050 | 0.12106 | 0.288 | 0.011007
1051 | 0.12104 | 0.28785 | 0.011092
1052 | 0.12104 | 0.28759 | 0.011176
1053 | 0.12102 | 0.28731 | 0.011285
1054 | 0.12103 | 0.28708 | 0.011385
1055 | 0.12105 | 0.287 | 0.011457
1056 | 0.12108 | 0.2871 | 0.011494
1057 | 0.12111 | 0.2874 | 0.011503
1058 | 0.12115 | 0.28772 | 0.011511
1059 | 0.1212 | 0.28796 | 0.011519
1060 | 0.12125 | 0.288 | 0.011578
1061 | 0.12131 | 0.28776 | 0.011666
1062 | 0.12138 | 0.28732 | 0.011813
1063 | 0.12145 | 0.28678 | 0.011972
1064 | 0.12152 | 0.2863 | 0.012127
1065 | 0.12159 | 0.286 | 0.012202
1066 | 0.12167 | 0.28587 | 0.01221
1067 | 0.12176 | 0.28583 | 0.012177
1068 | 0.12184 | 0.28584 | 0.012127
1069 | 0.12192 | 0.2859 | 0.012093
1070 | 0.122 | 0.286 | 0.012123
1071 | 0.12211 | 0.28615 | 0.012257
1072 | 0.1222 | 0.28638 | 0.012437
1073 | 0.12229 | 0.28663 | 0.012667
1074 | 0.12237 | 0.28685 | 0.012864
1075 | 0.12246 | 0.287 | 0.01301
1076 | 0.12257 | 0.28707 | 0.013073
1077 | 0.12266 | 0.28707 | 0.013086
1078 | 0.12276 | 0.28704 | 0.013077
1079 | 0.12286 | 0.28701 | 0.013065
1080 | 0.12294 | 0.287 | 0.013094
1081 | 0.12305 | 0.287 | 0.013161
1082 | 0.12314 | 0.28705 | 0.013263
1083 | 0.12325 | 0.28708 | 0.013358
1084 | 0.12336 | 0.28707 | 0.01345
1085 | 0.12345 | 0.287 | 0.013521
1086 | 0.12355 | 0.28685 | 0.013555
1087 | 0.12365 | 0.28663 | 0.013568
1088 | 0.12375 | 0.28637 | 0.013572
1089 | 0.12386 | 0.28615 | 0.013584
1090 | 0.12396 | 0.28599 | 0.013614
1091 | 0.12406 | 0.28592 | 0.013655
1092 | 0.12416 | 0.28593 | 0.013731
1093 | 0.12427 | 0.28595 | 0.01381
1094 | 0.12437 | 0.28598 | 0.013886
1095 | 0.12448 | 0.286 | 0.013961
1096 | 0.12459 | 0.28593 | 0.014045
1097 | 0.12469 | 0.28588 | 0.014129
1098 | 0.12479 | 0.28585 | 0.014229
1099 | 0.12489 | 0.28588 | 0.01433
1100 | 0.12501 | 0.28598 | 0.014443
1105 | 0.12615 | 0.28709 | 0.015133
1110 | 0.12728 | 0.28817 | 0.015804
1115 | 0.12838 | 0.28916 | 0.016319
1120 | 0.12954 | 0.29024 | 0.016146
1125 | 0.13061 | 0.29133 | 0.016955
1130 | 0.13177 | 0.29236 | 0.017201
1135 | 0.1329 | 0.29344 | 0.017676
1140 | 0.134 | 0.29453 | 0.018112
1145 | 0.13519 | 0.29567 | 0.01858
1150 | 0.13638 | 0.29681 | 0.019757
1155 | 0.13757 | 0.29797 | 0.019855
1160 | 0.13886 | 0.29914 | 0.021032
1165 | 0.14005 | 0.30034 | 0.021185
1170 | 0.1414 | 0.30161 | 0.022121
1175 | 0.14269 | 0.30289 | 0.022412
1180 | 0.14401 | 0.30419 | 0.022651
1185 | 0.14545 | 0.30555 | 0.024237
1190 | 0.14686 | 0.30691 | 0.024461
1195 | 0.14846 | 0.30843 | 0.02592
1200 | 0.15006 | 0.31 | 0.026
1250 | 0.1686 | 0.3276 | 0.0348
1300 | 0.198 | 0.3533 | 0.0483
1350 | 0.2268 | 0.3855 | 0.0623
1400 | 0.2538 | 0.4185 | 0.0781
1450 | 0.2814 | 0.4523 | 0.0919
1500 | 0.306 | 0.4764 | 0.1029
1550 | 0.3217 | 0.4948 | 0.1114
1600 | 0.3273 | 0.5065 | 0.115
1650 | 0.3292 | 0.5116 | 0.1143
1700 | 0.33 | 0.5109 | 0.1136
1750 | 0.33 | 0.5109 | 0.1136
1800 | 0.3317 | 0.5153 | 0.1157
1850 | 0.3368 | 0.5219 | 0.12
1900 | 0.3474 | 0.5292 | 0.1256
1950 | 0.3581 | 0.5417 | 0.1349
2000 | 0.37 | 0.5534 | 0.1455
2050 | 0.3845 | 0.5659 | 0.159
2100 | 0.4 | 0.5769 | 0.1718
2150 | 0.414 | 0.5871 | 0.1876
2200 | 0.43 | 0.5974 | 0.2037
2250 | 0.4442 | 0.6084 | 0.2186
2300 | 0.46 | 0.62 | 0.2378
2350 | 0.4731 | 0.6289 | 0.2555
2400 | 0.4857 | 0.6355 | 0.2719
2450 | 0.4982 | 0.6421 | 0.2853
2500 | 0.512 | 0.6472 | 0.3031
2550 | 0.5259 | 0.6516 | 0.3186
2600 | 0.534 | 0.6538 | 0.3308
2650 | 0.5347 | 0.6531 | 0.3354
2700 | 0.5 | 0.61 | 0.3265
2750 | 0.271 | 0.275 | 0.17
2800 | 0.15 | 0.1459 | 0.0894
2850 | 0.1404 | 0.1319 | 0.0823
2900 | 0.1514 | 0.1466 | 0.0916
2950 | 0.17 | 0.1679 | 0.1043
3000 | 0.19 | 0.19 | 0.1171
3050 | 0.2112 | 0.2184 | 0.1313
3100 | 0.2347 | 0.2448 | 0.1469
3150 | 0.2582 | 0.2727 | 0.1675
3200 | 0.28 | 0.3 | 0.186
3250 | 0.304 | 0.3262 | 0.204
3300 | 0.3149 | 0.3372 | 0.2129
3350 | 0.2923 | 0.32 | 0.2044
3400 | 0.25 | 0.27 | 0.1661
3450 | 0.19 | 0.2052 | 0.1384
3500 | 0.13 | 0.14 | 0.0944
3550 | 0.0877 | 0.0938 | 0.0625
3600 | 0.06 | 0.07 | 0.04
3650 | 0.0511 | 0.0564 | 0.0327
3700 | 0.06 | 0.07 | 0.0355
3750 | 0.081 | 0.1012 | 0.0483
3800 | 0.1055 | 0.1327 | 0.066
3850 | 0.1231 | 0.1539 | 0.0774
3900 | 0.1181 | 0.14 | 0.0759
3950 | 0.0986 | 0.1121 | 0.0666
4000 | 0.08 | 0.094406 | 0.0559
4050 | 0.0781 | 0.09 | 0.051286
4100 | 0.08 | 0.094406 | 0.053
4150 | 0.085901 | 0.097724 | 0.0561
4200 | 0.086497 | 0.1 | 0.0567
4250 | 0.08356 | 0.093325 | 0.054954
4300 | 0.064565 | 0.0733 | 0.047863
4350 | 0.049659 | 0.056754 | 0.035481
4400 | 0.036983 | 0.04 | 0.025704
4450 | 0.0259 | 0.025003 | 0.018621
4500 | 0.016368 | 0.016106 | 0.012882
4550 | 0.01 | 0.01 | 0.0077625
4600 | 0.006223 | 0.006166 | 0.0051286
4650 | 0.0037844 | 0.004 | 0.00302
4700 | 0.0020893 | 0.0021038 | 0.0018197
4750 | 0.0011858 | 0.0011535 | 0.001122
4800 | 0.00077268 | 0.00072611 | 0.00063096
4850 | 0.00051761 | 0.00046881 | 0.00041687
4900 | 0.00031477 | 0.00027861 | 0.00024547
4950 | 0.00014928 | 0.00013932 | 0.00012303
5000 | 5.7544e-05 | 5.2119e-05 | 5.4954e-05
5050 | 1.8197e-05 | 1.8197e-05 | 1.9953e-05
5100 | 1e-05 | 1e-05 | 1e-05
5150 | 1e-05 | 1e-05 | 1e-05
5200 | 1e-05 | 1e-05 | 1e-05
Change Log
2/19/2020 - Updated LCM damage threshold to 0.1 W/mm at 532 nm and 100 mJ/cm^2 at 1064 nm 10 ns.
9/27/2019 - Added response for 1310 nm.
7/11/2019 - Updated IR-BB wavelength response
6/18/2019 - Added TaperCamD option.
6/17/2019 - Updated 1550 nm sensitivity (preliminary).
4/26/2019 - Changed ND-5 back to using NG4 glass
4/25/2019 - Changed ND-5 filter to use NG1 glass
4/22/2019 - Added 1550 nm line to Standard sensor (Needs adjustment)
4/2019 - Added ability to change focus z location by changing cell D18

## IR-BB Saturation (hidden)
Wavelength (microns) | Relative Response | Saturation Irradiance (mW/mm^2) | Saturation Irradiance (mW/cm^2)
0.19 |  |  | 10000000000
1.55 |  |  | 10000000000
1.7 | 57.1258 | 2.5385 | 253.8454
3 | 57.1258 | 2.5385 | 253.8454
3.01 | 57.4423 | 2.5245 | 252.4465
3.01 | 57.7275 | 2.512 | 251.1993
3.02 | 57.8882 | 2.505 | 250.5023
3.03 | 58.017 | 2.4995 | 249.9458
3.04 | 58.2542 | 2.4893 | 248.9281
3.04 | 58.4826 | 2.4796 | 247.9562
3.05 | 58.6794 | 2.4712 | 247.1245
3.06 | 58.8046 | 2.466 | 246.5983
3.06 | 58.9778 | 2.4587 | 245.8742
3.07 | 59.2153 | 2.4489 | 244.8878
3.08 | 59.4471 | 2.4393 | 243.9333
3.09 | 59.6797 | 2.4298 | 242.9825
3.09 | 59.9221 | 2.42 | 241.9992
3.1 | 60.0938 | 2.4131 | 241.3082
3.11 | 60.2933 | 2.4051 | 240.5094
3.12 | 60.5574 | 2.3946 | 239.4608
3.12 | 60.7554 | 2.3868 | 238.6801
3.13 | 61.1057 | 2.3731 | 237.3118
3.14 | 61.4044 | 2.3616 | 236.1574
3.15 | 61.5747 | 2.355 | 235.5043
3.15 | 61.8842 | 2.3433 | 234.3266
3.16 | 62.091 | 2.3355 | 233.5461
3.17 | 62.3561 | 2.3255 | 232.5533
3.18 | 62.5835 | 2.3171 | 231.7084
3.19 | 62.9109 | 2.305 | 230.5024
3.19 | 63.1376 | 2.2967 | 229.6749
3.2 | 63.4691 | 2.2848 | 228.4753
3.21 | 63.8661 | 2.2705 | 227.055
3.22 | 64.1487 | 2.2605 | 226.0545
3.22 | 64.4287 | 2.2507 | 225.0723
3.23 | 64.6652 | 2.2425 | 224.2491
3.24 | 65.0152 | 2.2304 | 223.042
3.25 | 65.2776 | 2.2215 | 222.1454
3.26 | 65.678 | 2.2079 | 220.791
3.27 | 66.1187 | 2.1932 | 219.3195
3.27 | 66.3673 | 2.185 | 218.498
3.28 | 66.8239 | 2.17 | 217.0048
3.29 | 67.2358 | 2.1568 | 215.6755
3.3 | 67.6463 | 2.1437 | 214.3668
3.31 | 68.1411 | 2.1281 | 212.8102
3.32 | 68.6254 | 2.1131 | 211.3081
3.32 | 69.0232 | 2.1009 | 210.0905
3.33 | 69.645 | 2.0821 | 208.2148
3.34 | 70.1644 | 2.0667 | 206.6733
3.35 | 70.6719 | 2.0519 | 205.1891
3.36 | 71.1945 | 2.0368 | 203.6832
3.37 | 71.4477 | 2.0296 | 202.9611
3.38 | 71.6332 | 2.0244 | 202.4355
3.38 | 72.067 | 2.0122 | 201.2172
3.39 | 72.5056 | 2 | 200
3.4 | 72.9923 | 1.9867 | 198.6663
3.41 | 72.7159 | 1.9942 | 199.4216
3.42 | 71.6549 | 2.0237 | 202.3744
3.43 | 71.4392 | 2.0299 | 202.9854
3.44 | 73.8379 | 1.9639 | 196.3911
3.45 | 75.1003 | 1.9309 | 193.0899
3.46 | 75.9922 | 1.9082 | 190.8237
3.47 | 76.7657 | 1.889 | 188.9008
3.48 | 77.3546 | 1.8746 | 187.4628
3.48 | 77.7065 | 1.8661 | 186.6138
3.49 | 78.0344 | 1.8583 | 185.8299
3.5 | 77.6734 | 1.8669 | 186.6934
3.51 | 76.9147 | 1.8854 | 188.5351
3.52 | 78.5042 | 1.8472 | 184.7177
3.53 | 79.7213 | 1.819 | 181.8975
3.54 | 80.2507 | 1.807 | 180.6977
3.55 | 80.657 | 1.7979 | 179.7873
3.56 | 81.0354 | 1.7895 | 178.9478
3.57 | 81.2947 | 1.7838 | 178.3771
3.58 | 81.5926 | 1.7773 | 177.7259
3.59 | 81.7004 | 1.7749 | 177.4913
3.6 | 81.8769 | 1.7711 | 177.1088
3.61 | 82.0191 | 1.768 | 176.8016
3.62 | 82.1853 | 1.7644 | 176.4441
3.63 | 82.3649 | 1.7606 | 176.0595
3.64 | 82.5278 | 1.7571 | 175.7119
3.65 | 82.6991 | 1.7535 | 175.3479
3.66 | 82.8525 | 1.7502 | 175.0232
3.67 | 82.9752 | 1.7476 | 174.7645
3.68 | 83.0825 | 1.7454 | 174.5388
3.69 | 83.1957 | 1.743 | 174.3012
3.7 | 83.2362 | 1.7422 | 174.2165
3.71 | 83.1955 | 1.743 | 174.3017
3.73 | 83.2593 | 1.7417 | 174.1681
3.74 | 83.1475 | 1.744 | 174.4023
3.75 | 83.018 | 1.7467 | 174.6743
3.76 | 82.898 | 1.7493 | 174.9272
3.77 | 82.6118 | 1.7553 | 175.5332
3.78 | 82.3933 | 1.76 | 175.9987
3.79 | 82.0502 | 1.7673 | 176.7346
3.8 | 81.7024 | 1.7749 | 177.487
3.81 | 81.3796 | 1.7819 | 178.1911
3.82 | 80.9905 | 1.7905 | 179.0471
3.84 | 80.4803 | 1.8018 | 180.1822
3.85 | 79.9865 | 1.8129 | 181.2945
3.86 | 79.5026 | 1.824 | 182.3979
3.87 | 78.8931 | 1.8381 | 183.8071
3.88 | 78.2858 | 1.8523 | 185.2329
3.89 | 77.6117 | 1.8684 | 186.8419
3.9 | 76.9501 | 1.8845 | 188.4482
3.92 | 76.1458 | 1.9044 | 190.4387
3.93 | 75.2803 | 1.9263 | 192.6281
3.94 | 74.3657 | 1.95 | 194.9975
3.95 | 73.4 | 1.9756 | 197.563
3.96 | 72.3876 | 2.0033 | 200.326
3.98 | 71.2286 | 2.0359 | 203.5857
3.99 | 70.1252 | 2.0679 | 206.789
4 | 68.9057 | 2.1045 | 210.4487
4.01 | 67.6517 | 2.1435 | 214.3495
4.03 | 66.3016 | 2.1871 | 218.7144
4.04 | 64.9163 | 2.2338 | 223.3817
4.05 | 63.4363 | 2.2859 | 228.5935
4.06 | 61.9268 | 2.3417 | 234.1653
4.08 | 60.3565 | 2.4026 | 240.2578
4.09 | 58.7478 | 2.4684 | 246.8368
4.1 | 57.0482 | 2.5419 | 254.1907
4.12 | 55.3376 | 2.6205 | 262.048
4.13 | 53.5944 | 2.7057 | 270.5715
4.14 | 51.8166 | 2.7985 | 279.8546
4.16 | 50.0496 | 2.8974 | 289.7351
4.17 | 48.3331 | 3.0002 | 300.0245
4.18 | 46.7259 | 3.1034 | 310.344
4.2 | 44.9128 | 3.2287 | 322.8728
4.21 | 42.9674 | 3.3749 | 337.4911
4.22 | 40.5441 | 3.5766 | 357.6628
4.24 | 36.9539 | 3.9241 | 392.4105
4.25 | 34.9796 | 4.1456 | 414.5596
4.26 | 33.6502 | 4.3094 | 430.9369
4.28 | 31.5204 | 4.6005 | 460.0543
4.29 | 29.386 | 4.9347 | 493.4698
4.31 | 27.9616 | 5.1861 | 518.6089
4.32 | 26.4981 | 5.4725 | 547.2521
4.34 | 25.2308 | 5.7474 | 574.7378
4.35 | 23.8304 | 6.0851 | 608.5131
4.36 | 22.4365 | 6.4632 | 646.3168
4.38 | 21.0435 | 6.891 | 689.1033
4.39 | 19.5831 | 7.4049 | 740.4931
4.41 | 18.1861 | 7.9737 | 797.3715
4.42 | 16.7554 | 8.6546 | 865.4603
4.44 | 15.476 | 9.3701 | 937.0076
4.45 | 14.2389 | 10 | 1000
4.47 | 13.0611 | 10 | 1000
4.49 | 12.0123 | 10 | 1000
4.5 | 10.9824 | 10 | 1000
4.52 | 10.005 | 10 | 1000
4.53 | 9.1188 | 10 | 1000
4.55 | 8.3378 | 10 | 1000
4.56 | 7.5396 | 10 | 1000
4.58 | 6.866 | 10 | 1000
4.6 | 6.2063 | 10 | 1000
4.61 | 5.624 | 10 | 1000
4.63 | 5.0977 | 10 | 1000
4.65 | 4.707 | 10 | 1000
4.66 | 4.3986 | 10 | 1000
4.68 | 4.1446 | 10 | 1000
4.7 | 4.0146 | 10 | 1000
4.71 | 3.9403 | 10 | 1000
4.73 | 3.8947 | 10 | 1000
4.75 | 3.9483 | 10 | 1000
4.77 | 4.0841 | 10 | 1000
4.78 | 4.268 | 10 | 1000
4.8 | 4.5924 | 10 | 1000
4.82 | 4.9981 | 10 | 1000
4.84 | 5.4143 | 10 | 1000
4.86 | 5.886 | 10 | 1000
4.87 | 6.3667 | 10 | 1000
4.89 | 6.8435 | 10 | 1000
4.91 | 7.3027 | 10 | 1000
4.93 | 7.8061 | 10 | 1000
4.95 | 8.3466 | 10 | 1000
4.97 | 8.9198 | 10 | 1000
4.99 | 9.5857 | 10 | 1000
5.01 | 10.2897 | 10 | 1000
5.02 | 11.0206 | 10 | 1000
5.04 | 11.8062 | 10 | 1000
5.06 | 12.5639 | 10 | 1000
5.08 | 13.4015 | 10 | 1000
5.1 | 14.2293 | 10 | 1000
5.12 | 15.0238 | 9.6521 | 965.212
5.14 | 15.8992 | 9.1207 | 912.067
5.16 | 16.7388 | 8.6632 | 866.3188
5.19 | 17.5363 | 8.2692 | 826.9202
5.21 | 18.3576 | 7.8992 | 789.9242
5.23 | 19.2333 | 7.5396 | 753.9591
5.25 | 20.069 | 7.2256 | 722.5627
5.27 | 20.8847 | 6.9434 | 694.3407
5.29 | 21.7087 | 6.6799 | 667.9875
5.31 | 22.5344 | 6.4351 | 643.5105
5.33 | 23.3586 | 6.208 | 620.8034
5.36 | 24.2384 | 5.9827 | 598.2706
5.38 | 25.1668 | 5.762 | 576.1997
5.4 | 26.086 | 5.559 | 555.8965
5.42 | 27.094 | 5.3522 | 535.2154
5.45 | 28.1548 | 5.1505 | 515.0489
5.47 | 29.2855 | 4.9516 | 495.1635
5.49 | 30.4249 | 4.7662 | 476.6205
5.52 | 31.4764 | 4.607 | 460.6978
5.54 | 32.4647 | 4.4667 | 446.674
5.56 | 33.4105 | 4.3403 | 434.0289
5.59 | 34.3959 | 4.2159 | 421.5939
5.61 | 35.1849 | 4.1214 | 412.1407
5.64 | 35.9926 | 4.0289 | 402.8919
5.66 | 36.833 | 3.937 | 393.6987
5.69 | 37.6061 | 3.8561 | 385.6055
5.71 | 38.4119 | 3.7752 | 377.5157
5.74 | 39.3017 | 3.6897 | 368.9692
5.76 | 40.1803 | 3.609 | 360.901
5.79 | 41.0489 | 3.5326 | 353.2645
5.81 | 41.9459 | 3.4571 | 345.7101
5.84 | 42.8978 | 3.3804 | 338.0385
5.87 | 43.9487 | 3.2996 | 329.9554
5.89 | 45.2632 | 3.2037 | 320.3728
5.92 | 46.5674 | 3.114 | 311.4008
5.95 | 47.838 | 3.0313 | 303.1293
5.97 | 49.2944 | 2.9417 | 294.1734
6 | 50.4875 | 2.8722 | 287.2216
6.03 | 51.6275 | 2.8088 | 280.8796
6.06 | 52.8673 | 2.7429 | 274.2926
6.09 | 54.3863 | 2.6663 | 266.6315
6.11 | 55.6565 | 2.6055 | 260.5466
6.14 | 56.9537 | 2.5461 | 254.6122
6.17 | 58.3591 | 2.4848 | 248.4807
6.2 | 59.7986 | 2.425 | 242.4991
6.23 | 61.1079 | 2.373 | 237.3033
6.26 | 62.2391 | 2.3299 | 232.9905
6.29 | 63.2305 | 2.2934 | 229.3374
6.32 | 64.101 | 2.2622 | 226.2227
6.35 | 64.8916 | 2.2347 | 223.4668
6.39 | 65.508 | 2.2136 | 221.364
6.42 | 65.7302 | 2.2062 | 220.6159
6.45 | 65.5901 | 2.2109 | 221.0868
6.48 | 64.5652 | 2.246 | 224.5963
6.51 | 64.8479 | 2.2362 | 223.6173
6.55 | 67.4118 | 2.1511 | 215.1125
6.58 | 68.8537 | 2.1061 | 210.6077
6.61 | 70.1788 | 2.0663 | 206.6309
6.65 | 71.5177 | 2.0276 | 202.7626
6.68 | 72.5924 | 1.9976 | 199.7607
6.72 | 73.5093 | 1.9727 | 197.269
6.75 | 74.2975 | 1.9518 | 195.1763
6.79 | 74.5672 | 1.9447 | 194.4704
6.82 | 74.8015 | 1.9386 | 193.8612
6.86 | 75.5172 | 1.9202 | 192.024
6.9 | 76.4135 | 1.8977 | 189.7716
6.93 | 76.939 | 1.8848 | 188.4755
6.97 | 77.5223 | 1.8706 | 187.0572
7.01 | 78.1882 | 1.8546 | 185.4641
7.05 | 78.6893 | 1.8428 | 184.2833
7.08 | 79.3061 | 1.8285 | 182.85
7.12 | 79.8442 | 1.8162 | 181.6177
7.16 | 80.5783 | 1.7996 | 179.9629
7.2 | 81.9514 | 1.7695 | 176.9478
7.24 | 82.9197 | 1.7488 | 174.8815
7.28 | 83.7076 | 1.7324 | 173.2354
7.32 | 84.4187 | 1.7178 | 171.7761
7.37 | 84.9922 | 1.7062 | 170.617
7.41 | 85.5871 | 1.6943 | 169.4311
7.45 | 86.1387 | 1.6835 | 168.346
7.49 | 86.7975 | 1.6707 | 167.0684
7.54 | 87.34 | 1.6603 | 166.0307
7.58 | 87.7469 | 1.6526 | 165.2607
7.63 | 88.0595 | 1.6467 | 164.6741
7.67 | 88.3344 | 1.6416 | 164.1616
7.72 | 88.6453 | 1.6359 | 163.5859
7.76 | 88.9807 | 1.6297 | 162.9691
7.81 | 89.2031 | 1.6256 | 162.5629
7.86 | 89.2586 | 1.6246 | 162.4618
7.9 | 89.1791 | 1.6261 | 162.6067
7.95 | 89.0408 | 1.6286 | 162.8591
8 | 89.0389 | 1.6286 | 162.8628
8.05 | 89.1129 | 1.6273 | 162.7274
8.1 | 89.2858 | 1.6241 | 162.4123
8.15 | 89.4741 | 1.6207 | 162.0705
8.2 | 89.6011 | 1.6184 | 161.8408
8.26 | 89.7042 | 1.6165 | 161.6548
8.31 | 89.8346 | 1.6142 | 161.4201
8.36 | 89.9036 | 1.613 | 161.2962
8.42 | 89.9232 | 1.6126 | 161.261
8.47 | 89.9165 | 1.6127 | 161.2731
8.53 | 89.9635 | 1.6119 | 161.1889
8.59 | 89.9697 | 1.6118 | 161.1777
8.64 | 90.0451 | 1.6104 | 161.0428
8.7 | 90.1136 | 1.6092 | 160.9204
8.76 | 90.2967 | 1.6059 | 160.5941
8.82 | 90.4638 | 1.603 | 160.2974
8.88 | 90.8608 | 1.596 | 159.5971
8.94 | 91.3491 | 1.5874 | 158.7439
9 | 91.7416 | 1.5806 | 158.0648
9.07 | 91.9387 | 1.5773 | 157.7259
9.13 | 91.8071 | 1.5795 | 157.952
9.19 | 91.5701 | 1.5836 | 158.3608
9.26 | 91.1968 | 1.5901 | 159.009
9.33 | 90.8984 | 1.5953 | 159.5309
9.39 | 90.5538 | 1.6014 | 160.1381
9.46 | 90.2323 | 1.6071 | 160.7086
9.53 | 89.9046 | 1.6129 | 161.2944
9.6 | 89.6949 | 1.6167 | 161.6716
9.67 | 89.555 | 1.6192 | 161.9242
9.75 | 89.2121 | 1.6255 | 162.5465
9.82 | 88.9559 | 1.6301 | 163.0147
9.9 | 88.6051 | 1.6366 | 163.6601
9.97 | 88.2036 | 1.644 | 164.405
10.05 | 87.6073 | 1.6552 | 165.5241
10.13 | 87.0583 | 1.6657 | 166.5678
10.21 | 86.4597 | 1.6772 | 167.7211
10.29 | 85.7532 | 1.691 | 169.1029
10.37 | 85.0243 | 1.7055 | 170.5526
10.46 | 84.2173 | 1.7219 | 172.1869
10.54 | 83.2603 | 1.7417 | 174.1661
10.63 | 82.2781 | 1.7625 | 176.2452
10.71 | 81.1559 | 1.7868 | 178.6822
10.8 | 80.0765 | 1.8109 | 181.0908
10.89 | 79.0014 | 1.8356 | 183.5551
10.99 | 77.9785 | 1.8596 | 185.963
11.08 | 77.0777 | 1.8814 | 188.1363
11.18 | 76.0947 | 1.9057 | 190.5667
11.27 | 75.2722 | 1.9265 | 192.6489
11.37 | 74.3842 | 1.9495 | 194.9487
11.47 | 73.5323 | 1.9721 | 197.2075
11.58 | 72.5329 | 1.9992 | 199.9247
11.68 | 71.512 | 2.0278 | 202.7787
11.79 | 70.5113 | 2.0566 | 205.6566
11.89 | 69.8631 | 2.0756 | 207.5648
12 | 69.5856 | 2.0839 | 208.3925
12.12 | 69.5179 | 2.086 | 208.5954
12.23 | 69.497 | 2.0866 | 208.6581
12.35 | 69.521 | 2.0859 | 208.5862
12.47 | 69.6578 | 2.0818 | 208.1765
12.59 | 69.9366 | 2.0735 | 207.3466
12.71 | 70.3296 | 2.0619 | 206.1878
12.84 | 70.5192 | 2.0563 | 205.6336
12.96 | 70.7072 | 2.0509 | 205.0869
13.1 | 70.6499 | 2.0525 | 205.2531
13.23 | 70.3937 | 2.06 | 206.0003
13.37 | 70.0747 | 2.0694 | 206.938
13.5 | 69.8024 | 2.0775 | 207.7453
13.65 | 69.2641 | 2.0936 | 209.3598
13.79 | 68.4183 | 2.1195 | 211.948
13.94 | 67.506 | 2.1481 | 214.8122
14.09 | 66.4244 | 2.1831 | 218.3102
14.25 | 64.931 | 2.2333 | 223.331
14.4 | 63.2122 | 2.294 | 229.4037
14.57 | 61.1748 | 2.3704 | 237.044
14.73 | 59.3639 | 2.4428 | 244.2751
14.9 | 57.6011 | 2.5175 | 251.7506
15.07 | 56.1965 | 2.5804 | 258.0431
15.25 | 54.6422 | 2.6538 | 265.3831
15.43 | 52.8959 | 2.7414 | 274.1444
15.62 | 51.2945 | 2.827 | 282.7034
15.81 | 50.0181 | 2.8992 | 289.9171
16.01 | 48.642 | 2.9812 | 298.1194
16.21 | 47.1452 | 3.0758 | 307.5839
16.41 | 45.327 | 3.1992 | 319.9225
16.62 | 43.0593 | 3.3677 | 336.7708
16.84 | 39.9911 | 3.6261 | 362.6088
17.06 | 36.3659 | 3.9876 | 398.7558
17.29 | 32.4252 | 4.4722 | 447.2172
17.52 | 28.8088 | 5.0336 | 503.3565
17.76 | 25.9361 | 5.5911 | 559.11
18.01 | 24.2182 | 5.9877 | 598.7701
18.26 | 22.9648 | 6.3145 | 631.4498
18.52 | 21.4765 | 6.7521 | 675.2088
18.79 | 20.2179 | 7.1724 | 717.2409
19.06 | 19.141 | 7.5759 | 757.5933
19.35 | 18.8176 | 7.7061 | 770.6149
19.64 | 19.0691 | 7.6045 | 760.4498
19.94 | 19.559 | 7.414 | 741.4042

## CBS (hidden)
Wavelength (nm) | Reflection (0 to 1) at 10 degrees. AVG Pol. | Transmission
300 | 0.16677 | 1e-07
300.2 | 0.1215 | 1e-07 | First | Transmission | 1
300.4 | 0.068036 | 1e-07 | Second | Transmission | 1
300.6 | 0.027092 | 1e-07 | Third | Transmission | 1
300.8 | 0.025497 | 1e-07
301 | 0.079108 | 1e-07
301.2 | 0.17762 | 1e-07
301.4 | 0.29301 | 1e-07
301.6 | 0.40064 | 1e-07
301.8 | 0.48878 | 1e-07
302 | 0.55545 | 1e-07
302.2 | 0.60292 | 1e-07
302.4 | 0.63425 | 1e-07
302.6 | 0.65188 | 1e-07
302.8 | 0.65722 | 1e-07
303 | 0.65054 | 1e-07
303.2 | 0.63095 | 1e-07
303.4 | 0.59628 | 1e-07
303.6 | 0.54314 | 1e-07
303.8 | 0.46742 | 1e-07
304 | 0.36654 | 1e-07
304.2 | 0.2451 | 1e-07
304.4 | 0.12369 | 1e-07
304.6 | 0.040154 | 1e-07
304.8 | 0.02657 | 1e-07
305 | 0.079053 | 1e-07
305.2 | 0.16398 | 1e-07
305.4 | 0.24871 | 1e-07
305.6 | 0.31614 | 1e-07
305.8 | 0.36114 | 1e-07
306 | 0.38377 | 1e-07
306.2 | 0.38532 | 1e-07
306.4 | 0.3669 | 1e-07
306.6 | 0.32932 | 1e-07
306.8 | 0.27407 | 1e-07
307 | 0.20489 | 1e-07
307.2 | 0.12987 | 1e-07
307.4 | 0.062048 | 1e-07
307.6 | 0.016232 | 1e-07
307.8 | 0.0019823 | 1e-07
308 | 0.018085 | 1e-07
308.2 | 0.054099 | 1e-07
308.4 | 0.096943 | 1e-07
308.6 | 0.13637 | 1e-07
308.8 | 0.1667 | 1e-07
309 | 0.18618 | 1e-07
309.2 | 0.19581 | 1e-07
309.4 | 0.19854 | 1e-07
309.6 | 0.19853 | 1e-07
309.8 | 0.20034 | 1e-07
310 | 0.20756 | 1e-07
310.2 | 0.22159 | 1e-07
310.4 | 0.24094 | 1e-07
310.6 | 0.26188 | 1e-07
310.8 | 0.27963 | 1e-07
311 | 0.28964 | 1e-07
311.2 | 0.28808 | 1e-07
311.4 | 0.27208 | 1e-07
311.6 | 0.2399 | 1e-07
311.8 | 0.19199 | 1e-07
312 | 0.13335 | 1e-07
312.2 | 0.077165 | 1e-07
312.4 | 0.045877 | 1e-07
312.6 | 0.062675 | 1e-07
312.8 | 0.13401 | 1e-07
313 | 0.24195 | 1e-07
313.2 | 0.35806 | 1e-07
313.4 | 0.46142 | 1e-07
313.6 | 0.54359 | 1e-07
313.8 | 0.60427 | 1e-07
314 | 0.64623 | 1e-07
314.2 | 0.67245 | 1e-07
314.4 | 0.68501 | 1e-07
314.6 | 0.68472 | 1e-07
314.8 | 0.67097 | 1e-07
315 | 0.64147 | 1e-07
315.2 | 0.59189 | 1e-07
315.4 | 0.51603 | 1e-07
315.6 | 0.40822 | 1e-07
315.8 | 0.27326 | 1e-07
316 | 0.14667 | 1e-07
316.2 | 0.096334 | 1e-07
316.4 | 0.16006 | 1e-07
316.6 | 0.29697 | 1e-07
316.8 | 0.44014 | 1e-07
317 | 0.55465 | 1e-07
317.2 | 0.63606 | 1e-07
317.4 | 0.69073 | 1e-07
317.6 | 0.72567 | 1e-07
317.8 | 0.74592 | 1e-07
318 | 0.75458 | 1e-07
318.2 | 0.75327 | 1e-07
318.4 | 0.74259 | 1e-07
318.6 | 0.72275 | 1e-07
318.8 | 0.69451 | 1e-07
319 | 0.66095 | 1e-07
319.2 | 0.62976 | 1e-07
319.4 | 0.61317 | 1e-07
319.6 | 0.62156 | 1e-07
319.8 | 0.65452 | 1e-07
320 | 0.70098 | 1e-07
320.2 | 0.74839 | 1e-07
320.4 | 0.78946 | 1e-07
320.6 | 0.82213 | 1e-07
320.8 | 0.84707 | 1e-07
321 | 0.86578 | 1e-07
321.2 | 0.87976 | 1e-07
321.4 | 0.89022 | 1e-07
321.6 | 0.89811 | 1e-07
321.8 | 0.9041 | 1e-07
322 | 0.90867 | 1e-07
322.2 | 0.9121 | 1e-07
322.4 | 0.91455 | 1e-07
322.6 | 0.91602 | 1e-07
322.8 | 0.91638 | 1e-07
323 | 0.91539 | 1e-07
323.2 | 0.91268 | 1e-07
323.4 | 0.90768 | 1e-07
323.6 | 0.89962 | 1e-07
323.8 | 0.88737 | 1e-07
324 | 0.86956 | 1e-07
324.2 | 0.84508 | 1e-07
324.4 | 0.81514 | 1e-07
324.6 | 0.78739 | 1e-07
324.8 | 0.7768 | 1e-07
325 | 0.79337 | 1e-07
325.2 | 0.82935 | 1e-07
325.4 | 0.86787 | 1e-07
325.6 | 0.89887 | 1e-07
325.8 | 0.92089 | 1e-07
326 | 0.93583 | 1e-07
326.2 | 0.94579 | 1e-07
326.4 | 0.95232 | 1e-07
326.6 | 0.95642 | 1e-07
326.8 | 0.95871 | 1e-07
327 | 0.95952 | 1e-07
327.2 | 0.95896 | 1e-07
327.4 | 0.95695 | 1e-07
327.6 | 0.95322 | 1e-07
327.8 | 0.94723 | 1e-07
328 | 0.93816 | 1e-07
328.2 | 0.92525 | 1e-07
328.4 | 0.90925 | 1e-07
328.6 | 0.89557 | 1e-07
328.8 | 0.89411 | 1e-07
329 | 0.90887 | 1e-07
329.2 | 0.93146 | 1e-07
329.4 | 0.9514 | 1e-07
329.6 | 0.96537 | 1e-07
329.8 | 0.97447 | 1e-07
330 | 0.98039 | 1e-07
330.2 | 0.98434 | 1e-07
330.4 | 0.98706 | 1e-07
330.6 | 0.98898 | 1e-07
330.8 | 0.99036 | 1e-07
331 | 0.99137 | 1e-07
331.2 | 0.99208 | 1e-07
331.4 | 0.99256 | 1e-07
331.6 | 0.99285 | 1e-07
331.8 | 0.99295 | 1e-07
332 | 0.99286 | 1e-07
332.2 | 0.99255 | 1e-07
332.4 | 0.99194 | 1e-07
332.6 | 0.9909 | 1e-07
332.8 | 0.98915 | 1e-07
333 | 0.98623 | 1e-07
333.2 | 0.98147 | 1e-07
333.4 | 0.97494 | 1e-07
333.6 | 0.97039 | 1e-07
333.8 | 0.97322 | 1e-07
334 | 0.98074 | 1e-07
334.2 | 0.98704 | 1e-07
334.4 | 0.9909 | 1e-07
334.6 | 0.99314 | 1e-07
334.8 | 0.99449 | 1e-07
335 | 0.99534 | 1e-07
335.2 | 0.99591 | 1e-07
335.4 | 0.99631 | 1e-07
335.6 | 0.9966 | 1e-07
335.8 | 0.99681 | 1e-07
336 | 0.99695 | 1e-07
336.2 | 0.99701 | 1e-07
336.4 | 0.997 | 1e-07
336.6 | 0.9969 | 1e-07
336.8 | 0.99672 | 1e-07
337 | 0.99645 | 1e-07
337.2 | 0.99605 | 1e-07
337.4 | 0.99546 | 1e-07
337.6 | 0.99449 | 1e-07
337.8 | 0.99274 | 1e-07
338 | 0.98934 | 1e-07
338.2 | 0.98313 | 1e-07
338.4 | 0.97742 | 1e-07
338.6 | 0.97956 | 1e-07
338.8 | 0.98707 | 1e-07
339 | 0.99219 | 1e-07
339.2 | 0.99467 | 1e-07
339.4 | 0.99589 | 1e-07
339.6 | 0.9965 | 1e-07
339.8 | 0.99682 | 1e-07
340 | 0.99699 | 1e-07
340.2 | 0.99712 | 1e-07
340.4 | 0.99728 | 1e-07
340.6 | 0.99745 | 1e-07
340.8 | 0.9976 | 1e-07
341 | 0.99766 | 1e-07
341.2 | 0.9976 | 1e-07
341.4 | 0.99734 | 1e-07
341.6 | 0.99673 | 1e-07
341.8 | 0.99537 | 1e-07
342 | 0.99253 | 1e-07
342.2 | 0.98907 | 1e-07
342.4 | 0.98916 | 1e-07
342.6 | 0.99245 | 1e-07
342.8 | 0.99515 | 1e-07
343 | 0.99649 | 1e-07
343.2 | 0.99708 | 1e-07
343.4 | 0.99728 | 1e-07
343.6 | 0.99727 | 1e-07
343.8 | 0.99722 | 1e-07
344 | 0.99745 | 1e-07
344.2 | 0.99805 | 1e-07
344.4 | 0.99863 | 1e-07
344.6 | 0.99899 | 1e-07
344.8 | 0.99918 | 1e-07
345 | 0.99929 | 1e-07
345.2 | 0.99935 | 1e-07
345.4 | 0.99938 | 1e-07
345.6 | 0.99941 | 1e-07
345.8 | 0.99942 | 1e-07
346 | 0.99942 | 1e-07
346.2 | 0.99942 | 1e-07
346.4 | 0.99941 | 1e-07
346.6 | 0.9994 | 1e-07
346.8 | 0.99937 | 1e-07
347 | 0.99932 | 1e-07
347.2 | 0.99929 | 1e-07
347.4 | 0.9993 | 1e-07
347.6 | 0.99935 | 1e-07
347.8 | 0.99939 | 1e-07
348 | 0.99941 | 1e-07
348.2 | 0.99943 | 1e-07
348.4 | 0.99943 | 1e-07
348.6 | 0.99943 | 1e-07
348.8 | 0.99941 | 1e-07
349 | 0.99937 | 1e-07
349.2 | 0.99928 | 1e-07
349.4 | 0.99908 | 1e-07
349.6 | 0.99879 | 1e-07
349.8 | 0.99881 | 1e-07
350 | 0.99904 | 1e-07
350.2 | 0.99917 | 1e-07
350.4 | 0.99921 | 1e-07
350.6 | 0.99918 | 1e-07
350.8 | 0.99905 | 1e-07
351 | 0.9988 | 1e-07
351.2 | 0.99876 | 1e-07
351.4 | 0.99908 | 1e-07
351.6 | 0.99929 | 1e-07
351.8 | 0.99938 | 1e-07
352 | 0.99942 | 1e-07
352.2 | 0.99943 | 1e-07
352.4 | 0.99943 | 1e-07
352.6 | 0.99941 | 1e-07
352.8 | 0.99936 | 1e-07
353 | 0.99923 | 1e-07
353.2 | 0.99883 | 1e-07
353.4 | 0.99846 | 1e-07
353.6 | 0.99896 | 1e-07
353.8 | 0.99933 | 1e-07
354 | 0.99944 | 1e-07
354.2 | 0.99948 | 1e-07
354.4 | 0.9995 | 1e-07
354.6 | 0.99951 | 1e-07
354.8 | 0.99951 | 1e-07
355 | 0.99952 | 1e-07
355.2 | 0.99951 | 1e-07
355.4 | 0.99951 | 1e-07
355.6 | 0.9995 | 1e-07
355.8 | 0.99947 | 1e-07
356 | 0.99941 | 1e-07
356.2 | 0.99926 | 1e-07
356.4 | 0.99927 | 1e-07
356.6 | 0.99943 | 1e-07
356.8 | 0.99949 | 1e-07
357 | 0.99951 | 1e-07
357.2 | 0.99952 | 1e-07
357.4 | 0.99953 | 1e-07
357.6 | 0.99953 | 1e-07
357.8 | 0.99953 | 1e-07
358 | 0.99953 | 1e-07
358.2 | 0.99953 | 1e-07
358.4 | 0.99952 | 1e-07
358.6 | 0.99951 | 1e-07
358.8 | 0.99947 | 1e-07
359 | 0.99934 | 1.0635e-07
359.2 | 0.99928 | 1.9725e-07
359.4 | 0.99941 | 1.3946e-07
359.6 | 0.99945 | 1.2832e-07
359.8 | 0.99944 | 1.8154e-07
360 | 0.99936 | 3.9543e-07
360.2 | 0.99924 | 8.001e-07
360.4 | 0.9994 | 4.2888e-07
360.6 | 0.99948 | 2.1018e-07
360.8 | 0.99951 | 1.431e-07
361 | 0.99951 | 1.2421e-07
361.2 | 0.99951 | 1.2878e-07
361.4 | 0.99951 | 1.5684e-07
361.6 | 0.99949 | 2.3535e-07
361.8 | 0.99942 | 4.9885e-07
362 | 0.99922 | 1.1407e-06
362.2 | 0.99938 | 4.824e-07
362.4 | 0.99949 | 1.3838e-07
362.6 | 0.99951 | 1e-07
362.8 | 0.99952 | 1e-07
363 | 0.99952 | 1e-07
363.2 | 0.99951 | 1e-07
363.4 | 0.9995 | 1e-07
363.6 | 0.99945 | 1e-07
363.8 | 0.99933 | 2.047e-07
364 | 0.99945 | 1e-07
364.2 | 0.99951 | 1e-07
364.4 | 0.99952 | 1e-07
364.6 | 0.99952 | 1e-07
364.8 | 0.99952 | 1e-07
365 | 0.99952 | 1e-07
365.2 | 0.99952 | 1e-07
365.4 | 0.99952 | 1e-07
365.6 | 0.99952 | 1e-07
365.8 | 0.99952 | 1e-07
366 | 0.99952 | 1e-07
366.2 | 0.99952 | 1e-07
366.4 | 0.99951 | 1e-07
366.6 | 0.99951 | 1e-07
366.8 | 0.9995 | 1.6236e-07
367 | 0.99943 | 1.1121e-06
367.2 | 0.99937 | 3.4441e-06
367.4 | 0.99947 | 1.4746e-06
367.6 | 0.99949 | 1.1739e-06
367.8 | 0.99949 | 1.5875e-06
368 | 0.99947 | 3.182e-06
368.2 | 0.99944 | 6.3777e-06
368.4 | 0.99947 | 2.4254e-06
368.6 | 0.99949 | 6.1957e-07
368.8 | 0.9995 | 2.153e-07
369 | 0.9995 | 1e-07
369.2 | 0.9995 | 1e-07
369.4 | 0.99949 | 1e-07
369.6 | 0.99949 | 1e-07
369.8 | 0.99949 | 1e-07
370 | 0.99949 | 1e-07
370.2 | 0.99948 | 1e-07
370.4 | 0.99946 | 1.3073e-07
370.6 | 0.99934 | 5.6905e-07
370.8 | 0.99936 | 5.3629e-07
371 | 0.99947 | 1e-07
371.2 | 0.99947 | 1e-07
371.4 | 0.99947 | 1e-07
371.6 | 0.99947 | 1e-07
371.8 | 0.99947 | 1e-07
372 | 0.99947 | 1e-07
372.2 | 0.99947 | 1e-07
372.4 | 0.99946 | 1e-07
372.6 | 0.99945 | 1e-07
372.8 | 0.99942 | 1.731e-07
373 | 0.99945 | 1e-07
373.2 | 0.99945 | 1e-07
373.4 | 0.99945 | 1e-07
373.6 | 0.99945 | 1e-07
373.8 | 0.99945 | 1e-07
374 | 0.99945 | 1e-07
374.2 | 0.99944 | 1e-07
374.4 | 0.99944 | 1e-07
374.6 | 0.99944 | 1e-07
374.8 | 0.99943 | 1e-07
375 | 0.99943 | 1e-07
375.2 | 0.99942 | 1e-07
375.4 | 0.99936 | 1.5273e-07
375.6 | 0.99937 | 1.5013e-07
375.8 | 0.9994 | 1e-07
376 | 0.99941 | 1e-07
376.2 | 0.99941 | 1e-07
376.4 | 0.99941 | 1e-07
376.6 | 0.99941 | 1e-07
376.8 | 0.9994 | 1e-07
377 | 0.9994 | 1e-07
377.2 | 0.9994 | 1e-07
377.4 | 0.99939 | 1e-07
377.6 | 0.99939 | 1e-07
377.8 | 0.99939 | 1e-07
378 | 0.99938 | 1e-07
378.2 | 0.99938 | 1.6709e-07
378.4 | 0.99937 | 9.1067e-07
378.6 | 0.99932 | 7.2729e-06
378.8 | 0.99936 | 1.1843e-06
379 | 0.99936 | 4.3229e-07
379.2 | 0.99936 | 2.5803e-07
379.4 | 0.99936 | 2.0147e-07
379.6 | 0.99935 | 2.0045e-07
379.8 | 0.99935 | 2.7377e-07
380 | 0.99934 | 6.9826e-07
380.2 | 0.9993 | 3.7462e-06
380.4 | 0.99933 | 8.1814e-07
380.6 | 0.99933 | 1.3454e-07
380.8 | 0.99933 | 1e-07
381 | 0.99932 | 1e-07
381.2 | 0.99932 | 1e-07
381.4 | 0.99931 | 1e-07
381.6 | 0.99931 | 1e-07
381.8 | 0.9993 | 1e-07
382 | 0.9993 | 1e-07
382.2 | 0.99929 | 1e-07
382.4 | 0.99929 | 1e-07
382.6 | 0.99928 | 1e-07
382.8 | 0.99927 | 1e-07
383 | 0.99927 | 2.7107e-07
383.2 | 0.99923 | 4.2213e-06
383.4 | 0.99925 | 1.7693e-06
383.6 | 0.99925 | 3.4107e-07
383.8 | 0.99925 | 2.067e-07
384 | 0.99924 | 2.095e-07
384.2 | 0.99923 | 3.5209e-07
384.4 | 0.99922 | 1.5687e-06
384.6 | 0.99922 | 1.3704e-06
384.8 | 0.99921 | 1.7126e-07
385 | 0.99921 | 1e-07
385.2 | 0.9992 | 1e-07
385.4 | 0.99919 | 1e-07
385.6 | 0.99919 | 1e-07
385.8 | 0.99918 | 1e-07
386 | 0.99917 | 1e-07
386.2 | 0.99917 | 1e-07
386.4 | 0.99916 | 1e-07
386.6 | 0.99915 | 1e-07
386.8 | 0.99914 | 1.3812e-07
387 | 0.99912 | 5.3175e-07
387.2 | 0.99913 | 1e-07
387.4 | 0.99912 | 1e-07
387.6 | 0.99911 | 1e-07
387.8 | 0.9991 | 1e-07
388 | 0.9991 | 1e-07
388.2 | 0.99909 | 1e-07
388.4 | 0.99908 | 1e-07
388.6 | 0.99907 | 1e-07
388.8 | 0.99906 | 1e-07
389 | 0.99905 | 1e-07
389.2 | 0.99904 | 1e-07
389.4 | 0.99903 | 1e-07
389.6 | 0.99902 | 1e-07
389.8 | 0.99901 | 1e-07
390 | 0.99899 | 1e-07
390.2 | 0.99899 | 1e-07
390.4 | 0.99899 | 1e-07
390.6 | 0.99898 | 1e-07
390.8 | 0.99897 | 1e-07
391 | 0.99896 | 1e-07
391.2 | 0.99894 | 1e-07
391.4 | 0.99893 | 1e-07
391.6 | 0.99892 | 1e-07
391.8 | 0.99891 | 1e-07
392 | 0.9989 | 1e-07
392.2 | 0.99889 | 1e-07
392.4 | 0.99888 | 1e-07
392.6 | 0.99887 | 1e-07
392.8 | 0.99884 | 1e-07
393 | 0.99881 | 2.97e-07
393.2 | 0.99883 | 1e-07
393.4 | 0.99882 | 1e-07
393.6 | 0.99881 | 1e-07
393.8 | 0.9988 | 1e-07
394 | 0.99878 | 1e-07
394.2 | 0.99877 | 1e-07
394.4 | 0.99876 | 1e-07
394.6 | 0.99875 | 1e-07
394.8 | 0.99873 | 1e-07
395 | 0.99872 | 1e-07
395.2 | 0.99871 | 1e-07
395.4 | 0.9987 | 1e-07
395.6 | 0.99868 | 1e-07
395.8 | 0.99866 | 5.568e-07
396 | 0.99866 | 1e-07
396.2 | 0.99864 | 1e-07
396.4 | 0.99863 | 1e-07
396.6 | 0.99862 | 1e-07
396.8 | 0.99861 | 1e-07
397 | 0.99859 | 1e-07
397.2 | 0.99858 | 1e-07
397.4 | 0.99857 | 1e-07
397.6 | 0.99855 | 1e-07
397.8 | 0.99854 | 1e-07
398 | 0.99853 | 1e-07
398.2 | 0.99852 | 1e-07
398.4 | 0.9985 | 6.6798e-07
398.6 | 0.99849 | 1.7219e-07
398.8 | 0.99848 | 1e-07
399 | 0.99847 | 1e-07
399.2 | 0.99845 | 1e-07
399.4 | 0.99844 | 1e-07
399.6 | 0.99843 | 1e-07
399.8 | 0.99842 | 1e-07
400 | 0.99841 | 1e-07
400.2 | 0.9984 | 1e-07
400.4 | 0.99839 | 1e-07
400.6 | 0.99838 | 1e-07
400.8 | 0.99837 | 4.0747e-07
401 | 0.99836 | 1.0421e-07
401.2 | 0.99835 | 1e-07
401.4 | 0.99834 | 1e-07
401.6 | 0.99833 | 1e-07
401.8 | 0.99832 | 1e-07
402 | 0.99832 | 1e-07
402.2 | 0.99831 | 1e-07
402.4 | 0.9983 | 1e-07
402.6 | 0.9983 | 1e-07
402.8 | 0.99829 | 1e-07
403 | 0.99829 | 1e-07
403.2 | 0.99828 | 1e-07
403.4 | 0.99828 | 1e-07
403.6 | 0.99827 | 1e-07
403.8 | 0.99827 | 1e-07
404 | 0.99827 | 1e-07
404.2 | 0.99827 | 1e-07
404.4 | 0.99826 | 1e-07
404.6 | 0.99826 | 1e-07
404.8 | 0.99826 | 1e-07
405 | 0.99826 | 1e-07
405.2 | 0.99826 | 1e-07
405.4 | 0.99826 | 1e-07
405.6 | 0.99827 | 1e-07
405.8 | 0.99827 | 1e-07
406 | 0.99827 | 1e-07
406.2 | 0.99827 | 1e-07
406.4 | 0.99827 | 1e-07
406.6 | 0.99828 | 1e-07
406.8 | 0.99828 | 1e-07
407 | 0.99829 | 1e-07
407.2 | 0.99829 | 1e-07
407.4 | 0.9983 | 1e-07
407.6 | 0.99831 | 1e-07
407.8 | 0.99831 | 1e-07
408 | 0.99832 | 1e-07
408.2 | 0.99833 | 1e-07
408.4 | 0.99833 | 1e-07
408.6 | 0.99834 | 1e-07
408.8 | 0.99835 | 1e-07
409 | 0.99836 | 1e-07
409.2 | 0.99836 | 1e-07
409.4 | 0.99837 | 1e-07
409.6 | 0.99838 | 1e-07
409.8 | 0.99839 | 1e-07
410 | 0.9984 | 1e-07
410.2 | 0.99841 | 1e-07
410.4 | 0.99841 | 1e-07
410.6 | 0.99842 | 1e-07
410.8 | 0.99843 | 1e-07
411 | 0.99844 | 1e-07
411.2 | 0.99845 | 1e-07
411.4 | 0.99846 | 1e-07
411.6 | 0.99847 | 1e-07
411.8 | 0.99847 | 1e-07
412 | 0.99848 | 1e-07
412.2 | 0.99849 | 1e-07
412.4 | 0.9985 | 1e-07
412.6 | 0.99851 | 1e-07
412.8 | 0.99851 | 1e-07
413 | 0.99852 | 1e-07
413.2 | 0.99853 | 1e-07
413.4 | 0.99853 | 1e-07
413.6 | 0.99854 | 2.7301e-07
413.8 | 0.99855 | 1e-07
414 | 0.99855 | 1e-07
414.2 | 0.99856 | 1e-07
414.4 | 0.99856 | 1e-07
414.6 | 0.99857 | 1e-07
414.8 | 0.99857 | 1e-07
415 | 0.99858 | 1e-07
415.2 | 0.99858 | 1e-07
415.4 | 0.99858 | 1e-07
415.6 | 0.99858 | 8.4973e-07
415.8 | 0.99859 | 1e-07
416 | 0.99859 | 1e-07
416.2 | 0.99859 | 1e-07
416.4 | 0.99859 | 1e-07
416.6 | 0.9986 | 1e-07
416.8 | 0.9986 | 1e-07
417 | 0.9986 | 1e-07
417.2 | 0.9986 | 1e-07
417.4 | 0.99859 | 1e-07
417.6 | 0.99859 | 1e-07
417.8 | 0.99859 | 1e-07
418 | 0.99859 | 1e-07
418.2 | 0.99859 | 1e-07
418.4 | 0.99858 | 1e-07
418.6 | 0.99854 | 1.9761e-06
418.8 | 0.99858 | 1e-07
419 | 0.99857 | 1e-07
419.2 | 0.99857 | 1e-07
419.4 | 0.99856 | 1e-07
419.6 | 0.99856 | 1e-07
419.8 | 0.99855 | 1e-07
420 | 0.99854 | 1e-07
420.2 | 0.99854 | 1e-07
420.4 | 0.99853 | 1e-07
420.6 | 0.99852 | 1e-07
420.8 | 0.99851 | 1e-07
421 | 0.9985 | 1e-07
421.2 | 0.99849 | 1e-07
421.4 | 0.99848 | 1e-07
421.6 | 0.99847 | 1e-07
421.8 | 0.99846 | 1e-07
422 | 0.99845 | 1.186e-07
422.2 | 0.99844 | 1e-07
422.4 | 0.99843 | 1e-07
422.6 | 0.99841 | 1e-07
422.8 | 0.9984 | 1e-07
423 | 0.99839 | 1e-07
423.2 | 0.99837 | 1e-07
423.4 | 0.99836 | 1e-07
423.6 | 0.99835 | 1e-07
423.8 | 0.99833 | 1e-07
424 | 0.99832 | 1e-07
424.2 | 0.9983 | 1e-07
424.4 | 0.99829 | 1e-07
424.6 | 0.99828 | 1e-07
424.8 | 0.99826 | 1e-07
425 | 0.99825 | 1e-07
425.2 | 0.99823 | 1e-07
425.4 | 0.99822 | 1e-07
425.6 | 0.99821 | 1e-07
425.8 | 0.99819 | 1e-07
426 | 0.99818 | 1e-07
426.2 | 0.99817 | 1e-07
426.4 | 0.99815 | 1e-07
426.6 | 0.99814 | 1e-07
426.8 | 0.99813 | 1e-07
427 | 0.99812 | 1e-07
427.2 | 0.99811 | 1e-07
427.4 | 0.9981 | 1e-07
427.6 | 0.99809 | 1e-07
427.8 | 0.99808 | 1e-07
428 | 0.99808 | 1e-07
428.2 | 0.99807 | 1e-07
428.4 | 0.99807 | 1e-07
428.6 | 0.99806 | 1e-07
428.8 | 0.99806 | 1e-07
429 | 0.99806 | 1e-07
429.2 | 0.99805 | 1e-07
429.4 | 0.99805 | 1e-07
429.6 | 0.99806 | 1e-07
429.8 | 0.99806 | 1e-07
430 | 0.99806 | 1e-07
430.2 | 0.99806 | 1e-07
430.4 | 0.99807 | 1e-07
430.6 | 0.99808 | 1e-07
430.8 | 0.99808 | 1e-07
431 | 0.99809 | 1e-07
431.2 | 0.9981 | 1e-07
431.4 | 0.99811 | 1e-07
431.6 | 0.99812 | 1e-07
431.8 | 0.99813 | 1e-07
432 | 0.99814 | 1e-07
432.2 | 0.99816 | 5.4085e-07
432.4 | 0.99817 | 1e-07
432.6 | 0.99818 | 1e-07
432.8 | 0.9982 | 1e-07
433 | 0.99821 | 1e-07
433.2 | 0.99823 | 1e-07
433.4 | 0.99824 | 1e-07
433.6 | 0.99826 | 1e-07
433.8 | 0.99828 | 1e-07
434 | 0.99829 | 1e-07
434.2 | 0.99831 | 9.6062e-07
434.4 | 0.99832 | 3.5661e-07
434.6 | 0.99834 | 1e-07
434.8 | 0.99835 | 1e-07
435 | 0.99837 | 1e-07
435.2 | 0.99839 | 1e-07
435.4 | 0.9984 | 1e-07
435.6 | 0.99842 | 1e-07
435.8 | 0.99843 | 1e-07
436 | 0.99844 | 1e-07
436.2 | 0.99846 | 1e-07
436.4 | 0.99847 | 1e-07
436.6 | 0.99848 | 1e-07
436.8 | 0.9985 | 1e-07
437 | 0.99851 | 1e-07
437.2 | 0.99852 | 2.1585e-07
437.4 | 0.99853 | 6.5932e-07
437.6 | 0.99854 | 3.6836e-07
437.8 | 0.99855 | 1e-07
438 | 0.99856 | 1e-07
438.2 | 0.99857 | 1e-07
438.4 | 0.99858 | 1e-07
438.6 | 0.99858 | 1e-07
438.8 | 0.99859 | 1e-07
439 | 0.9986 | 1e-07
439.2 | 0.9986 | 1e-07
439.4 | 0.9986 | 1e-07
439.6 | 0.99861 | 1e-07
439.8 | 0.99861 | 2.7951e-07
440 | 0.99861 | 9.4945e-07
440.2 | 0.99861 | 1.0134e-07
440.4 | 0.99862 | 1e-07
440.6 | 0.99862 | 1e-07
440.8 | 0.99861 | 1e-07
441 | 0.99861 | 1e-07
441.2 | 0.99861 | 1e-07
441.4 | 0.99861 | 1e-07
441.6 | 0.9986 | 1e-07
441.8 | 0.9986 | 1e-07
442 | 0.99859 | 1e-07
442.2 | 0.99859 | 1e-07
442.4 | 0.99858 | 1e-07
442.6 | 0.99857 | 1e-07
442.8 | 0.99856 | 1e-07
443 | 0.99855 | 1e-07
443.2 | 0.99854 | 1e-07
443.4 | 0.99853 | 4.2234e-07
443.6 | 0.99852 | 6.3845e-07
443.8 | 0.9985 | 1e-07
444 | 0.99849 | 1e-07
444.2 | 0.99848 | 1e-07
444.4 | 0.99846 | 1e-07
444.6 | 0.99844 | 1e-07
444.8 | 0.99843 | 1e-07
445 | 0.99841 | 1e-07
445.2 | 0.99839 | 1e-07
445.4 | 0.99837 | 1e-07
445.6 | 0.99835 | 1e-07
445.8 | 0.99833 | 1e-07
446 | 0.99831 | 1e-07
446.2 | 0.99829 | 1e-07
446.4 | 0.99827 | 1e-07
446.6 | 0.99825 | 1e-07
446.8 | 0.99823 | 1e-07
447 | 0.99821 | 1e-07
447.2 | 0.99819 | 1e-07
447.4 | 0.99816 | 1.5656e-07
447.6 | 0.99814 | 4.0199e-07
447.8 | 0.99812 | 5.377e-07
448 | 0.9981 | 3.2479e-07
448.2 | 0.99809 | 1.7834e-07
448.4 | 0.99807 | 1.1565e-07
448.6 | 0.99805 | 1e-07
448.8 | 0.99804 | 1e-07
449 | 0.99802 | 1e-07
449.2 | 0.99801 | 1e-07
449.4 | 0.998 | 1e-07
449.6 | 0.99799 | 1e-07
449.8 | 0.99798 | 1e-07
450 | 0.99797 | 1e-07
450.2 | 0.99797 | 1e-07
450.4 | 0.99796 | 1e-07
450.6 | 0.99796 | 1e-07
450.8 | 0.99796 | 1e-07
451 | 0.99797 | 1e-07
451.2 | 0.99797 | 1e-07
451.4 | 0.99798 | 1e-07
451.6 | 0.99799 | 1.0435e-07
451.8 | 0.998 | 1.5224e-07
452 | 0.99801 | 2.5847e-07
452.2 | 0.99802 | 4.8424e-07
452.4 | 0.99804 | 6.8527e-07
452.6 | 0.99806 | 4.9887e-07
452.8 | 0.99807 | 2.2106e-07
453 | 0.99809 | 1.0905e-07
453.2 | 0.99811 | 1e-07
453.4 | 0.99813 | 1e-07
453.6 | 0.99815 | 1e-07
453.8 | 0.99817 | 1e-07
454 | 0.9982 | 1e-07
454.2 | 0.99822 | 1e-07
454.4 | 0.99824 | 1e-07
454.6 | 0.99826 | 1e-07
454.8 | 0.99828 | 1e-07
455 | 0.9983 | 1e-07
455.2 | 0.99833 | 1e-07
455.4 | 0.99835 | 1e-07
455.6 | 0.99837 | 1e-07
455.8 | 0.99839 | 1e-07
456 | 0.9984 | 1e-07
456.2 | 0.99842 | 1.0214e-07
456.4 | 0.99844 | 1.6812e-07
456.6 | 0.99846 | 1.9606e-07
456.8 | 0.99847 | 1.3817e-07
457 | 0.99849 | 1e-07
457.2 | 0.9985 | 1e-07
457.4 | 0.99852 | 1e-07
457.6 | 0.99853 | 1e-07
457.8 | 0.99854 | 1e-07
458 | 0.99855 | 1e-07
458.2 | 0.99856 | 1e-07
458.4 | 0.99857 | 1e-07
458.6 | 0.99857 | 1e-07
458.8 | 0.99858 | 1e-07
459 | 0.99859 | 1e-07
459.2 | 0.99859 | 1e-07
459.4 | 0.99859 | 1e-07
459.6 | 0.99859 | 1.1864e-07
459.8 | 0.9986 | 2.8247e-07
460 | 0.99859 | 5.8616e-07
460.2 | 0.99859 | 3.8162e-07
460.4 | 0.99859 | 1.4513e-07
460.6 | 0.99859 | 1e-07
460.8 | 0.99858 | 1e-07
461 | 0.99858 | 1e-07
461.2 | 0.99857 | 1e-07
461.4 | 0.99856 | 1e-07
461.6 | 0.99855 | 1e-07
461.8 | 0.99854 | 1e-07
462 | 0.99853 | 1e-07
462.2 | 0.99852 | 1e-07
462.4 | 0.99851 | 1e-07
462.6 | 0.99849 | 1e-07
462.8 | 0.99848 | 1e-07
463 | 0.99846 | 1e-07
463.2 | 0.99844 | 1e-07
463.4 | 0.99843 | 1e-07
463.6 | 0.99841 | 1e-07
463.8 | 0.99839 | 1.1467e-07
464 | 0.99836 | 2.7704e-07
464.2 | 0.99834 | 8.5629e-07
464.4 | 0.99832 | 9.8124e-07
464.6 | 0.9983 | 3.5084e-07
464.8 | 0.99828 | 1.5333e-07
465 | 0.99825 | 1e-07
465.2 | 0.99823 | 1e-07
465.4 | 0.99821 | 1e-07
465.6 | 0.99819 | 1e-07
465.8 | 0.99816 | 1e-07
466 | 0.99814 | 1e-07
466.2 | 0.99812 | 1e-07
466.4 | 0.9981 | 1e-07
466.6 | 0.99808 | 1e-07
466.8 | 0.99806 | 1e-07
467 | 0.99805 | 1e-07
467.2 | 0.99803 | 1e-07
467.4 | 0.99802 | 1e-07
467.6 | 0.99801 | 1.3893e-07
467.8 | 0.998 | 2.6306e-07
468 | 0.99799 | 6.8464e-07
468.2 | 0.99799 | 2.0499e-06
468.4 | 0.99799 | 2.4756e-06
468.6 | 0.99799 | 8.8788e-07
468.8 | 0.99799 | 3.4864e-07
469 | 0.998 | 1.8878e-07
469.2 | 0.998 | 1.2345e-07
469.4 | 0.99801 | 1e-07
469.6 | 0.99802 | 1e-07
469.8 | 0.99804 | 1e-07
470 | 0.99805 | 1e-07
470.2 | 0.99807 | 1e-07
470.4 | 0.99809 | 1e-07
470.6 | 0.99811 | 1e-07
470.8 | 0.99813 | 1e-07
471 | 0.99815 | 1e-07
471.2 | 0.99817 | 1e-07
471.4 | 0.99819 | 1.2066e-07
471.6 | 0.99821 | 1.7079e-07
471.8 | 0.99824 | 2.6808e-07
472 | 0.99826 | 4.8265e-07
472.2 | 0.99828 | 9.9398e-07
472.4 | 0.9983 | 1.6857e-06
472.6 | 0.99832 | 1.7044e-06
472.8 | 0.99834 | 1.0682e-06
473 | 0.99837 | 6.0451e-07
473.2 | 0.99839 | 3.8725e-07
473.4 | 0.9984 | 2.8181e-07
473.6 | 0.99842 | 2.2731e-07
473.8 | 0.99844 | 1.9947e-07
474 | 0.99846 | 1.8827e-07
474.2 | 0.99847 | 1.8995e-07
474.4 | 0.99848 | 2.0435e-07
474.6 | 0.9985 | 2.3434e-07
474.8 | 0.99851 | 2.8649e-07
475 | 0.99852 | 3.7146e-07
475.2 | 0.99853 | 4.9834e-07
475.4 | 0.99854 | 6.4351e-07
475.6 | 0.99854 | 7.0675e-07
475.8 | 0.99855 | 6.114e-07
476 | 0.99855 | 4.4656e-07
476.2 | 0.99856 | 3.1239e-07
476.4 | 0.99856 | 2.2513e-07
476.6 | 0.99856 | 1.7112e-07
476.8 | 0.99856 | 1.3762e-07
477 | 0.99856 | 1.1679e-07
477.2 | 0.99855 | 1.0421e-07
477.4 | 0.99855 | 1e-07
477.6 | 0.99854 | 1e-07
477.8 | 0.99854 | 1e-07
478 | 0.99853 | 1.0529e-07
478.2 | 0.99852 | 1.1888e-07
478.4 | 0.99851 | 1.4155e-07
478.6 | 0.9985 | 1.7851e-07
478.8 | 0.99849 | 2.3892e-07
479 | 0.99847 | 3.3583e-07
479.2 | 0.99846 | 4.7089e-07
479.4 | 0.99844 | 5.7877e-07
479.6 | 0.99843 | 5.4164e-07
479.8 | 0.99841 | 3.9874e-07
480 | 0.99839 | 2.7044e-07
480.2 | 0.99837 | 1.8726e-07
480.4 | 0.99836 | 1.3674e-07
480.6 | 0.99834 | 1.0569e-07
480.8 | 0.99832 | 1e-07
481 | 0.9983 | 1e-07
481.2 | 0.99828 | 1e-07
481.4 | 0.99826 | 1e-07
481.6 | 0.99824 | 1e-07
481.8 | 0.99822 | 1e-07
482 | 0.9982 | 1e-07
482.2 | 0.99819 | 1e-07
482.4 | 0.99817 | 1e-07
482.6 | 0.99816 | 1e-07
482.8 | 0.99814 | 1.1611e-07
483 | 0.99813 | 1.6082e-07
483.2 | 0.99812 | 2.495e-07
483.4 | 0.99811 | 4.5739e-07
483.6 | 0.99811 | 1.0282e-06
483.8 | 0.9981 | 1.8618e-06
484 | 0.9981 | 1.5441e-06
484.2 | 0.9981 | 6.2227e-07
484.4 | 0.9981 | 2.8665e-07
484.6 | 0.99811 | 1.6094e-07
484.8 | 0.99811 | 1.034e-07
485 | 0.99812 | 1e-07
485.2 | 0.99813 | 1e-07
485.4 | 0.99814 | 1e-07
485.6 | 0.99816 | 1e-07
485.8 | 0.99817 | 1e-07
486 | 0.99819 | 1e-07
486.2 | 0.99821 | 1e-07
486.4 | 0.99822 | 1e-07
486.6 | 0.99824 | 1e-07
486.8 | 0.99826 | 1e-07
487 | 0.99828 | 1e-07
487.2 | 0.9983 | 1e-07
487.4 | 0.99832 | 1e-07
487.6 | 0.99834 | 1e-07
487.8 | 0.99836 | 1e-07
488 | 0.99838 | 1e-07
488.2 | 0.9984 | 1e-07
488.4 | 0.99841 | 1e-07
488.6 | 0.99843 | 1.2691e-07
488.8 | 0.99845 | 2.1357e-07
489 | 0.99846 | 4.1588e-07
489.2 | 0.99848 | 8.502e-07
489.4 | 0.99849 | 1.0402e-06
489.6 | 0.9985 | 6.0517e-07
489.8 | 0.99851 | 3.0646e-07
490 | 0.99853 | 1.7739e-07
490.2 | 0.99853 | 1.1684e-07
490.4 | 0.99854 | 1e-07
490.6 | 0.99855 | 1e-07
490.8 | 0.99856 | 1e-07
491 | 0.99856 | 1e-07
491.2 | 0.99856 | 1e-07
491.4 | 0.99856 | 1e-07
491.6 | 0.99857 | 1e-07
491.8 | 0.99857 | 1e-07
492 | 0.99856 | 1e-07
492.2 | 0.99856 | 1e-07
492.4 | 0.99856 | 1e-07
492.6 | 0.99855 | 1e-07
492.8 | 0.99855 | 1e-07
493 | 0.99854 | 1.0942e-07
493.2 | 0.99853 | 1.5733e-07
493.4 | 0.99852 | 2.474e-07
493.6 | 0.99851 | 4.2292e-07
493.8 | 0.9985 | 6.9169e-07
494 | 0.99849 | 7.5598e-07
494.2 | 0.99848 | 4.9994e-07
494.4 | 0.99846 | 2.8736e-07
494.6 | 0.99845 | 1.7588e-07
494.8 | 0.99844 | 1.1779e-07
495 | 0.99842 | 1e-07
495.2 | 0.99841 | 1e-07
495.4 | 0.99839 | 1e-07
495.6 | 0.99838 | 1e-07
495.8 | 0.99836 | 1e-07
496 | 0.99835 | 1e-07
496.2 | 0.99834 | 1e-07
496.4 | 0.99832 | 1e-07
496.6 | 0.99831 | 1e-07
496.8 | 0.9983 | 1e-07
497 | 0.99829 | 1e-07
497.2 | 0.99828 | 1e-07
497.4 | 0.99828 | 1e-07
497.6 | 0.99827 | 1e-07
497.8 | 0.99827 | 1e-07
498 | 0.99826 | 1e-07
498.2 | 0.99826 | 1e-07
498.4 | 0.99826 | 1e-07
498.6 | 0.99827 | 1.2091e-07
498.8 | 0.99827 | 1.8349e-07
499 | 0.99828 | 3.2076e-07
499.2 | 0.99828 | 7.1935e-07
499.4 | 0.99829 | 2.3586e-06
499.6 | 0.9983 | 3.7763e-06
499.8 | 0.99831 | 2.1292e-06
500 | 0.99833 | 6.6684e-07
500.2 | 0.99834 | 3.061e-07
500.4 | 0.99836 | 1.7624e-07
500.6 | 0.99837 | 1.1601e-07
500.8 | 0.99839 | 1e-07
501 | 0.9984 | 1e-07
501.2 | 0.99842 | 1e-07
501.4 | 0.99844 | 1e-07
501.6 | 0.99845 | 1e-07
501.8 | 0.99847 | 1e-07
502 | 0.99849 | 1e-07
502.2 | 0.9985 | 1e-07
502.4 | 0.99852 | 1e-07
502.6 | 0.99853 | 1e-07
502.8 | 0.99854 | 1e-07
503 | 0.99856 | 1e-07
503.2 | 0.99857 | 1e-07
503.4 | 0.99858 | 1e-07
503.6 | 0.99859 | 1e-07
503.8 | 0.9986 | 1e-07
504 | 0.99861 | 1e-07
504.2 | 0.99861 | 1e-07
504.4 | 0.99862 | 1e-07
504.6 | 0.99862 | 1e-07
504.8 | 0.99863 | 1e-07
505 | 0.99863 | 1e-07
505.2 | 0.99863 | 1e-07
505.4 | 0.99863 | 1e-07
505.6 | 0.99863 | 1.1583e-07
505.8 | 0.99863 | 1.491e-07
506 | 0.99863 | 1.9627e-07
506.2 | 0.99862 | 2.6375e-07
506.4 | 0.99862 | 3.5987e-07
506.6 | 0.99861 | 4.9257e-07
506.8 | 0.9986 | 6.6124e-07
507 | 0.99859 | 8.4205e-07
507.2 | 0.99858 | 9.8423e-07
507.4 | 0.99857 | 1.044e-06
507.6 | 0.99856 | 1.0241e-06
507.8 | 0.99855 | 9.6128e-07
508 | 0.99853 | 8.9046e-07
508.2 | 0.99852 | 8.3061e-07
508.4 | 0.9985 | 7.8879e-07
508.6 | 0.99849 | 7.6646e-07
508.8 | 0.99847 | 7.6346e-07
509 | 0.99846 | 7.7964e-07
509.2 | 0.99844 | 8.1559e-07
509.4 | 0.99842 | 8.7274e-07
509.6 | 0.99841 | 9.5319e-07
509.8 | 0.99839 | 1.0591e-06
510 | 0.99838 | 1.1911e-06
510.2 | 0.99837 | 1.3459e-06
510.4 | 0.99835 | 1.5119e-06
510.6 | 0.99834 | 1.6658e-06
510.8 | 0.99833 | 1.7743e-06
511 | 0.99832 | 1.8062e-06
511.2 | 0.99831 | 1.7499e-06
511.4 | 0.99831 | 1.6211e-06
511.6 | 0.9983 | 1.4526e-06
511.8 | 0.9983 | 1.2762e-06
512 | 0.9983 | 1.1124e-06
512.2 | 0.9983 | 9.7067e-07
512.4 | 0.99831 | 8.5264e-07
512.6 | 0.99831 | 7.5647e-07
512.8 | 0.99832 | 6.7901e-07
513 | 0.99833 | 6.1703e-07
513.2 | 0.99834 | 5.677e-07
513.4 | 0.99835 | 5.2869e-07
513.6 | 0.99837 | 4.9814e-07
513.8 | 0.99838 | 4.7464e-07
514 | 0.9984 | 4.5708e-07
514.2 | 0.99842 | 4.4465e-07
514.4 | 0.99843 | 4.3673e-07
514.6 | 0.99845 | 4.3291e-07
514.8 | 0.99847 | 4.329e-07
515 | 0.99849 | 4.3654e-07
515.2 | 0.9985 | 4.438e-07
515.4 | 0.99852 | 4.5476e-07
515.6 | 0.99854 | 4.6959e-07
515.8 | 0.99856 | 4.886e-07
516 | 0.99857 | 5.1226e-07
516.2 | 0.99859 | 5.4117e-07
516.4 | 0.9986 | 5.7613e-07
516.6 | 0.99862 | 6.182e-07
516.8 | 0.99863 | 6.6868e-07
517 | 0.99864 | 7.2921e-07
517.2 | 0.99866 | 8.0178e-07
517.4 | 0.99867 | 8.8874e-07
517.6 | 0.99867 | 9.9265e-07
517.8 | 0.99868 | 1.116e-06
518 | 0.99869 | 1.2604e-06
518.2 | 0.9987 | 1.4254e-06
518.4 | 0.9987 | 1.6062e-06
518.6 | 0.9987 | 1.7909e-06
518.8 | 0.99871 | 1.9591e-06
519 | 0.99871 | 2.0833e-06
519.2 | 0.99871 | 2.1373e-06
519.4 | 0.99871 | 2.1074e-06
519.6 | 0.99871 | 2.0002e-06
519.8 | 0.9987 | 1.8393e-06
520 | 0.9987 | 1.6539e-06
520.2 | 0.99869 | 1.4682e-06
520.4 | 0.99869 | 1.2975e-06
520.6 | 0.99868 | 1.1486e-06
520.8 | 0.99867 | 1.0231e-06
521 | 0.99866 | 9.1936e-07
521.2 | 0.99865 | 8.3499e-07
521.4 | 0.99864 | 7.6722e-07
521.6 | 0.99863 | 7.1353e-07
521.8 | 0.99862 | 6.7182e-07
522 | 0.99861 | 6.4042e-07
522.2 | 0.9986 | 6.1805e-07
522.4 | 0.99858 | 6.0383e-07
522.6 | 0.99857 | 5.9716e-07
522.8 | 0.99856 | 5.9779e-07
523 | 0.99854 | 6.0573e-07
523.2 | 0.99853 | 6.2131e-07
523.4 | 0.99852 | 6.4522e-07
523.6 | 0.9985 | 6.785e-07
523.8 | 0.99849 | 7.2273e-07
524 | 0.99848 | 7.8011e-07
524.2 | 0.99847 | 8.537e-07
524.4 | 0.99846 | 9.4772e-07
524.6 | 0.99845 | 1.068e-06
524.8 | 0.99844 | 1.2228e-06
525 | 0.99844 | 1.4232e-06
525.2 | 0.99843 | 1.6849e-06
525.4 | 0.99843 | 2.0286e-06
525.6 | 0.99842 | 2.4808e-06
525.8 | 0.99842 | 3.069e-06
526 | 0.99842 | 3.8079e-06
526.2 | 0.99842 | 4.6645e-06
526.4 | 0.99843 | 5.5075e-06
526.6 | 0.99843 | 6.0988e-06
526.8 | 0.99844 | 6.2129e-06
527 | 0.99844 | 5.8195e-06
527.2 | 0.99845 | 5.1055e-06
527.4 | 0.99846 | 4.3098e-06
527.6 | 0.99847 | 3.5855e-06
527.8 | 0.99849 | 2.9878e-06
528 | 0.9985 | 2.5163e-06
528.2 | 0.99851 | 2.1511e-06
528.4 | 0.99852 | 1.8693e-06
528.6 | 0.99853 | 1.6516e-06
528.8 | 0.99855 | 1.4827e-06
529 | 0.99856 | 1.3514e-06
529.2 | 0.99857 | 1.2493e-06
529.4 | 0.99858 | 1.1702e-06
529.6 | 0.99859 | 1.1097e-06
529.8 | 0.9986 | 1.0644e-06
530 | 0.99861 | 1.0317e-06
530.2 | 0.99862 | 1.0097e-06
530.4 | 0.99863 | 9.9684e-07
530.6 | 0.99864 | 9.9203e-07
530.8 | 0.99864 | 9.9426e-07
531 | 0.99865 | 1.0027e-06
531.2 | 0.99865 | 1.0165e-06
531.4 | 0.99866 | 1.035e-06
531.6 | 0.99866 | 1.0573e-06
531.8 | 0.99866 | 1.0826e-06
532 | 0.99866 | 1.1101e-06
532.2 | 0.99866 | 1.1388e-06
532.4 | 0.99866 | 1.168e-06
532.6 | 0.99866 | 1.1969e-06
532.8 | 0.99866 | 1.2247e-06
533 | 0.99865 | 1.2511e-06
533.2 | 0.99865 | 1.2759e-06
533.4 | 0.99864 | 1.2991e-06
533.6 | 0.99864 | 1.3212e-06
533.8 | 0.99863 | 1.3427e-06
534 | 0.99863 | 1.3644e-06
534.2 | 0.99862 | 1.3875e-06
534.4 | 0.99861 | 1.413e-06
534.6 | 0.9986 | 1.4421e-06
534.8 | 0.99859 | 1.4762e-06
535 | 0.99859 | 1.5166e-06
535.2 | 0.99858 | 1.5647e-06
535.4 | 0.99857 | 1.6223e-06
535.6 | 0.99856 | 1.6911e-06
535.8 | 0.99855 | 1.7731e-06
536 | 0.99855 | 1.8708e-06
536.2 | 0.99854 | 1.9868e-06
536.4 | 0.99853 | 2.1245e-06
536.6 | 0.99853 | 2.2878e-06
536.8 | 0.99852 | 2.4813e-06
537 | 0.99852 | 2.7103e-06
537.2 | 0.99851 | 2.9813e-06
537.4 | 0.99851 | 3.3013e-06
537.6 | 0.99851 | 3.6779e-06
537.8 | 0.99851 | 4.1187e-06
538 | 0.99851 | 4.6299e-06
538.2 | 0.99851 | 5.2142e-06
538.4 | 0.99852 | 5.868e-06
538.6 | 0.99852 | 6.5777e-06
538.8 | 0.99853 | 7.3156e-06
539 | 0.99853 | 8.039e-06
539.2 | 0.99854 | 8.6939e-06
539.4 | 0.99855 | 9.2248e-06
539.6 | 0.99855 | 9.5888e-06
539.8 | 0.99856 | 9.7683e-06
540 | 0.99857 | 9.7749e-06
540.2 | 0.99858 | 9.6433e-06
540.4 | 0.99859 | 9.4203e-06
540.6 | 0.99861 | 9.1526e-06
540.8 | 0.99862 | 8.8803e-06
541 | 0.99863 | 8.6337e-06
541.2 | 0.99864 | 8.4341e-06
541.4 | 0.99865 | 8.296e-06
541.6 | 0.99866 | 8.2293e-06
541.8 | 0.99867 | 8.2411e-06
542 | 0.99868 | 8.3379e-06
542.2 | 0.99869 | 8.526e-06
542.4 | 0.9987 | 8.8122e-06
542.6 | 0.9987 | 9.2038e-06
542.8 | 0.99871 | 9.7075e-06
543 | 0.99872 | 1.0328e-05
543.2 | 0.99872 | 1.1065e-05
543.4 | 0.99873 | 1.1907e-05
543.6 | 0.99873 | 1.2826e-05
543.8 | 0.99873 | 1.3769e-05
544 | 0.99874 | 1.4654e-05
544.2 | 0.99874 | 1.5376e-05
544.4 | 0.99874 | 1.5826e-05
544.6 | 0.99874 | 1.592e-05
544.8 | 0.99874 | 1.5633e-05
545 | 0.99874 | 1.5006e-05
545.2 | 0.99874 | 1.4132e-05
545.4 | 0.99875 | 1.3126e-05
545.6 | 0.99875 | 1.2089e-05
545.8 | 0.99874 | 1.1097e-05
546 | 0.99874 | 1.0196e-05
546.2 | 0.99874 | 9.4077e-06
546.4 | 0.99874 | 8.738e-06
546.6 | 0.99874 | 8.1834e-06
546.8 | 0.99874 | 7.7361e-06
547 | 0.99873 | 7.387e-06
547.2 | 0.99873 | 7.1273e-06
547.4 | 0.99873 | 6.9495e-06
547.6 | 0.99872 | 6.8474e-06
547.8 | 0.99872 | 6.8162e-06
548 | 0.99872 | 6.8525e-06
548.2 | 0.99871 | 6.9531e-06
548.4 | 0.99871 | 7.1152e-06
548.6 | 0.99871 | 7.3345e-06
548.8 | 0.99871 | 7.6042e-06
549 | 0.99871 | 7.9129e-06
549.2 | 0.99871 | 8.2421e-06
549.4 | 0.9987 | 8.5641e-06
549.6 | 0.9987 | 8.8413e-06
549.8 | 0.9987 | 9.0286e-06
550 | 0.99871 | 9.08e-06
550.2 | 0.99871 | 8.9604e-06
550.4 | 0.99871 | 8.6571e-06
550.6 | 0.99871 | 8.186e-06
550.8 | 0.99872 | 7.5883e-06
551 | 0.99872 | 6.9182e-06
551.2 | 0.99872 | 6.2287e-06
551.4 | 0.99873 | 5.562e-06
551.6 | 0.99873 | 4.9458e-06
551.8 | 0.99874 | 4.394e-06
552 | 0.99874 | 3.911e-06
552.2 | 0.99875 | 3.4944e-06
552.4 | 0.99875 | 3.1389e-06
552.6 | 0.99876 | 2.8375e-06
552.8 | 0.99877 | 2.5831e-06
553 | 0.99877 | 2.3692e-06
553.2 | 0.99878 | 2.1901e-06
553.4 | 0.99878 | 2.0406e-06
553.6 | 0.99878 | 1.9167e-06
553.8 | 0.99879 | 1.8148e-06
554 | 0.99879 | 1.7323e-06
554.2 | 0.9988 | 1.667e-06
554.4 | 0.9988 | 1.617e-06
554.6 | 0.9988 | 1.5812e-06
554.8 | 0.9988 | 1.5587e-06
555 | 0.9988 | 1.5489e-06
555.2 | 0.99881 | 1.5516e-06
555.4 | 0.99881 | 1.567e-06
555.6 | 0.99881 | 1.5955e-06
555.8 | 0.99881 | 1.6381e-06
556 | 0.9988 | 1.696e-06
556.2 | 0.9988 | 1.7711e-06
556.4 | 0.9988 | 1.8658e-06
556.6 | 0.9988 | 1.9835e-06
556.8 | 0.99879 | 2.1284e-06
557 | 0.99879 | 2.3063e-06
557.2 | 0.99879 | 2.5247e-06
557.4 | 0.99878 | 2.7938e-06
557.6 | 0.99878 | 3.1271e-06
557.8 | 0.99877 | 3.5434e-06
558 | 0.99877 | 4.0689e-06
558.2 | 0.99876 | 4.7406e-06
558.4 | 0.99876 | 5.6123e-06
558.6 | 0.99875 | 6.7642e-06
558.8 | 0.99875 | 8.3186e-06
559 | 0.99874 | 1.0468e-05
559.2 | 0.99873 | 1.3525e-05
559.4 | 0.99872 | 1.8008e-05
559.6 | 0.99871 | 2.4796e-05
559.8 | 0.9987 | 3.5331e-05
560 | 0.99868 | 5.1593e-05
560.2 | 0.99866 | 7.4438e-05
560.4 | 0.99863 | 9.8018e-05
560.6 | 0.99862 | 0.00010825
560.8 | 0.99863 | 9.8397e-05
561 | 0.99866 | 7.7747e-05
561.2 | 0.99868 | 5.8149e-05
561.4 | 0.9987 | 4.3665e-05
561.6 | 0.99871 | 3.3716e-05
561.8 | 0.99872 | 2.6917e-05
562 | 0.99873 | 2.2187e-05
562.2 | 0.99874 | 1.8822e-05
562.4 | 0.99875 | 1.6379e-05
562.6 | 0.99876 | 1.4577e-05
562.8 | 0.99877 | 1.3237e-05
563 | 0.99878 | 1.224e-05
563.2 | 0.99879 | 1.1507e-05
563.4 | 0.99881 | 1.0986e-05
563.6 | 0.99882 | 1.064e-05
563.8 | 0.99883 | 1.0448e-05
564 | 0.99884 | 1.0395e-05
564.2 | 0.99885 | 1.0476e-05
564.4 | 0.99886 | 1.0691e-05
564.6 | 0.99887 | 1.1047e-05
564.8 | 0.99888 | 1.1557e-05
565 | 0.99889 | 1.2241e-05
565.2 | 0.9989 | 1.3129e-05
565.4 | 0.99891 | 1.4259e-05
565.6 | 0.99892 | 1.5683e-05
565.8 | 0.99893 | 1.7466e-05
566 | 0.99893 | 1.9691e-05
566.2 | 0.99894 | 2.2456e-05
566.4 | 0.99894 | 2.5866e-05
566.6 | 0.99894 | 3.0016e-05
566.8 | 0.99895 | 3.4944e-05
567 | 0.99895 | 4.055e-05
567.2 | 0.99895 | 4.649e-05
567.4 | 0.99894 | 5.2095e-05
567.6 | 0.99894 | 5.6447e-05
567.8 | 0.99895 | 5.8687e-05
568 | 0.99895 | 5.843e-05
568.2 | 0.99895 | 5.5954e-05
568.4 | 0.99896 | 5.2009e-05
568.6 | 0.99896 | 4.7431e-05
568.8 | 0.99897 | 4.2866e-05
569 | 0.99897 | 3.8694e-05
569.2 | 0.99897 | 3.5081e-05
569.4 | 0.99898 | 3.206e-05
569.6 | 0.99898 | 2.9596e-05
569.8 | 0.99897 | 2.7631e-05
570 | 0.99897 | 2.6102e-05
570.2 | 0.99897 | 2.4954e-05
570.4 | 0.99897 | 2.4141e-05
570.6 | 0.99896 | 2.3627e-05
570.8 | 0.99896 | 2.3388e-05
571 | 0.99895 | 2.3407e-05
571.2 | 0.99895 | 2.3676e-05
571.4 | 0.99894 | 2.4194e-05
571.6 | 0.99893 | 2.4965e-05
571.8 | 0.99893 | 2.6002e-05
572 | 0.99892 | 2.7317e-05
572.2 | 0.99891 | 2.8928e-05
572.4 | 0.9989 | 3.0852e-05
572.6 | 0.9989 | 3.31e-05
572.8 | 0.99889 | 3.5672e-05
573 | 0.99888 | 3.8547e-05
573.2 | 0.99887 | 4.1673e-05
573.4 | 0.99886 | 4.495e-05
573.6 | 0.99886 | 4.8224e-05
573.8 | 0.99885 | 5.1287e-05
574 | 0.99885 | 5.3892e-05
574.2 | 0.99884 | 5.5795e-05
574.4 | 0.99884 | 5.6806e-05
574.6 | 0.99884 | 5.6835e-05
574.8 | 0.99884 | 5.5919e-05
575 | 0.99885 | 5.4206e-05
575.2 | 0.99885 | 5.1915e-05
575.4 | 0.99886 | 4.9282e-05
575.6 | 0.99887 | 4.652e-05
575.8 | 0.99887 | 4.38e-05
576 | 0.99888 | 4.124e-05
576.2 | 0.99889 | 3.8915e-05
576.4 | 0.9989 | 3.6865e-05
576.6 | 0.99891 | 3.5108e-05
576.8 | 0.99892 | 3.3646e-05
577 | 0.99893 | 3.2475e-05
577.2 | 0.99894 | 3.1589e-05
577.4 | 0.99895 | 3.0982e-05
577.6 | 0.99896 | 3.065e-05
577.8 | 0.99897 | 3.0595e-05
578 | 0.99898 | 3.0824e-05
578.2 | 0.99899 | 3.135e-05
578.4 | 0.99899 | 3.2195e-05
578.6 | 0.999 | 3.3386e-05
578.8 | 0.99901 | 3.4961e-05
579 | 0.99902 | 3.6966e-05
579.2 | 0.99902 | 3.9452e-05
579.4 | 0.99903 | 4.2471e-05
579.6 | 0.99903 | 4.6064e-05
579.8 | 0.99903 | 5.0246e-05
580 | 0.99903 | 5.4967e-05
580.2 | 0.99904 | 6.0069e-05
580.4 | 0.99904 | 6.5235e-05
580.6 | 0.99904 | 6.995e-05
580.8 | 0.99904 | 7.3542e-05
581 | 0.99904 | 7.5325e-05
581.2 | 0.99904 | 7.4836e-05
581.4 | 0.99905 | 7.2044e-05
581.6 | 0.99905 | 6.7379e-05
581.8 | 0.99906 | 6.1551e-05
582 | 0.99907 | 5.5298e-05
582.2 | 0.99907 | 4.9196e-05
582.4 | 0.99908 | 4.3596e-05
582.6 | 0.99908 | 3.8661e-05
582.8 | 0.99909 | 3.4422e-05
583 | 0.99909 | 3.0839e-05
583.2 | 0.99909 | 2.7843e-05
583.4 | 0.99909 | 2.5352e-05
583.6 | 0.99909 | 2.3293e-05
583.8 | 0.99908 | 2.1598e-05
584 | 0.99908 | 2.0211e-05
584.2 | 0.99908 | 1.9086e-05
584.4 | 0.99907 | 1.8188e-05
584.6 | 0.99907 | 1.7485e-05
584.8 | 0.99906 | 1.6957e-05
585 | 0.99906 | 1.6585e-05
585.2 | 0.99905 | 1.6359e-05
585.4 | 0.99904 | 1.6271e-05
585.6 | 0.99904 | 1.6318e-05
585.8 | 0.99903 | 1.6499e-05
586 | 0.99902 | 1.6821e-05
586.2 | 0.99902 | 1.7291e-05
586.4 | 0.99901 | 1.7925e-05
586.6 | 0.999 | 1.8743e-05
586.8 | 0.999 | 1.9773e-05
587 | 0.99899 | 2.1053e-05
587.2 | 0.99898 | 2.2634e-05
587.4 | 0.99898 | 2.4584e-05
587.6 | 0.99897 | 2.6999e-05
587.8 | 0.99896 | 3.0007e-05
588 | 0.99896 | 3.3792e-05
588.2 | 0.99895 | 3.8614e-05
588.4 | 0.99895 | 4.4858e-05
588.6 | 0.99894 | 5.3106e-05
588.8 | 0.99893 | 6.4268e-05
589 | 0.99891 | 7.9815e-05
589.2 | 0.99889 | 0.00010223
589.4 | 0.99886 | 0.00013583
589.6 | 0.99882 | 0.00018828
589.8 | 0.99874 | 0.0002722
590 | 0.99861 | 0.00040064
590.2 | 0.99847 | 0.00055374
590.4 | 0.99837 | 0.0006532
590.6 | 0.99837 | 0.00066454
590.8 | 0.99847 | 0.00057094
591 | 0.99864 | 0.00041387
591.2 | 0.99878 | 0.00028231
591.4 | 0.99887 | 0.00019589
591.6 | 0.99893 | 0.00014157
591.8 | 0.99898 | 0.00010668
592 | 0.99901 | 8.341e-05
592.2 | 0.99903 | 6.73e-05
592.4 | 0.99905 | 5.5773e-05
592.6 | 0.99907 | 4.7293e-05
592.8 | 0.99908 | 4.0909e-05
593 | 0.99909 | 3.6013e-05
593.2 | 0.99911 | 3.2205e-05
593.4 | 0.99911 | 2.921e-05
593.6 | 0.99912 | 2.6839e-05
593.8 | 0.99913 | 2.4959e-05
594 | 0.99914 | 2.3472e-05
594.2 | 0.99914 | 2.2309e-05
594.4 | 0.99915 | 2.1417e-05
594.6 | 0.99915 | 2.0761e-05
594.8 | 0.99916 | 2.0314e-05
595 | 0.99916 | 2.0057e-05
595.2 | 0.99916 | 1.9981e-05
595.4 | 0.99917 | 2.008e-05
595.6 | 0.99917 | 2.0357e-05
595.8 | 0.99917 | 2.0817e-05
596 | 0.99917 | 2.1475e-05
596.2 | 0.99917 | 2.2351e-05
596.4 | 0.99917 | 2.3472e-05
596.6 | 0.99917 | 2.4877e-05
596.8 | 0.99917 | 2.6616e-05
597 | 0.99916 | 2.8757e-05
597.2 | 0.99916 | 3.1386e-05
597.4 | 0.99916 | 3.4621e-05
597.6 | 0.99915 | 3.8615e-05
597.8 | 0.99915 | 4.3573e-05
598 | 0.99914 | 4.9764e-05
598.2 | 0.99913 | 5.755e-05
598.4 | 0.99912 | 6.7397e-05
598.6 | 0.9991 | 7.9893e-05
598.8 | 0.99909 | 9.5718e-05
599 | 0.99907 | 0.0001155
599.2 | 0.99904 | 0.00013944
599.4 | 0.99902 | 0.00016657
599.6 | 0.99899 | 0.00019374
599.8 | 0.99897 | 0.00021536
600 | 0.99896 | 0.00022527
600.2 | 0.99896 | 0.00022035
600.4 | 0.99898 | 0.00020274
600.6 | 0.99901 | 0.00017811
600.8 | 0.99904 | 0.00015212
601 | 0.99906 | 0.00012835
601.2 | 0.99909 | 0.00010818
601.4 | 0.99911 | 9.1728e-05
601.6 | 0.99913 | 7.8535e-05
601.8 | 0.99914 | 6.8014e-05
602 | 0.99915 | 5.9616e-05
602.2 | 0.99917 | 5.2886e-05
602.4 | 0.99918 | 4.7467e-05
602.6 | 0.99919 | 4.3084e-05
602.8 | 0.9992 | 3.9527e-05
603 | 0.99921 | 3.6638e-05
603.2 | 0.99921 | 3.4296e-05
603.4 | 0.99922 | 3.2409e-05
603.6 | 0.99923 | 3.0909e-05
603.8 | 0.99924 | 2.9743e-05
604 | 0.99925 | 2.8874e-05
604.2 | 0.99925 | 2.8275e-05
604.4 | 0.99926 | 2.793e-05
604.6 | 0.99927 | 2.7829e-05
604.8 | 0.99928 | 2.7972e-05
605 | 0.99928 | 2.8369e-05
605.2 | 0.99929 | 2.9035e-05
605.4 | 0.99929 | 2.9998e-05
605.6 | 0.9993 | 3.1299e-05
605.8 | 0.9993 | 3.2994e-05
606 | 0.99931 | 3.5161e-05
606.2 | 0.99931 | 3.7906e-05
606.4 | 0.99931 | 4.1376e-05
606.6 | 0.99931 | 4.5776e-05
606.8 | 0.99931 | 5.1391e-05
607 | 0.99931 | 5.8631e-05
607.2 | 0.9993 | 6.809e-05
607.4 | 0.99929 | 8.0639e-05
607.6 | 0.99928 | 9.7569e-05
607.8 | 0.99926 | 0.00012076
608 | 0.99923 | 0.00015278
608.2 | 0.99919 | 0.00019649
608.4 | 0.99913 | 0.00025284
608.6 | 0.99907 | 0.00031576
608.8 | 0.99902 | 0.00036919
609 | 0.99899 | 0.00039692
609.2 | 0.999 | 0.00039024
609.4 | 0.99904 | 0.00034734
609.6 | 0.99911 | 0.0002833
609.8 | 0.99917 | 0.00022029
610 | 0.99922 | 0.00016953
610.2 | 0.99925 | 0.0001319
610.4 | 0.99928 | 0.00010467
610.6 | 0.9993 | 8.4891e-05
610.8 | 0.99931 | 7.0331e-05
611 | 0.99932 | 5.9422e-05
611.2 | 0.99932 | 5.1103e-05
611.4 | 0.99933 | 4.4656e-05
611.6 | 0.99933 | 3.9585e-05
611.8 | 0.99933 | 3.5547e-05
612 | 0.99933 | 3.2298e-05
612.2 | 0.99933 | 2.9661e-05
612.4 | 0.99933 | 2.7507e-05
612.6 | 0.99932 | 2.5741e-05
612.8 | 0.99932 | 2.4291e-05
613 | 0.99932 | 2.3101e-05
613.2 | 0.99931 | 2.2129e-05
613.4 | 0.99931 | 2.1343e-05
613.6 | 0.99931 | 2.0718e-05
613.8 | 0.9993 | 2.0234e-05
614 | 0.9993 | 1.9875e-05
614.2 | 0.9993 | 1.9631e-05
614.4 | 0.99929 | 1.9491e-05
614.6 | 0.99929 | 1.9451e-05
614.8 | 0.99929 | 1.9506e-05
615 | 0.99929 | 1.9654e-05
615.2 | 0.99928 | 1.9893e-05
615.4 | 0.99928 | 2.0226e-05
615.6 | 0.99928 | 2.0655e-05
615.8 | 0.99928 | 2.1183e-05
616 | 0.99928 | 2.1818e-05
616.2 | 0.99928 | 2.2565e-05
616.4 | 0.99928 | 2.3436e-05
616.6 | 0.99928 | 2.4439e-05
616.8 | 0.99928 | 2.5589e-05
617 | 0.99928 | 2.6901e-05
617.2 | 0.99928 | 2.839e-05
617.4 | 0.99929 | 3.0076e-05
617.6 | 0.99929 | 3.1978e-05
617.8 | 0.99929 | 3.4115e-05
618 | 0.99929 | 3.6506e-05
618.2 | 0.9993 | 3.9165e-05
618.4 | 0.9993 | 4.2097e-05
618.6 | 0.9993 | 4.5295e-05
618.8 | 0.9993 | 4.873e-05
619 | 0.99931 | 5.2345e-05
619.2 | 0.99931 | 5.6044e-05
619.4 | 0.99931 | 5.9684e-05
619.6 | 0.99932 | 6.3078e-05
619.8 | 0.99932 | 6.6005e-05
620 | 0.99932 | 6.8232e-05
620.2 | 0.99933 | 6.9553e-05
620.4 | 0.99933 | 6.9831e-05
620.6 | 0.99934 | 6.9032e-05
620.8 | 0.99935 | 6.7228e-05
621 | 0.99936 | 6.4587e-05
621.2 | 0.99937 | 6.1332e-05
621.4 | 0.99938 | 5.7702e-05
621.6 | 0.99938 | 5.3914e-05
621.8 | 0.99939 | 5.0144e-05
622 | 0.9994 | 4.6522e-05
622.2 | 0.99941 | 4.3131e-05
622.4 | 0.99942 | 4.0017e-05
622.6 | 0.99942 | 3.7198e-05
622.8 | 0.99943 | 3.4675e-05
623 | 0.99943 | 3.2435e-05
623.2 | 0.99944 | 3.0462e-05
623.4 | 0.99944 | 2.8733e-05
623.6 | 0.99944 | 2.7227e-05
623.8 | 0.99945 | 2.5922e-05
624 | 0.99945 | 2.4801e-05
624.2 | 0.99945 | 2.3844e-05
624.4 | 0.99945 | 2.3038e-05
624.6 | 0.99945 | 2.2368e-05
624.8 | 0.99945 | 2.1825e-05
625 | 0.99945 | 2.1398e-05
625.2 | 0.99945 | 2.108e-05
625.4 | 0.99945 | 2.0865e-05
625.6 | 0.99945 | 2.075e-05
625.8 | 0.99945 | 2.0731e-05
626 | 0.99945 | 2.0806e-05
626.2 | 0.99944 | 2.0976e-05
626.4 | 0.99944 | 2.1241e-05
626.6 | 0.99944 | 2.1604e-05
626.8 | 0.99943 | 2.2068e-05
627 | 0.99943 | 2.2637e-05
627.2 | 0.99943 | 2.332e-05
627.4 | 0.99942 | 2.4122e-05
627.6 | 0.99942 | 2.5055e-05
627.8 | 0.99941 | 2.6129e-05
628 | 0.99941 | 2.7359e-05
628.2 | 0.9994 | 2.8759e-05
628.4 | 0.9994 | 3.0349e-05
628.6 | 0.99939 | 3.2151e-05
628.8 | 0.99939 | 3.4188e-05
629 | 0.99938 | 3.649e-05
629.2 | 0.99938 | 3.9089e-05
629.4 | 0.99937 | 4.2022e-05
629.6 | 0.99937 | 4.5331e-05
629.8 | 0.99936 | 4.9062e-05
630 | 0.99936 | 5.3264e-05
630.2 | 0.99935 | 5.7992e-05
630.4 | 0.99935 | 6.3298e-05
630.6 | 0.99934 | 6.9234e-05
630.8 | 0.99933 | 7.5842e-05
631 | 0.99933 | 8.3149e-05
631.2 | 0.99932 | 9.1149e-05
631.4 | 0.99932 | 9.9795e-05
631.6 | 0.99931 | 0.00010897
631.8 | 0.9993 | 0.00011849
632 | 0.9993 | 0.00012807
632.2 | 0.99929 | 0.0001373
632.4 | 0.99929 | 0.00014574
632.6 | 0.99929 | 0.00015288
632.8 | 0.99929 | 0.00015826
633 | 0.99929 | 0.00016153
633.2 | 0.99929 | 0.00016253
633.4 | 0.9993 | 0.00016129
633.6 | 0.99931 | 0.00015806
633.8 | 0.99932 | 0.00015321
634 | 0.99933 | 0.00014718
634.2 | 0.99935 | 0.00014042
634.4 | 0.99936 | 0.00013332
634.6 | 0.99937 | 0.00012621
634.8 | 0.99939 | 0.00011933
635 | 0.9994 | 0.00011283
635.2 | 0.99941 | 0.00010682
635.4 | 0.99942 | 0.00010135
635.6 | 0.99943 | 9.6451e-05
635.8 | 0.99944 | 9.2107e-05
636 | 0.99945 | 8.8307e-05
636.2 | 0.99946 | 8.5029e-05
636.4 | 0.99947 | 8.2245e-05
636.6 | 0.99947 | 7.9931e-05
636.8 | 0.99948 | 7.8061e-05
637 | 0.99949 | 7.6616e-05
637.2 | 0.99949 | 7.5578e-05
637.4 | 0.99949 | 7.4935e-05
637.6 | 0.99949 | 7.468e-05
637.8 | 0.9995 | 7.4809e-05
638 | 0.9995 | 7.5326e-05
638.2 | 0.9995 | 7.6237e-05
638.4 | 0.99949 | 7.7558e-05
638.6 | 0.99949 | 7.9306e-05
638.8 | 0.99949 | 8.1508e-05
639 | 0.99948 | 8.4197e-05
639.2 | 0.99948 | 8.7413e-05
639.4 | 0.99947 | 9.1205e-05
639.6 | 0.99946 | 9.5631e-05
639.8 | 0.99946 | 0.00010076
640 | 0.99944 | 0.00010667
640.2 | 0.99943 | 0.00011345
640.4 | 0.99942 | 0.00012121
640.6 | 0.9994 | 0.00013008
640.8 | 0.99938 | 0.00014017
641 | 0.99936 | 0.00015166
641.2 | 0.99934 | 0.00016469
641.4 | 0.99932 | 0.00017947
641.6 | 0.99929 | 0.00019617
641.8 | 0.99926 | 0.00021501
642 | 0.99923 | 0.00023617
642.2 | 0.9992 | 0.00025987
642.4 | 0.99916 | 0.00028625
642.6 | 0.99912 | 0.00031546
642.8 | 0.99908 | 0.00034753
643 | 0.99903 | 0.00038244
643.2 | 0.99899 | 0.00041999
643.4 | 0.99894 | 0.00045985
643.6 | 0.99889 | 0.00050145
643.8 | 0.99884 | 0.00054401
644 | 0.9988 | 0.00058646
644.2 | 0.99875 | 0.00062752
644.4 | 0.99871 | 0.0006657
644.6 | 0.99868 | 0.0006994
644.8 | 0.99866 | 0.00072706
645 | 0.99864 | 0.00074732
645.2 | 0.99864 | 0.00075915
645.4 | 0.99865 | 0.00076205
645.6 | 0.99866 | 0.00075605
645.8 | 0.99869 | 0.00074179
646 | 0.99872 | 0.00072033
646.2 | 0.99876 | 0.00069309
646.4 | 0.99881 | 0.00066162
646.6 | 0.99885 | 0.00062747
646.8 | 0.9989 | 0.00059204
647 | 0.99895 | 0.00055652
647.2 | 0.999 | 0.00052187
647.4 | 0.99904 | 0.00048877
647.6 | 0.99909 | 0.0004577
647.8 | 0.99913 | 0.00042895
648 | 0.99916 | 0.00040265
648.2 | 0.99919 | 0.00037884
648.4 | 0.99922 | 0.00035748
648.6 | 0.99925 | 0.00033846
648.8 | 0.99927 | 0.00032168
649 | 0.99929 | 0.000307
649.2 | 0.99931 | 0.00029428
649.4 | 0.99933 | 0.0002834
649.6 | 0.99934 | 0.00027424
649.8 | 0.99935 | 0.00026669
650 | 0.99936 | 0.00026067
650.2 | 0.99936 | 0.00025609
650.4 | 0.99937 | 0.00025291
650.6 | 0.99937 | 0.00025109
650.8 | 0.99937 | 0.00025061
651 | 0.99937 | 0.00025147
651.2 | 0.99936 | 0.00025368
651.4 | 0.99936 | 0.00025729
651.6 | 0.99935 | 0.00026236
651.8 | 0.99934 | 0.00026897
652 | 0.99933 | 0.00027724
652.2 | 0.99931 | 0.00028729
652.4 | 0.9993 | 0.00029931
652.6 | 0.99928 | 0.00031351
652.8 | 0.99926 | 0.00033013
653 | 0.99923 | 0.00034948
653.2 | 0.9992 | 0.00037191
653.4 | 0.99917 | 0.00039785
653.6 | 0.99913 | 0.00042778
653.8 | 0.99909 | 0.00046229
654 | 0.99905 | 0.00050202
654.2 | 0.99899 | 0.00054773
654.4 | 0.99893 | 0.00060025
654.6 | 0.99887 | 0.00066047
654.8 | 0.99879 | 0.00072931
655 | 0.99871 | 0.00080763
655.2 | 0.99861 | 0.0008961
655.4 | 0.99851 | 0.00099496
655.6 | 0.99839 | 0.0011038
655.8 | 0.99827 | 0.001221
656 | 0.99814 | 0.0013435
656.2 | 0.99802 | 0.0014664
656.4 | 0.9979 | 0.0015828
656.6 | 0.99779 | 0.0016842
656.8 | 0.99771 | 0.0017619
657 | 0.99767 | 0.0018087
657.2 | 0.99766 | 0.0018198
657.4 | 0.99768 | 0.0017949
657.6 | 0.99774 | 0.0017373
657.8 | 0.99783 | 0.0016537
658 | 0.99793 | 0.0015522
658.2 | 0.99805 | 0.0014409
658.4 | 0.99817 | 0.0013268
658.6 | 0.99829 | 0.0012151
658.8 | 0.9984 | 0.0011094
659 | 0.9985 | 0.0010119
659.2 | 0.9986 | 0.00092341
659.4 | 0.99868 | 0.00084414
659.6 | 0.99876 | 0.00077374
659.8 | 0.99883 | 0.0007116
660 | 0.99889 | 0.000657
660.2 | 0.99895 | 0.00060916
660.4 | 0.99899 | 0.00056736
660.6 | 0.99904 | 0.00053089
660.8 | 0.99908 | 0.00049913
661 | 0.99911 | 0.00047155
661.2 | 0.99914 | 0.00044765
661.4 | 0.99917 | 0.00042702
661.6 | 0.99919 | 0.00040928
661.8 | 0.99921 | 0.00039413
662 | 0.99923 | 0.0003813
662.2 | 0.99924 | 0.00037054
662.4 | 0.99925 | 0.00036165
662.6 | 0.99926 | 0.00035446
662.8 | 0.99927 | 0.00034879
663 | 0.99928 | 0.00034449
663.2 | 0.99928 | 0.00034143
663.4 | 0.99929 | 0.00033947
663.6 | 0.99929 | 0.00033847
663.8 | 0.99929 | 0.00033829
664 | 0.99929 | 0.00033878
664.2 | 0.99929 | 0.00033976
664.4 | 0.99929 | 0.00034107
664.6 | 0.99928 | 0.00034251
664.8 | 0.99928 | 0.00034387
665 | 0.99928 | 0.00034492
665.2 | 0.99927 | 0.00034544
665.4 | 0.99927 | 0.00034521
665.6 | 0.99927 | 0.00034404
665.8 | 0.99927 | 0.00034173
666 | 0.99927 | 0.00033818
666.2 | 0.99927 | 0.00033329
666.4 | 0.99927 | 0.00032704
666.6 | 0.99927 | 0.0003195
666.8 | 0.99927 | 0.00031075
667 | 0.99928 | 0.00030094
667.2 | 0.99928 | 0.00029026
667.4 | 0.99929 | 0.00027892
667.6 | 0.9993 | 0.00026711
667.8 | 0.9993 | 0.00025505
668 | 0.99931 | 0.00024291
668.2 | 0.99931 | 0.00023085
668.4 | 0.99932 | 0.000219
668.6 | 0.99933 | 0.00020747
668.8 | 0.99933 | 0.00019634
669 | 0.99934 | 0.00018565
669.2 | 0.99934 | 0.00017545
669.4 | 0.99935 | 0.00016575
669.6 | 0.99936 | 0.00015657
669.8 | 0.99936 | 0.00014788
670 | 0.99937 | 0.0001397
670.2 | 0.99938 | 0.00013199
670.4 | 0.99938 | 0.00012475
670.6 | 0.99939 | 0.00011795
670.8 | 0.9994 | 0.00011157
671 | 0.99941 | 0.0001056
671.2 | 0.99942 | 0.00010002
671.4 | 0.99942 | 9.4805e-05
671.6 | 0.99943 | 8.9939e-05
671.8 | 0.99944 | 8.5404e-05
672 | 0.99945 | 8.1186e-05
672.2 | 0.99946 | 7.7268e-05
672.4 | 0.99947 | 7.3636e-05
672.6 | 0.99949 | 7.0276e-05
672.8 | 0.9995 | 6.7175e-05
673 | 0.99951 | 6.4321e-05
673.2 | 0.99952 | 6.17e-05
673.4 | 0.99953 | 5.9303e-05
673.6 | 0.99954 | 5.7117e-05
673.8 | 0.99955 | 5.5135e-05
674 | 0.99956 | 5.3345e-05
674.2 | 0.99957 | 5.1741e-05
674.4 | 0.99958 | 5.0314e-05
674.6 | 0.99959 | 4.9057e-05
674.8 | 0.9996 | 4.7965e-05
675 | 0.9996 | 4.7033e-05
675.2 | 0.99961 | 4.6256e-05
675.4 | 0.99962 | 4.5631e-05
675.6 | 0.99963 | 4.5158e-05
675.8 | 0.99963 | 4.4833e-05
676 | 0.99964 | 4.4659e-05
676.2 | 0.99965 | 4.4636e-05
676.4 | 0.99965 | 4.4768e-05
676.6 | 0.99966 | 4.5059e-05
676.8 | 0.99966 | 4.5516e-05
677 | 0.99967 | 4.6146e-05
677.2 | 0.99967 | 4.6959e-05
677.4 | 0.99967 | 4.7968e-05
677.6 | 0.99968 | 4.9188e-05
677.8 | 0.99968 | 5.0636e-05
678 | 0.99968 | 5.2335e-05
678.2 | 0.99968 | 5.4308e-05
678.4 | 0.99968 | 5.6587e-05
678.6 | 0.99968 | 5.9205e-05
678.8 | 0.99968 | 6.2204e-05
679 | 0.99968 | 6.5632e-05
679.2 | 0.99968 | 6.9545e-05
679.4 | 0.99967 | 7.4009e-05
679.6 | 0.99967 | 7.9099e-05
679.8 | 0.99966 | 8.4906e-05
680 | 0.99966 | 9.1532e-05
680.2 | 0.99965 | 9.9096e-05
680.4 | 0.99964 | 0.00010774
680.6 | 0.99963 | 0.00011761
680.8 | 0.99962 | 0.00012889
681 | 0.99961 | 0.00014178
681.2 | 0.99959 | 0.00015648
681.4 | 0.99958 | 0.00017323
681.6 | 0.99955 | 0.00019224
681.8 | 0.99953 | 0.00021374
682 | 0.99951 | 0.00023788
682.2 | 0.99948 | 0.00026478
682.4 | 0.99944 | 0.00029443
682.6 | 0.99941 | 0.00032667
682.8 | 0.99937 | 0.00036113
683 | 0.99933 | 0.00039723
683.2 | 0.99929 | 0.00043419
683.4 | 0.99925 | 0.000471
683.6 | 0.99921 | 0.0005066
683.8 | 0.99917 | 0.00053991
684 | 0.99914 | 0.00057004
684.2 | 0.9991 | 0.00059634
684.4 | 0.99908 | 0.00061848
684.6 | 0.99905 | 0.00063652
684.8 | 0.99903 | 0.00065076
685 | 0.99901 | 0.00066177
685.2 | 0.99899 | 0.00067023
685.4 | 0.99898 | 0.00067686
685.6 | 0.99897 | 0.00068238
685.8 | 0.99895 | 0.00068743
686 | 0.99894 | 0.00069255
686.2 | 0.99892 | 0.00069819
686.4 | 0.9989 | 0.0007047
686.6 | 0.99889 | 0.00071234
686.8 | 0.99887 | 0.00072129
687 | 0.99884 | 0.00073165
687.2 | 0.99882 | 0.00074349
687.4 | 0.9988 | 0.0007568
687.6 | 0.99877 | 0.00077154
687.8 | 0.99875 | 0.00078763
688 | 0.99872 | 0.00080498
688.2 | 0.9987 | 0.00082345
688.4 | 0.99867 | 0.00084289
688.6 | 0.99864 | 0.00086317
688.8 | 0.99862 | 0.00088414
689 | 0.9986 | 0.00090566
689.2 | 0.99857 | 0.0009276
689.4 | 0.99855 | 0.00094987
689.6 | 0.99853 | 0.0009724
689.8 | 0.99851 | 0.00099514
690 | 0.9985 | 0.0010181
690.2 | 0.99848 | 0.0010412
690.4 | 0.99846 | 0.0010646
690.6 | 0.99845 | 0.0010883
690.8 | 0.99844 | 0.0011123
691 | 0.99842 | 0.0011367
691.2 | 0.99841 | 0.0011613
691.4 | 0.9984 | 0.0011862
691.6 | 0.99838 | 0.001211
691.8 | 0.99837 | 0.0012354
692 | 0.99836 | 0.001259
692.2 | 0.99835 | 0.0012811
692.4 | 0.99834 | 0.0013007
692.6 | 0.99833 | 0.001317
692.8 | 0.99833 | 0.0013288
693 | 0.99833 | 0.001335
693.2 | 0.99834 | 0.0013343
693.4 | 0.99836 | 0.0013259
693.6 | 0.99838 | 0.0013091
693.8 | 0.99842 | 0.0012837
694 | 0.99846 | 0.0012499
694.2 | 0.9985 | 0.0012084
694.4 | 0.99856 | 0.0011604
694.6 | 0.99861 | 0.0011073
694.8 | 0.99868 | 0.0010506
695 | 0.99874 | 0.00099203
695.2 | 0.9988 | 0.00093309
695.4 | 0.99886 | 0.00087509
695.6 | 0.99892 | 0.00081909
695.8 | 0.99897 | 0.00076587
696 | 0.99902 | 0.00071598
696.2 | 0.99907 | 0.00066971
696.4 | 0.99911 | 0.00062719
696.6 | 0.99915 | 0.00058842
696.8 | 0.99919 | 0.00055328
697 | 0.99922 | 0.0005216
697.2 | 0.99924 | 0.00049317
697.4 | 0.99927 | 0.00046774
697.6 | 0.99929 | 0.00044509
697.8 | 0.9993 | 0.00042497
698 | 0.99932 | 0.00040718
698.2 | 0.99933 | 0.00039149
698.4 | 0.99934 | 0.00037773
698.6 | 0.99935 | 0.00036572
698.8 | 0.99936 | 0.0003553
699 | 0.99936 | 0.00034633
699.2 | 0.99936 | 0.00033869
699.4 | 0.99936 | 0.00033226
699.6 | 0.99937 | 0.00032695
699.8 | 0.99936 | 0.00032266
700 | 0.99936 | 0.00031931
700.2 | 0.99936 | 0.00031683
700.4 | 0.99935 | 0.00031514
700.6 | 0.99935 | 0.00031419
700.8 | 0.99935 | 0.00031391
701 | 0.99934 | 0.00031426
701.2 | 0.99933 | 0.00031517
701.4 | 0.99933 | 0.0003166
701.6 | 0.99932 | 0.0003185
701.8 | 0.99931 | 0.00032082
702 | 0.99931 | 0.00032351
702.2 | 0.9993 | 0.00032653
702.4 | 0.99929 | 0.00032983
702.6 | 0.99929 | 0.00033337
702.8 | 0.99928 | 0.00033709
703 | 0.99928 | 0.00034095
703.2 | 0.99927 | 0.00034491
703.4 | 0.99927 | 0.00034891
703.6 | 0.99926 | 0.0003529
703.8 | 0.99926 | 0.00035684
704 | 0.99926 | 0.00036066
704.2 | 0.99926 | 0.00036431
704.4 | 0.99926 | 0.00036773
704.6 | 0.99926 | 0.00037085
704.8 | 0.99926 | 0.0003736
705 | 0.99926 | 0.0003759
705.2 | 0.99927 | 0.00037769
705.4 | 0.99927 | 0.00037886
705.6 | 0.99927 | 0.00037935
705.8 | 0.99928 | 0.00037906
706 | 0.99929 | 0.00037792
706.2 | 0.9993 | 0.00037586
706.4 | 0.99931 | 0.00037283
706.6 | 0.99932 | 0.00036877
706.8 | 0.99933 | 0.00036368
707 | 0.99934 | 0.00035756
707.2 | 0.99935 | 0.00035044
707.4 | 0.99937 | 0.00034239
707.6 | 0.99938 | 0.0003335
707.8 | 0.9994 | 0.00032386
708 | 0.99941 | 0.00031362
708.2 | 0.99943 | 0.00030291
708.4 | 0.99944 | 0.00029188
708.6 | 0.99946 | 0.00028068
708.8 | 0.99948 | 0.00026944
709 | 0.99949 | 0.00025829
709.2 | 0.99951 | 0.00024736
709.4 | 0.99952 | 0.00023674
709.6 | 0.99953 | 0.00022651
709.8 | 0.99955 | 0.00021674
710 | 0.99956 | 0.00020747
710.2 | 0.99957 | 0.00019874
710.4 | 0.99958 | 0.00019058
710.6 | 0.99959 | 0.00018298
710.8 | 0.9996 | 0.00017596
711 | 0.99961 | 0.0001695
711.2 | 0.99962 | 0.0001636
711.4 | 0.99962 | 0.00015824
711.6 | 0.99963 | 0.00015341
711.8 | 0.99964 | 0.00014909
712 | 0.99964 | 0.00014525
712.2 | 0.99964 | 0.00014189
712.4 | 0.99965 | 0.00013899
712.6 | 0.99965 | 0.00013653
712.8 | 0.99965 | 0.0001345
713 | 0.99965 | 0.00013288
713.2 | 0.99965 | 0.00013167
713.4 | 0.99965 | 0.00013087
713.6 | 0.99965 | 0.00013046
713.8 | 0.99965 | 0.00013046
714 | 0.99965 | 0.00013085
714.2 | 0.99965 | 0.00013164
714.4 | 0.99964 | 0.00013284
714.6 | 0.99964 | 0.00013446
714.8 | 0.99964 | 0.00013651
715 | 0.99963 | 0.000139
715.2 | 0.99962 | 0.00014196
715.4 | 0.99962 | 0.00014541
715.6 | 0.99961 | 0.00014936
715.8 | 0.9996 | 0.00015386
716 | 0.99959 | 0.00015892
716.2 | 0.99958 | 0.00016459
716.4 | 0.99957 | 0.00017089
716.6 | 0.99956 | 0.00017787
716.8 | 0.99955 | 0.00018557
717 | 0.99953 | 0.00019401
717.2 | 0.99952 | 0.00020326
717.4 | 0.9995 | 0.00021333
717.6 | 0.99948 | 0.00022426
717.8 | 0.99947 | 0.00023609
718 | 0.99945 | 0.00024882
718.2 | 0.99943 | 0.00026246
718.4 | 0.99941 | 0.000277
718.6 | 0.99938 | 0.00029241
718.8 | 0.99936 | 0.00030862
719 | 0.99934 | 0.00032553
719.2 | 0.99932 | 0.00034303
719.4 | 0.99929 | 0.00036093
719.6 | 0.99927 | 0.00037901
719.8 | 0.99925 | 0.00039702
720 | 0.99922 | 0.00041463
720.2 | 0.9992 | 0.0004315
720.4 | 0.99919 | 0.00044724
720.6 | 0.99917 | 0.00046144
720.8 | 0.99916 | 0.0004737
721 | 0.99914 | 0.00048364
721.2 | 0.99914 | 0.00049091
721.4 | 0.99913 | 0.00049523
721.6 | 0.99913 | 0.00049644
721.8 | 0.99914 | 0.00049445
722 | 0.99915 | 0.00048929
722.2 | 0.99916 | 0.0004811
722.4 | 0.99918 | 0.00047013
722.6 | 0.99919 | 0.00045671
722.8 | 0.99922 | 0.00044123
723 | 0.99924 | 0.0004241
723.2 | 0.99926 | 0.00040576
723.4 | 0.99929 | 0.00038664
723.6 | 0.99932 | 0.00036712
723.8 | 0.99934 | 0.00034755
724 | 0.99937 | 0.00032823
724.2 | 0.9994 | 0.00030941
724.4 | 0.99942 | 0.00029126
724.6 | 0.99945 | 0.00027393
724.8 | 0.99947 | 0.00025752
725 | 0.99949 | 0.00024209
725.2 | 0.99951 | 0.00022766
725.4 | 0.99953 | 0.00021424
725.6 | 0.99955 | 0.00020181
725.8 | 0.99957 | 0.00019033
726 | 0.99958 | 0.00017977
726.2 | 0.9996 | 0.00017009
726.4 | 0.99961 | 0.00016122
726.6 | 0.99962 | 0.00015312
726.8 | 0.99963 | 0.00014575
727 | 0.99964 | 0.00013904
727.2 | 0.99965 | 0.00013294
727.4 | 0.99966 | 0.00012743
727.6 | 0.99967 | 0.00012244
727.8 | 0.99968 | 0.00011795
728 | 0.99968 | 0.00011391
728.2 | 0.99969 | 0.0001103
728.4 | 0.99969 | 0.00010707
728.6 | 0.9997 | 0.00010422
728.8 | 0.9997 | 0.0001017
729 | 0.9997 | 9.9509e-05
729.2 | 0.99971 | 9.7615e-05
729.4 | 0.99971 | 9.6003e-05
729.6 | 0.99971 | 9.466e-05
729.8 | 0.99971 | 9.357e-05
730 | 0.99971 | 9.2723e-05
730.2 | 0.99971 | 9.2109e-05
730.4 | 0.99971 | 9.1718e-05
730.6 | 0.99971 | 9.1543e-05
730.8 | 0.99971 | 9.1577e-05
731 | 0.99971 | 9.1815e-05
731.2 | 0.99971 | 9.2251e-05
731.4 | 0.9997 | 9.288e-05
731.6 | 0.9997 | 9.37e-05
731.8 | 0.9997 | 9.4704e-05
732 | 0.99969 | 9.5891e-05
732.2 | 0.99969 | 9.7255e-05
732.4 | 0.99968 | 9.8791e-05
732.6 | 0.99968 | 0.00010049
732.8 | 0.99967 | 0.00010236
733 | 0.99967 | 0.00010438
733.2 | 0.99966 | 0.00010654
733.4 | 0.99966 | 0.00010883
733.6 | 0.99965 | 0.00011124
733.8 | 0.99964 | 0.00011376
734 | 0.99964 | 0.00011636
734.2 | 0.99963 | 0.00011902
734.4 | 0.99962 | 0.00012172
734.6 | 0.99961 | 0.00012443
734.8 | 0.99961 | 0.00012712
735 | 0.9996 | 0.00012976
735.2 | 0.99959 | 0.0001323
735.4 | 0.99958 | 0.00013472
735.6 | 0.99958 | 0.00013696
735.8 | 0.99957 | 0.00013899
736 | 0.99956 | 0.00014077
736.2 | 0.99956 | 0.00014226
736.4 | 0.99955 | 0.00014342
736.6 | 0.99955 | 0.00014422
736.8 | 0.99954 | 0.00014463
737 | 0.99954 | 0.00014464
737.2 | 0.99954 | 0.00014423
737.4 | 0.99954 | 0.00014339
737.6 | 0.99954 | 0.00014212
737.8 | 0.99954 | 0.00014044
738 | 0.99954 | 0.00013837
738.2 | 0.99954 | 0.00013592
738.4 | 0.99955 | 0.00013314
738.6 | 0.99955 | 0.00013005
738.8 | 0.99956 | 0.0001267
739 | 0.99956 | 0.00012314
739.2 | 0.99957 | 0.0001194
739.4 | 0.99958 | 0.00011553
739.6 | 0.99959 | 0.00011158
739.8 | 0.9996 | 0.00010758
740 | 0.99961 | 0.00010357
740.2 | 0.99961 | 9.9589e-05
740.4 | 0.99962 | 9.5665e-05
740.6 | 0.99963 | 9.1823e-05
740.8 | 0.99964 | 8.8086e-05
741 | 0.99965 | 8.4473e-05
741.2 | 0.99966 | 8.0997e-05
741.4 | 0.99967 | 7.7669e-05
741.6 | 0.99968 | 7.4498e-05
741.8 | 0.99969 | 7.1488e-05
742 | 0.9997 | 6.8642e-05
742.2 | 0.99971 | 6.596e-05
742.4 | 0.99971 | 6.3443e-05
742.6 | 0.99972 | 6.1087e-05
742.8 | 0.99973 | 5.8889e-05
743 | 0.99973 | 5.6846e-05
743.2 | 0.99974 | 5.4953e-05
743.4 | 0.99975 | 5.3205e-05
743.6 | 0.99975 | 5.1596e-05
743.8 | 0.99976 | 5.0123e-05
744 | 0.99976 | 4.878e-05
744.2 | 0.99977 | 4.7562e-05
744.4 | 0.99977 | 4.6465e-05
744.6 | 0.99978 | 4.5485e-05
744.8 | 0.99978 | 4.4617e-05
745 | 0.99979 | 4.3859e-05
745.2 | 0.99979 | 4.3208e-05
745.4 | 0.99979 | 4.266e-05
745.6 | 0.9998 | 4.2214e-05
745.8 | 0.9998 | 4.1869e-05
746 | 0.9998 | 4.1624e-05
746.2 | 0.9998 | 4.1478e-05
746.4 | 0.9998 | 4.1432e-05
746.6 | 0.9998 | 4.1487e-05
746.8 | 0.99981 | 4.1644e-05
747 | 0.99981 | 4.1905e-05
747.2 | 0.99981 | 4.2275e-05
747.4 | 0.99981 | 4.2756e-05
747.6 | 0.99981 | 4.3354e-05
747.8 | 0.99981 | 4.4074e-05
748 | 0.99981 | 4.4923e-05
748.2 | 0.9998 | 4.591e-05
748.4 | 0.9998 | 4.7044e-05
748.6 | 0.9998 | 4.8335e-05
748.8 | 0.9998 | 4.9796e-05
749 | 0.9998 | 5.1442e-05
749.2 | 0.99979 | 5.3287e-05
749.4 | 0.99979 | 5.5351e-05
749.6 | 0.99978 | 5.7654e-05
749.8 | 0.99978 | 6.022e-05
750 | 0.99977 | 6.3075e-05
750.2 | 0.99977 | 6.625e-05
750.4 | 0.99976 | 6.9779e-05
750.6 | 0.99975 | 7.3699e-05
750.8 | 0.99975 | 7.8054e-05
751 | 0.99974 | 8.2891e-05
751.2 | 0.99973 | 8.8263e-05
751.4 | 0.99972 | 9.423e-05
751.6 | 0.9997 | 0.00010085
751.8 | 0.99969 | 0.00010821
752 | 0.99968 | 0.00011637
752.2 | 0.99966 | 0.00012541
752.4 | 0.99964 | 0.00013543
752.6 | 0.99962 | 0.00014649
752.8 | 0.9996 | 0.00015869
753 | 0.99958 | 0.00017211
753.2 | 0.99956 | 0.0001868
753.4 | 0.99953 | 0.00020281
753.6 | 0.9995 | 0.00022017
753.8 | 0.99947 | 0.00023885
754 | 0.99944 | 0.00025879
754.2 | 0.99941 | 0.00027987
754.4 | 0.99938 | 0.00030193
754.6 | 0.99935 | 0.00032472
754.8 | 0.99932 | 0.00034795
755 | 0.99929 | 0.00037125
755.2 | 0.99926 | 0.00039425
755.4 | 0.99923 | 0.0004165
755.6 | 0.99921 | 0.00043759
755.8 | 0.99919 | 0.00045709
756 | 0.99917 | 0.00047464
756.2 | 0.99916 | 0.00048992
756.4 | 0.99916 | 0.00050267
756.6 | 0.99915 | 0.00051273
756.8 | 0.99916 | 0.00052004
757 | 0.99916 | 0.00052463
757.2 | 0.99917 | 0.00052659
757.4 | 0.99918 | 0.00052609
757.6 | 0.9992 | 0.00052336
757.8 | 0.99921 | 0.00051865
758 | 0.99923 | 0.0005122
758.2 | 0.99925 | 0.00050426
758.4 | 0.99927 | 0.00049508
758.6 | 0.99929 | 0.00048486
758.8 | 0.99931 | 0.00047379
759 | 0.99933 | 0.00046204
759.2 | 0.99935 | 0.00044973
759.4 | 0.99938 | 0.00043701
759.6 | 0.9994 | 0.00042397
759.8 | 0.99942 | 0.00041071
760 | 0.99944 | 0.00039732
760.2 | 0.99945 | 0.00038389
760.4 | 0.99947 | 0.00037048
760.6 | 0.99949 | 0.00035718
760.8 | 0.99951 | 0.00034404
761 | 0.99953 | 0.00033114
761.2 | 0.99954 | 0.00031853
761.4 | 0.99956 | 0.00030627
761.6 | 0.99957 | 0.0002944
761.8 | 0.99959 | 0.00028296
762 | 0.9996 | 0.00027198
762.2 | 0.99961 | 0.0002615
762.4 | 0.99962 | 0.00025153
762.6 | 0.99963 | 0.00024208
762.8 | 0.99965 | 0.00023316
763 | 0.99965 | 0.00022478
763.2 | 0.99966 | 0.00021693
763.4 | 0.99967 | 0.00020961
763.6 | 0.99968 | 0.0002028
763.8 | 0.99968 | 0.0001965
764 | 0.99969 | 0.00019069
764.2 | 0.9997 | 0.00018536
764.4 | 0.9997 | 0.0001805
764.6 | 0.9997 | 0.00017608
764.8 | 0.99971 | 0.00017209
765 | 0.99971 | 0.00016852
765.2 | 0.99971 | 0.00016535
765.4 | 0.99971 | 0.00016257
765.6 | 0.99971 | 0.00016016
765.8 | 0.99971 | 0.00015812
766 | 0.99971 | 0.00015644
766.2 | 0.99971 | 0.00015509
766.4 | 0.99971 | 0.00015409
766.6 | 0.99971 | 0.00015341
766.8 | 0.9997 | 0.00015305
767 | 0.9997 | 0.00015301
767.2 | 0.9997 | 0.00015329
767.4 | 0.99969 | 0.00015387
767.6 | 0.99969 | 0.00015475
767.8 | 0.99968 | 0.00015594
768 | 0.99968 | 0.00015742
768.2 | 0.99967 | 0.0001592
768.4 | 0.99966 | 0.00016126
768.6 | 0.99965 | 0.00016361
768.8 | 0.99964 | 0.00016623
769 | 0.99963 | 0.00016911
769.2 | 0.99962 | 0.00017224
769.4 | 0.99961 | 0.0001756
769.6 | 0.9996 | 0.00017918
769.8 | 0.99959 | 0.00018293
770 | 0.99958 | 0.00018684
770.2 | 0.99957 | 0.00019085
770.4 | 0.99956 | 0.00019493
770.6 | 0.99954 | 0.00019901
770.8 | 0.99953 | 0.00020303
771 | 0.99952 | 0.00020693
771.2 | 0.99951 | 0.00021064
771.4 | 0.9995 | 0.00021407
771.6 | 0.99949 | 0.00021715
771.8 | 0.99948 | 0.0002198
772 | 0.99947 | 0.00022195
772.2 | 0.99946 | 0.00022353
772.4 | 0.99946 | 0.00022448
772.6 | 0.99946 | 0.00022477
772.8 | 0.99946 | 0.00022437
773 | 0.99946 | 0.00022326
773.2 | 0.99946 | 0.00022147
773.4 | 0.99946 | 0.00021901
773.6 | 0.99947 | 0.00021594
773.8 | 0.99948 | 0.0002123
774 | 0.99949 | 0.00020816
774.2 | 0.9995 | 0.0002036
774.4 | 0.99951 | 0.0001987
774.6 | 0.99952 | 0.00019354
774.8 | 0.99953 | 0.00018821
775 | 0.99955 | 0.00018276
775.2 | 0.99956 | 0.00017727
775.4 | 0.99958 | 0.00017181
775.6 | 0.99959 | 0.00016642
775.8 | 0.9996 | 0.00016115
776 | 0.99962 | 0.00015603
776.2 | 0.99963 | 0.0001511
776.4 | 0.99964 | 0.00014637
776.6 | 0.99965 | 0.00014186
776.8 | 0.99967 | 0.00013758
777 | 0.99968 | 0.00013354
777.2 | 0.99969 | 0.00012973
777.4 | 0.9997 | 0.00012617
777.6 | 0.99971 | 0.00012283
777.8 | 0.99971 | 0.00011973
778 | 0.99972 | 0.00011686
778.2 | 0.99973 | 0.0001142
778.4 | 0.99974 | 0.00011175
778.6 | 0.99974 | 0.0001095
778.8 | 0.99975 | 0.00010744
779 | 0.99976 | 0.00010558
779.2 | 0.99976 | 0.00010389
779.4 | 0.99977 | 0.00010237
779.6 | 0.99977 | 0.00010102
779.8 | 0.99977 | 9.9823e-05
780 | 0.99978 | 9.8778e-05
780.2 | 0.99978 | 9.7877e-05
780.4 | 0.99978 | 9.7115e-05
780.6 | 0.99979 | 9.6485e-05
780.8 | 0.99979 | 9.5983e-05
781 | 0.99979 | 9.5603e-05
781.2 | 0.99979 | 9.534e-05
781.4 | 0.9998 | 9.5191e-05
781.6 | 0.9998 | 9.5151e-05
781.8 | 0.9998 | 9.5216e-05
782 | 0.9998 | 9.5381e-05
782.2 | 0.9998 | 9.5644e-05
782.4 | 0.9998 | 9.6001e-05
782.6 | 0.9998 | 9.6448e-05
782.8 | 0.9998 | 9.6982e-05
783 | 0.9998 | 9.76e-05
783.2 | 0.9998 | 9.8299e-05
783.4 | 0.9998 | 9.9077e-05
783.6 | 0.9998 | 9.993e-05
783.8 | 0.9998 | 0.00010086
784 | 0.9998 | 0.00010185
784.2 | 0.9998 | 0.00010292
784.4 | 0.9998 | 0.00010405
784.6 | 0.9998 | 0.00010525
784.8 | 0.99979 | 0.00010651
785 | 0.99979 | 0.00010784
785.2 | 0.99979 | 0.00010924
785.4 | 0.99979 | 0.0001107
785.6 | 0.99979 | 0.00011222
785.8 | 0.99978 | 0.00011382
786 | 0.99978 | 0.00011548
786.2 | 0.99978 | 0.00011723
786.4 | 0.99977 | 0.00011905
786.6 | 0.99977 | 0.00012096
786.8 | 0.99977 | 0.00012296
787 | 0.99976 | 0.00012506
787.2 | 0.99976 | 0.00012727
787.4 | 0.99976 | 0.00012959
787.6 | 0.99975 | 0.00013205
787.8 | 0.99975 | 0.00013464
788 | 0.99974 | 0.00013739
788.2 | 0.99973 | 0.0001403
788.4 | 0.99973 | 0.0001434
788.6 | 0.99972 | 0.00014669
788.8 | 0.99971 | 0.0001502
789 | 0.99971 | 0.00015393
789.2 | 0.9997 | 0.00015792
789.4 | 0.99969 | 0.00016216
789.6 | 0.99968 | 0.00016669
789.8 | 0.99967 | 0.00017152
790 | 0.99966 | 0.00017666
790.2 | 0.99965 | 0.00018212
790.4 | 0.99963 | 0.00018791
790.6 | 0.99962 | 0.00019404
790.8 | 0.99961 | 0.00020049
791 | 0.99959 | 0.00020726
791.2 | 0.99958 | 0.00021432
791.4 | 0.99956 | 0.00022163
791.6 | 0.99954 | 0.00022914
791.8 | 0.99953 | 0.00023678
792 | 0.99951 | 0.00024445
792.2 | 0.99949 | 0.00025205
792.4 | 0.99948 | 0.00025943
792.6 | 0.99946 | 0.00026645
792.8 | 0.99944 | 0.00027294
793 | 0.99943 | 0.00027872
793.2 | 0.99942 | 0.00028361
793.4 | 0.99941 | 0.00028745
793.6 | 0.9994 | 0.00029008
793.8 | 0.99939 | 0.00029139
794 | 0.99939 | 0.00029131
794.2 | 0.99939 | 0.00028979
794.4 | 0.9994 | 0.00028684
794.6 | 0.9994 | 0.00028254
794.8 | 0.99941 | 0.00027698
795 | 0.99943 | 0.0002703
795.2 | 0.99944 | 0.00026266
795.4 | 0.99946 | 0.00025426
795.6 | 0.99947 | 0.00024529
795.8 | 0.99949 | 0.00023594
796 | 0.99951 | 0.00022638
796.2 | 0.99953 | 0.00021678
796.4 | 0.99955 | 0.00020728
796.6 | 0.99957 | 0.000198
796.8 | 0.99959 | 0.00018903
797 | 0.99961 | 0.00018044
797.2 | 0.99963 | 0.00017229
797.4 | 0.99965 | 0.00016461
797.6 | 0.99966 | 0.00015742
797.8 | 0.99968 | 0.00015071
798 | 0.99969 | 0.00014449
798.2 | 0.9997 | 0.00013876
798.4 | 0.99972 | 0.00013348
798.6 | 0.99973 | 0.00012865
798.8 | 0.99974 | 0.00012424
799 | 0.99975 | 0.00012024
799.2 | 0.99976 | 0.00011662
799.4 | 0.99976 | 0.00011336
799.6 | 0.99977 | 0.00011045
799.8 | 0.99978 | 0.00010787
800 | 0.99978 | 0.0001056
800.2 | 0.99979 | 0.00010362
800.4 | 0.99979 | 0.00010192
800.6 | 0.9998 | 0.0001005
800.8 | 0.9998 | 9.9337e-05
801 | 0.9998 | 9.8427e-05
801.2 | 0.99981 | 9.7764e-05
801.4 | 0.99981 | 9.7342e-05
801.6 | 0.99981 | 9.7159e-05
801.8 | 0.99981 | 9.7213e-05
802 | 0.99981 | 9.7503e-05
802.2 | 0.99981 | 9.8032e-05
802.4 | 0.99981 | 9.8803e-05
802.6 | 0.99981 | 9.982e-05
802.8 | 0.99981 | 0.00010109
803 | 0.99981 | 0.00010262
803.2 | 0.99981 | 0.00010443
803.4 | 0.99981 | 0.00010652
803.6 | 0.9998 | 0.00010891
803.8 | 0.9998 | 0.00011162
804 | 0.9998 | 0.00011467
804.2 | 0.99979 | 0.00011808
804.4 | 0.99979 | 0.00012187
804.6 | 0.99979 | 0.00012607
804.8 | 0.99978 | 0.00013072
805 | 0.99977 | 0.00013585
805.2 | 0.99977 | 0.0001415
805.4 | 0.99976 | 0.00014772
805.6 | 0.99975 | 0.00015455
805.8 | 0.99974 | 0.00016204
806 | 0.99973 | 0.00017026
806.2 | 0.99972 | 0.00017927
806.4 | 0.99971 | 0.00018915
806.6 | 0.9997 | 0.00019996
806.8 | 0.99968 | 0.00021179
807 | 0.99967 | 0.00022474
807.2 | 0.99965 | 0.00023889
807.4 | 0.99963 | 0.00025435
807.6 | 0.99961 | 0.00027122
807.8 | 0.99959 | 0.00028961
808 | 0.99957 | 0.00030963
808.2 | 0.99954 | 0.00033139
808.4 | 0.99951 | 0.00035498
808.6 | 0.99948 | 0.0003805
808.8 | 0.99945 | 0.00040801
809 | 0.99941 | 0.00043757
809.2 | 0.99938 | 0.00046916
809.4 | 0.99934 | 0.00050276
809.6 | 0.99929 | 0.00053826
809.8 | 0.99925 | 0.00057547
810 | 0.99921 | 0.00061413
810.2 | 0.99916 | 0.00065386
810.4 | 0.99911 | 0.00069417
810.6 | 0.99906 | 0.00073445
810.8 | 0.99902 | 0.00077399
811 | 0.99897 | 0.00081196
811.2 | 0.99893 | 0.00084747
811.4 | 0.99889 | 0.00087959
811.6 | 0.99886 | 0.00090741
811.8 | 0.99884 | 0.00093008
812 | 0.99882 | 0.00094689
812.2 | 0.9988 | 0.00095727
812.4 | 0.9988 | 0.00096091
812.6 | 0.9988 | 0.00095769
812.8 | 0.99881 | 0.00094776
813 | 0.99883 | 0.00093149
813.2 | 0.99886 | 0.00090946
813.4 | 0.99889 | 0.0008824
813.6 | 0.99893 | 0.00085115
813.8 | 0.99897 | 0.00081662
814 | 0.99901 | 0.00077971
814.2 | 0.99906 | 0.00074132
814.4 | 0.9991 | 0.00070224
814.6 | 0.99915 | 0.00066318
814.8 | 0.9992 | 0.00062472
815 | 0.99924 | 0.00058736
815.2 | 0.99928 | 0.00055145
815.4 | 0.99932 | 0.00051725
815.6 | 0.99936 | 0.00048493
815.8 | 0.9994 | 0.0004546
816 | 0.99943 | 0.00042627
816.2 | 0.99947 | 0.00039994
816.4 | 0.99949 | 0.00037557
816.6 | 0.99952 | 0.00035306
816.8 | 0.99955 | 0.00033233
817 | 0.99957 | 0.00031329
817.2 | 0.99959 | 0.00029582
817.4 | 0.99961 | 0.00027981
817.6 | 0.99963 | 0.00026517
817.8 | 0.99964 | 0.00025178
818 | 0.99966 | 0.00023956
818.2 | 0.99967 | 0.00022841
818.4 | 0.99968 | 0.00021825
818.6 | 0.99969 | 0.00020899
818.8 | 0.9997 | 0.00020057
819 | 0.99971 | 0.00019293
819.2 | 0.99972 | 0.000186
819.4 | 0.99973 | 0.00017972
819.6 | 0.99973 | 0.00017406
819.8 | 0.99974 | 0.00016896
820 | 0.99974 | 0.00016438
820.2 | 0.99975 | 0.0001603
820.4 | 0.99975 | 0.00015668
820.6 | 0.99975 | 0.00015349
820.8 | 0.99976 | 0.0001507
821 | 0.99976 | 0.00014831
821.2 | 0.99976 | 0.00014628
821.4 | 0.99976 | 0.00014459
821.6 | 0.99976 | 0.00014325
821.8 | 0.99976 | 0.00014223
822 | 0.99976 | 0.00014152
822.2 | 0.99976 | 0.00014112
822.4 | 0.99976 | 0.00014101
822.6 | 0.99976 | 0.0001412
822.8 | 0.99975 | 0.00014168
823 | 0.99975 | 0.00014245
823.2 | 0.99975 | 0.0001435
823.4 | 0.99974 | 0.00014484
823.6 | 0.99974 | 0.00014645
823.8 | 0.99973 | 0.00014836
824 | 0.99973 | 0.00015054
824.2 | 0.99972 | 0.00015301
824.4 | 0.99972 | 0.00015577
824.6 | 0.99971 | 0.00015881
824.8 | 0.9997 | 0.00016213
825 | 0.9997 | 0.00016573
825.2 | 0.99969 | 0.0001696
825.4 | 0.99968 | 0.00017375
825.6 | 0.99967 | 0.00017815
825.8 | 0.99966 | 0.00018279
826 | 0.99965 | 0.00018767
826.2 | 0.99964 | 0.00019276
826.4 | 0.99964 | 0.00019803
826.6 | 0.99963 | 0.00020345
826.8 | 0.99962 | 0.000209
827 | 0.99961 | 0.00021463
827.2 | 0.9996 | 0.00022031
827.4 | 0.99959 | 0.00022598
827.6 | 0.99958 | 0.00023161
827.8 | 0.99957 | 0.00023713
828 | 0.99956 | 0.00024251
828.2 | 0.99955 | 0.00024769
828.4 | 0.99955 | 0.00025263
828.6 | 0.99954 | 0.00025728
828.8 | 0.99954 | 0.00026161
829 | 0.99953 | 0.00026559
829.2 | 0.99953 | 0.0002692
829.4 | 0.99953 | 0.00027242
829.6 | 0.99952 | 0.00027526
829.8 | 0.99952 | 0.00027773
830 | 0.99952 | 0.00027983
830.2 | 0.99952 | 0.00028159
830.4 | 0.99953 | 0.00028305
830.6 | 0.99953 | 0.00028424
830.8 | 0.99953 | 0.00028521
831 | 0.99953 | 0.000286
831.2 | 0.99954 | 0.00028666
831.4 | 0.99954 | 0.00028725
831.6 | 0.99955 | 0.00028781
831.8 | 0.99955 | 0.0002884
832 | 0.99955 | 0.00028906
832.2 | 0.99956 | 0.00028984
832.4 | 0.99956 | 0.00029078
832.6 | 0.99956 | 0.00029192
832.8 | 0.99957 | 0.00029329
833 | 0.99957 | 0.00029493
833.2 | 0.99957 | 0.00029686
833.4 | 0.99957 | 0.00029911
833.6 | 0.99957 | 0.00030171
833.8 | 0.99957 | 0.00030466
834 | 0.99957 | 0.00030798
834.2 | 0.99957 | 0.00031168
834.4 | 0.99957 | 0.00031576
834.6 | 0.99957 | 0.00032022
834.8 | 0.99957 | 0.00032505
835 | 0.99956 | 0.00033023
835.2 | 0.99956 | 0.00033575
835.4 | 0.99956 | 0.00034157
835.6 | 0.99955 | 0.00034765
835.8 | 0.99955 | 0.00035395
836 | 0.99954 | 0.0003604
836.2 | 0.99954 | 0.00036695
836.4 | 0.99953 | 0.0003735
836.6 | 0.99953 | 0.00037998
836.8 | 0.99952 | 0.00038629
837 | 0.99952 | 0.00039233
837.2 | 0.99951 | 0.00039801
837.4 | 0.99951 | 0.00040322
837.6 | 0.99951 | 0.00040786
837.8 | 0.9995 | 0.00041184
838 | 0.9995 | 0.00041508
838.2 | 0.9995 | 0.00041752
838.4 | 0.9995 | 0.00041911
838.6 | 0.9995 | 0.00041982
838.8 | 0.9995 | 0.00041963
839 | 0.9995 | 0.00041858
839.2 | 0.9995 | 0.00041668
839.4 | 0.9995 | 0.00041399
839.6 | 0.99951 | 0.00041058
839.8 | 0.99951 | 0.00040652
840 | 0.99951 | 0.00040192
840.2 | 0.99952 | 0.00039687
840.4 | 0.99952 | 0.00039147
840.6 | 0.99953 | 0.00038582
840.8 | 0.99953 | 0.00038002
841 | 0.99954 | 0.00037416
841.2 | 0.99954 | 0.00036832
841.4 | 0.99955 | 0.00036258
841.6 | 0.99955 | 0.00035701
841.8 | 0.99956 | 0.00035166
842 | 0.99956 | 0.00034658
842.2 | 0.99956 | 0.00034182
842.4 | 0.99957 | 0.0003374
842.6 | 0.99957 | 0.00033334
842.8 | 0.99957 | 0.00032967
843 | 0.99957 | 0.00032638
843.2 | 0.99957 | 0.00032349
843.4 | 0.99957 | 0.000321
843.6 | 0.99957 | 0.00031889
843.8 | 0.99957 | 0.00031714
844 | 0.99957 | 0.00031575
844.2 | 0.99957 | 0.00031469
844.4 | 0.99956 | 0.00031393
844.6 | 0.99956 | 0.00031343
844.8 | 0.99956 | 0.00031316
845 | 0.99955 | 0.00031306
845.2 | 0.99955 | 0.0003131
845.4 | 0.99955 | 0.00031321
845.6 | 0.99954 | 0.00031333
845.8 | 0.99954 | 0.00031339
846 | 0.99954 | 0.00031333
846.2 | 0.99953 | 0.00031306
846.4 | 0.99953 | 0.00031253
846.6 | 0.99953 | 0.00031164
846.8 | 0.99953 | 0.00031034
847 | 0.99952 | 0.00030856
847.2 | 0.99952 | 0.00030625
847.4 | 0.99953 | 0.00030336
847.6 | 0.99953 | 0.00029986
847.8 | 0.99953 | 0.00029574
848 | 0.99953 | 0.00029098
848.2 | 0.99954 | 0.00028562
848.4 | 0.99955 | 0.00027968
848.6 | 0.99955 | 0.00027319
848.8 | 0.99956 | 0.00026623
849 | 0.99957 | 0.00025886
849.2 | 0.99958 | 0.00025115
849.4 | 0.99959 | 0.00024318
849.6 | 0.9996 | 0.00023504
849.8 | 0.99961 | 0.00022679
850 | 0.99963 | 0.00021853
850.2 | 0.99964 | 0.00021031
850.4 | 0.99965 | 0.00020221
850.6 | 0.99966 | 0.00019428
850.8 | 0.99967 | 0.00018656
851 | 0.99969 | 0.00017911
851.2 | 0.9997 | 0.00017193
851.4 | 0.99971 | 0.00016507
851.6 | 0.99972 | 0.00015853
851.8 | 0.99973 | 0.00015233
852 | 0.99974 | 0.00014647
852.2 | 0.99975 | 0.00014094
852.4 | 0.99976 | 0.00013575
852.6 | 0.99976 | 0.00013088
852.8 | 0.99977 | 0.00012633
853 | 0.99978 | 0.00012209
853.2 | 0.99978 | 0.00011815
853.4 | 0.99979 | 0.00011449
853.6 | 0.9998 | 0.0001111
853.8 | 0.9998 | 0.00010797
854 | 0.99981 | 0.00010508
854.2 | 0.99981 | 0.00010243
854.4 | 0.99981 | 9.9995e-05
854.6 | 0.99982 | 9.7774e-05
854.8 | 0.99982 | 9.5754e-05
855 | 0.99982 | 9.3924e-05
855.2 | 0.99983 | 9.2275e-05
855.4 | 0.99983 | 9.0799e-05
855.6 | 0.99983 | 8.9488e-05
855.8 | 0.99983 | 8.8335e-05
856 | 0.99983 | 8.7333e-05
856.2 | 0.99984 | 8.6477e-05
856.4 | 0.99984 | 8.576e-05
856.6 | 0.99984 | 8.5178e-05
856.8 | 0.99984 | 8.4727e-05
857 | 0.99984 | 8.4402e-05
857.2 | 0.99984 | 8.4201e-05
857.4 | 0.99984 | 8.412e-05
857.6 | 0.99984 | 8.4157e-05
857.8 | 0.99984 | 8.4309e-05
858 | 0.99984 | 8.4574e-05
858.2 | 0.99984 | 8.4951e-05
858.4 | 0.99984 | 8.5437e-05
858.6 | 0.99984 | 8.6033e-05
858.8 | 0.99983 | 8.6735e-05
859 | 0.99983 | 8.7545e-05
859.2 | 0.99983 | 8.8459e-05
859.4 | 0.99983 | 8.9478e-05
859.6 | 0.99983 | 9.06e-05
859.8 | 0.99982 | 9.1825e-05
860 | 0.99982 | 9.315e-05
860.2 | 0.99982 | 9.4575e-05
860.4 | 0.99982 | 9.6097e-05
860.6 | 0.99981 | 9.7714e-05
860.8 | 0.99981 | 9.9423e-05
861 | 0.99981 | 0.00010122
861.2 | 0.9998 | 0.0001031
861.4 | 0.9998 | 0.00010507
861.6 | 0.99979 | 0.0001071
861.8 | 0.99979 | 0.00010921
862 | 0.99978 | 0.00011137
862.2 | 0.99978 | 0.00011358
862.4 | 0.99978 | 0.00011583
862.6 | 0.99977 | 0.00011811
862.8 | 0.99977 | 0.00012039
863 | 0.99976 | 0.00012268
863.2 | 0.99976 | 0.00012495
863.4 | 0.99975 | 0.00012717
863.6 | 0.99975 | 0.00012934
863.8 | 0.99974 | 0.00013143
864 | 0.99974 | 0.00013341
864.2 | 0.99973 | 0.00013527
864.4 | 0.99973 | 0.00013697
864.6 | 0.99972 | 0.00013849
864.8 | 0.99972 | 0.0001398
865 | 0.99972 | 0.00014089
865.2 | 0.99971 | 0.00014173
865.4 | 0.99971 | 0.0001423
865.6 | 0.99971 | 0.00014257
865.8 | 0.99971 | 0.00014254
866 | 0.9997 | 0.0001422
866.2 | 0.9997 | 0.00014153
866.4 | 0.9997 | 0.00014053
866.6 | 0.99971 | 0.00013922
866.8 | 0.99971 | 0.00013758
867 | 0.99971 | 0.00013564
867.2 | 0.99971 | 0.00013341
867.4 | 0.99971 | 0.00013092
867.6 | 0.99972 | 0.00012818
867.8 | 0.99972 | 0.00012523
868 | 0.99973 | 0.00012208
868.2 | 0.99973 | 0.00011878
868.4 | 0.99974 | 0.00011536
868.6 | 0.99974 | 0.00011183
868.8 | 0.99975 | 0.00010825
869 | 0.99975 | 0.00010462
869.2 | 0.99976 | 0.00010098
869.4 | 0.99977 | 9.7359e-05
869.6 | 0.99977 | 9.377e-05
869.8 | 0.99978 | 9.0236e-05
870 | 0.99979 | 8.6773e-05
870.2 | 0.99979 | 8.3396e-05
870.4 | 0.9998 | 8.0116e-05
870.6 | 0.9998 | 7.6942e-05
870.8 | 0.99981 | 7.3881e-05
871 | 0.99982 | 7.0939e-05
871.2 | 0.99982 | 6.8119e-05
871.4 | 0.99983 | 6.5424e-05
871.6 | 0.99983 | 6.2852e-05
871.8 | 0.99984 | 6.0404e-05
872 | 0.99984 | 5.8079e-05
872.2 | 0.99985 | 5.5873e-05
872.4 | 0.99985 | 5.3784e-05
872.6 | 0.99985 | 5.181e-05
872.8 | 0.99986 | 4.9945e-05
873 | 0.99986 | 4.8187e-05
873.2 | 0.99987 | 4.653e-05
873.4 | 0.99987 | 4.4972e-05
873.6 | 0.99987 | 4.3507e-05
873.8 | 0.99987 | 4.2132e-05
874 | 0.99988 | 4.0843e-05
874.2 | 0.99988 | 3.9636e-05
874.4 | 0.99988 | 3.8506e-05
874.6 | 0.99989 | 3.7451e-05
874.8 | 0.99989 | 3.6467e-05
875 | 0.99989 | 3.555e-05
875.2 | 0.99989 | 3.4698e-05
875.4 | 0.99989 | 3.3907e-05
875.6 | 0.99989 | 3.3176e-05
875.8 | 0.9999 | 3.25e-05
876 | 0.9999 | 3.1878e-05
876.2 | 0.9999 | 3.1308e-05
876.4 | 0.9999 | 3.0787e-05
876.6 | 0.9999 | 3.0314e-05
876.8 | 0.9999 | 2.9887e-05
877 | 0.9999 | 2.9505e-05
877.2 | 0.9999 | 2.9166e-05
877.4 | 0.9999 | 2.8869e-05
877.6 | 0.9999 | 2.8612e-05
877.8 | 0.9999 | 2.8396e-05
878 | 0.9999 | 2.8219e-05
878.2 | 0.9999 | 2.808e-05
878.4 | 0.9999 | 2.798e-05
878.6 | 0.9999 | 2.7917e-05
878.8 | 0.9999 | 2.7892e-05
879 | 0.9999 | 2.7903e-05
879.2 | 0.9999 | 2.7953e-05
879.4 | 0.9999 | 2.8039e-05
879.6 | 0.9999 | 2.8163e-05
879.8 | 0.9999 | 2.8326e-05
880 | 0.9999 | 2.8527e-05
880.2 | 0.9999 | 2.8767e-05
880.4 | 0.9999 | 2.9047e-05
880.6 | 0.9999 | 2.9367e-05
880.8 | 0.9999 | 2.973e-05
881 | 0.9999 | 3.0135e-05
881.2 | 0.9999 | 3.0584e-05
881.4 | 0.99989 | 3.1078e-05
881.6 | 0.99989 | 3.1619e-05
881.8 | 0.99989 | 3.2207e-05
882 | 0.99989 | 3.2845e-05
882.2 | 0.99989 | 3.3534e-05
882.4 | 0.99988 | 3.4275e-05
882.6 | 0.99988 | 3.507e-05
882.8 | 0.99988 | 3.5921e-05
883 | 0.99988 | 3.6828e-05
883.2 | 0.99987 | 3.7793e-05
883.4 | 0.99987 | 3.8817e-05
883.6 | 0.99987 | 3.9901e-05
883.8 | 0.99986 | 4.1046e-05
884 | 0.99986 | 4.2251e-05
884.2 | 0.99986 | 4.3517e-05
884.4 | 0.99985 | 4.4843e-05
884.6 | 0.99985 | 4.6226e-05
884.8 | 0.99985 | 4.7666e-05
885 | 0.99984 | 4.9158e-05
885.2 | 0.99984 | 5.0699e-05
885.4 | 0.99984 | 5.2283e-05
885.6 | 0.99983 | 5.3905e-05
885.8 | 0.99983 | 5.5557e-05
886 | 0.99982 | 5.723e-05
886.2 | 0.99982 | 5.8914e-05
886.4 | 0.99981 | 6.0598e-05
886.6 | 0.99981 | 6.227e-05
886.8 | 0.99981 | 6.3917e-05
887 | 0.9998 | 6.5524e-05
887.2 | 0.9998 | 6.7078e-05
887.4 | 0.9998 | 6.8564e-05
887.6 | 0.99979 | 6.9967e-05
887.8 | 0.99979 | 7.1273e-05
888 | 0.99979 | 7.247e-05
888.2 | 0.99979 | 7.3545e-05
888.4 | 0.99979 | 7.4489e-05
888.6 | 0.99979 | 7.5294e-05
888.8 | 0.99979 | 7.5954e-05
889 | 0.99979 | 7.6464e-05
889.2 | 0.99979 | 7.6824e-05
889.4 | 0.99979 | 7.7036e-05
889.6 | 0.99979 | 7.7102e-05
889.8 | 0.99979 | 7.7027e-05
890 | 0.99979 | 7.6821e-05
890.2 | 0.9998 | 7.6491e-05
890.4 | 0.9998 | 7.6047e-05
890.6 | 0.9998 | 7.5501e-05
890.8 | 0.99981 | 7.4864e-05
891 | 0.99981 | 7.4149e-05
891.2 | 0.99981 | 7.3366e-05
891.4 | 0.99982 | 7.2528e-05
891.6 | 0.99982 | 7.1645e-05
891.8 | 0.99982 | 7.0728e-05
892 | 0.99983 | 6.9786e-05
892.2 | 0.99983 | 6.8827e-05
892.4 | 0.99983 | 6.7861e-05
892.6 | 0.99984 | 6.6892e-05
892.8 | 0.99984 | 6.5928e-05
893 | 0.99984 | 6.4973e-05
893.2 | 0.99985 | 6.4031e-05
893.4 | 0.99985 | 6.3106e-05
893.6 | 0.99985 | 6.2202e-05
893.8 | 0.99986 | 6.1319e-05
894 | 0.99986 | 6.046e-05
894.2 | 0.99986 | 5.9626e-05
894.4 | 0.99986 | 5.8818e-05
894.6 | 0.99987 | 5.8035e-05
894.8 | 0.99987 | 5.728e-05
895 | 0.99987 | 5.655e-05
895.2 | 0.99987 | 5.5847e-05
895.4 | 0.99987 | 5.5169e-05
895.6 | 0.99988 | 5.4517e-05
895.8 | 0.99988 | 5.3888e-05
896 | 0.99988 | 5.3284e-05
896.2 | 0.99988 | 5.2703e-05
896.4 | 0.99988 | 5.2144e-05
896.6 | 0.99988 | 5.1606e-05
896.8 | 0.99989 | 5.109e-05
897 | 0.99989 | 5.0594e-05
897.2 | 0.99989 | 5.0118e-05
897.4 | 0.99989 | 4.9661e-05
897.6 | 0.99989 | 4.9223e-05
897.8 | 0.99989 | 4.8803e-05
898 | 0.99989 | 4.8402e-05
898.2 | 0.99989 | 4.8018e-05
898.4 | 0.99989 | 4.7652e-05
898.6 | 0.99989 | 4.7304e-05
898.8 | 0.99989 | 4.6974e-05
899 | 0.99989 | 4.6661e-05
899.2 | 0.9999 | 4.6366e-05
899.4 | 0.9999 | 4.609e-05
899.6 | 0.9999 | 4.5831e-05
899.8 | 0.9999 | 4.5592e-05
900 | 0.9999 | 4.5371e-05
900.2 | 0.9999 | 4.5171e-05
900.4 | 0.9999 | 4.499e-05
900.6 | 0.9999 | 4.4829e-05
900.8 | 0.9999 | 4.469e-05
901 | 0.9999 | 4.4573e-05
901.2 | 0.99989 | 4.4477e-05
901.4 | 0.99989 | 4.4404e-05
901.6 | 0.99989 | 4.4355e-05
901.8 | 0.99989 | 4.4329e-05
902 | 0.99989 | 4.4328e-05
902.2 | 0.99989 | 4.4352e-05
902.4 | 0.99989 | 4.4402e-05
902.6 | 0.99989 | 4.4477e-05
902.8 | 0.99989 | 4.4579e-05
903 | 0.99989 | 4.4707e-05
903.2 | 0.99989 | 4.4863e-05
903.4 | 0.99989 | 4.5046e-05
903.6 | 0.99988 | 4.5256e-05
903.8 | 0.99988 | 4.5494e-05
904 | 0.99988 | 4.5759e-05
904.2 | 0.99988 | 4.6052e-05
904.4 | 0.99988 | 4.637e-05
904.6 | 0.99988 | 4.6715e-05
904.8 | 0.99987 | 4.7084e-05
905 | 0.99987 | 4.7477e-05
905.2 | 0.99987 | 4.7892e-05
905.4 | 0.99987 | 4.8327e-05
905.6 | 0.99986 | 4.8781e-05
905.8 | 0.99986 | 4.9249e-05
906 | 0.99986 | 4.973e-05
906.2 | 0.99986 | 5.0219e-05
906.4 | 0.99985 | 5.0713e-05
906.6 | 0.99985 | 5.1207e-05
906.8 | 0.99985 | 5.1696e-05
907 | 0.99985 | 5.2174e-05
907.2 | 0.99984 | 5.2636e-05
907.4 | 0.99984 | 5.3074e-05
907.6 | 0.99984 | 5.3483e-05
907.8 | 0.99984 | 5.3855e-05
908 | 0.99983 | 5.4183e-05
908.2 | 0.99983 | 5.4459e-05
908.4 | 0.99983 | 5.4677e-05
908.6 | 0.99983 | 5.483e-05
908.8 | 0.99983 | 5.4912e-05
909 | 0.99982 | 5.4916e-05
909.2 | 0.99982 | 5.4838e-05
909.4 | 0.99982 | 5.4675e-05
909.6 | 0.99982 | 5.4423e-05
909.8 | 0.99982 | 5.4081e-05
910 | 0.99982 | 5.3649e-05
910.2 | 0.99982 | 5.3129e-05
910.4 | 0.99982 | 5.2524e-05
910.6 | 0.99983 | 5.1836e-05
910.8 | 0.99983 | 5.1071e-05
911 | 0.99983 | 5.0235e-05
911.2 | 0.99983 | 4.9336e-05
911.4 | 0.99983 | 4.8381e-05
911.6 | 0.99984 | 4.7378e-05
911.8 | 0.99984 | 4.6335e-05
912 | 0.99984 | 4.5262e-05
912.2 | 0.99984 | 4.4166e-05
912.4 | 0.99985 | 4.3055e-05
912.6 | 0.99985 | 4.1938e-05
912.8 | 0.99985 | 4.0821e-05
913 | 0.99986 | 3.9712e-05
913.2 | 0.99986 | 3.8615e-05
913.4 | 0.99986 | 3.7537e-05
913.6 | 0.99987 | 3.6481e-05
913.8 | 0.99987 | 3.5452e-05
914 | 0.99987 | 3.4453e-05
914.2 | 0.99988 | 3.3486e-05
914.4 | 0.99988 | 3.2554e-05
914.6 | 0.99988 | 3.1659e-05
914.8 | 0.99989 | 3.08e-05
915 | 0.99989 | 2.9979e-05
915.2 | 0.99989 | 2.9197e-05
915.4 | 0.99989 | 2.8452e-05
915.6 | 0.9999 | 2.7746e-05
915.8 | 0.9999 | 2.7078e-05
916 | 0.9999 | 2.6446e-05
916.2 | 0.9999 | 2.5852e-05
916.4 | 0.99991 | 2.5293e-05
916.6 | 0.99991 | 2.4768e-05
916.8 | 0.99991 | 2.4278e-05
917 | 0.99991 | 2.3821e-05
917.2 | 0.99991 | 2.3397e-05
917.4 | 0.99991 | 2.3003e-05
917.6 | 0.99992 | 2.264e-05
917.8 | 0.99992 | 2.2307e-05
918 | 0.99992 | 2.2003e-05
918.2 | 0.99992 | 2.1727e-05
918.4 | 0.99992 | 2.1478e-05
918.6 | 0.99992 | 2.1256e-05
918.8 | 0.99992 | 2.106e-05
919 | 0.99992 | 2.089e-05
919.2 | 0.99992 | 2.0745e-05
919.4 | 0.99992 | 2.0625e-05
919.6 | 0.99992 | 2.053e-05
919.8 | 0.99992 | 2.0458e-05
920 | 0.99992 | 2.0411e-05
920.2 | 0.99993 | 2.0387e-05
920.4 | 0.99993 | 2.0388e-05
920.6 | 0.99993 | 2.0412e-05
920.8 | 0.99993 | 2.046e-05
921 | 0.99993 | 2.0531e-05
921.2 | 0.99993 | 2.0627e-05
921.4 | 0.99993 | 2.0748e-05
921.6 | 0.99993 | 2.0893e-05
921.8 | 0.99992 | 2.1064e-05
922 | 0.99992 | 2.126e-05
922.2 | 0.99992 | 2.1483e-05
922.4 | 0.99992 | 2.1732e-05
922.6 | 0.99992 | 2.2009e-05
922.8 | 0.99992 | 2.2315e-05
923 | 0.99992 | 2.265e-05
923.2 | 0.99992 | 2.3015e-05
923.4 | 0.99992 | 2.3411e-05
923.6 | 0.99992 | 2.3839e-05
923.8 | 0.99992 | 2.4301e-05
924 | 0.99992 | 2.4798e-05
924.2 | 0.99992 | 2.533e-05
924.4 | 0.99991 | 2.5899e-05
924.6 | 0.99991 | 2.6507e-05
924.8 | 0.99991 | 2.7155e-05
925 | 0.99991 | 2.7845e-05
925.2 | 0.99991 | 2.8577e-05
925.4 | 0.99991 | 2.9354e-05
925.6 | 0.9999 | 3.0176e-05
925.8 | 0.9999 | 3.1045e-05
926 | 0.9999 | 3.1963e-05
926.2 | 0.9999 | 3.293e-05
926.4 | 0.99989 | 3.3948e-05
926.6 | 0.99989 | 3.5018e-05
926.8 | 0.99989 | 3.6139e-05
927 | 0.99989 | 3.7313e-05
927.2 | 0.99988 | 3.8538e-05
927.4 | 0.99988 | 3.9816e-05
927.6 | 0.99988 | 4.1145e-05
927.8 | 0.99988 | 4.2523e-05
928 | 0.99987 | 4.3949e-05
928.2 | 0.99987 | 4.542e-05
928.4 | 0.99987 | 4.6932e-05
928.6 | 0.99986 | 4.8481e-05
928.8 | 0.99986 | 5.0063e-05
929 | 0.99986 | 5.1672e-05
929.2 | 0.99985 | 5.3302e-05
929.4 | 0.99985 | 5.4945e-05
929.6 | 0.99985 | 5.6593e-05
929.8 | 0.99984 | 5.8238e-05
930 | 0.99984 | 5.9872e-05
930.2 | 0.99984 | 6.1484e-05
930.4 | 0.99984 | 6.3066e-05
930.6 | 0.99983 | 6.4607e-05
930.8 | 0.99983 | 6.6099e-05
931 | 0.99983 | 6.7531e-05
931.2 | 0.99983 | 6.8897e-05
931.4 | 0.99983 | 7.0187e-05
931.6 | 0.99983 | 7.1396e-05
931.8 | 0.99982 | 7.2519e-05
932 | 0.99982 | 7.355e-05
932.2 | 0.99982 | 7.4487e-05
932.4 | 0.99982 | 7.5329e-05
932.6 | 0.99982 | 7.6075e-05
932.8 | 0.99982 | 7.6727e-05
933 | 0.99982 | 7.7287e-05
933.2 | 0.99983 | 7.7761e-05
933.4 | 0.99983 | 7.8151e-05
933.6 | 0.99983 | 7.8465e-05
933.8 | 0.99983 | 7.8709e-05
934 | 0.99983 | 7.8891e-05
934.2 | 0.99983 | 7.9018e-05
934.4 | 0.99983 | 7.9098e-05
934.6 | 0.99983 | 7.9139e-05
934.8 | 0.99984 | 7.9149e-05
935 | 0.99984 | 7.9137e-05
935.2 | 0.99984 | 7.911e-05
935.4 | 0.99984 | 7.9075e-05
935.6 | 0.99984 | 7.904e-05
935.8 | 0.99985 | 7.9012e-05
936 | 0.99985 | 7.8997e-05
936.2 | 0.99985 | 7.9e-05
936.4 | 0.99985 | 7.9028e-05
936.6 | 0.99985 | 7.9086e-05
936.8 | 0.99985 | 7.9179e-05
937 | 0.99985 | 7.931e-05
937.2 | 0.99986 | 7.9485e-05
937.4 | 0.99986 | 7.9707e-05
937.6 | 0.99986 | 7.998e-05
937.8 | 0.99986 | 8.0307e-05
938 | 0.99986 | 8.0692e-05
938.2 | 0.99986 | 8.1138e-05
938.4 | 0.99986 | 8.1647e-05
938.6 | 0.99986 | 8.2224e-05
938.8 | 0.99986 | 8.287e-05
939 | 0.99986 | 8.3589e-05
939.2 | 0.99986 | 8.4384e-05
939.4 | 0.99986 | 8.5257e-05
939.6 | 0.99986 | 8.6211e-05
939.8 | 0.99986 | 8.725e-05
940 | 0.99986 | 8.8377e-05
940.2 | 0.99986 | 8.9594e-05
940.4 | 0.99986 | 9.0906e-05
940.6 | 0.99986 | 9.2315e-05
940.8 | 0.99985 | 9.3825e-05
941 | 0.99985 | 9.544e-05
941.2 | 0.99985 | 9.7164e-05
941.4 | 0.99985 | 9.9002e-05
941.6 | 0.99985 | 0.00010096
941.8 | 0.99985 | 0.00010303
942 | 0.99984 | 0.00010524
942.2 | 0.99984 | 0.00010757
942.4 | 0.99984 | 0.00011004
942.6 | 0.99984 | 0.00011266
942.8 | 0.99983 | 0.00011542
943 | 0.99983 | 0.00011834
943.2 | 0.99983 | 0.00012142
943.4 | 0.99982 | 0.00012466
943.6 | 0.99982 | 0.00012808
943.8 | 0.99982 | 0.00013168
944 | 0.99981 | 0.00013546
944.2 | 0.99981 | 0.00013944
944.4 | 0.9998 | 0.00014362
944.6 | 0.9998 | 0.00014801
944.8 | 0.99979 | 0.00015261
945 | 0.99979 | 0.00015744
945.2 | 0.99978 | 0.00016249
945.4 | 0.99978 | 0.00016779
945.6 | 0.99977 | 0.00017333
945.8 | 0.99976 | 0.00017912
946 | 0.99976 | 0.00018516
946.2 | 0.99975 | 0.00019147
946.4 | 0.99974 | 0.00019804
946.6 | 0.99973 | 0.00020489
946.8 | 0.99973 | 0.000212
947 | 0.99972 | 0.00021939
947.2 | 0.99971 | 0.00022705
947.4 | 0.9997 | 0.00023497
947.6 | 0.99969 | 0.00024316
947.8 | 0.99968 | 0.0002516
948 | 0.99967 | 0.00026028
948.2 | 0.99966 | 0.00026919
948.4 | 0.99965 | 0.00027831
948.6 | 0.99964 | 0.00028761
948.8 | 0.99963 | 0.00029707
949 | 0.99962 | 0.00030666
949.2 | 0.99961 | 0.00031632
949.4 | 0.99959 | 0.00032603
949.6 | 0.99958 | 0.00033573
949.8 | 0.99957 | 0.00034538
950 | 0.99956 | 0.0003549
950.2 | 0.99955 | 0.00036424
950.4 | 0.99954 | 0.00037333
950.6 | 0.99953 | 0.00038211
950.8 | 0.99952 | 0.0003905
951 | 0.99951 | 0.00039843
951.2 | 0.9995 | 0.00040583
951.4 | 0.9995 | 0.00041263
951.6 | 0.99949 | 0.00041878
951.8 | 0.99948 | 0.00042421
952 | 0.99948 | 0.00042888
952.2 | 0.99947 | 0.00043273
952.4 | 0.99947 | 0.00043576
952.6 | 0.99947 | 0.00043792
952.8 | 0.99947 | 0.00043921
953 | 0.99947 | 0.00043964
953.2 | 0.99947 | 0.00043922
953.4 | 0.99947 | 0.00043797
953.6 | 0.99947 | 0.00043593
953.8 | 0.99948 | 0.00043314
954 | 0.99948 | 0.00042965
954.2 | 0.99949 | 0.00042553
954.4 | 0.99949 | 0.00042083
954.6 | 0.9995 | 0.00041562
954.8 | 0.99951 | 0.00040997
955 | 0.99952 | 0.00040395
955.2 | 0.99952 | 0.00039762
955.4 | 0.99953 | 0.00039106
955.6 | 0.99954 | 0.00038433
955.8 | 0.99955 | 0.00037748
956 | 0.99956 | 0.00037058
956.2 | 0.99957 | 0.00036367
956.4 | 0.99957 | 0.00035679
956.6 | 0.99958 | 0.00035
956.8 | 0.99959 | 0.00034333
957 | 0.9996 | 0.0003368
957.2 | 0.99961 | 0.00033046
957.4 | 0.99962 | 0.00032431
957.6 | 0.99962 | 0.00031839
957.8 | 0.99963 | 0.0003127
958 | 0.99964 | 0.00030727
958.2 | 0.99964 | 0.0003021
958.4 | 0.99965 | 0.0002972
958.6 | 0.99965 | 0.00029258
958.8 | 0.99966 | 0.00028824
959 | 0.99966 | 0.00028419
959.2 | 0.99967 | 0.00028043
959.4 | 0.99967 | 0.00027695
959.6 | 0.99968 | 0.00027377
959.8 | 0.99968 | 0.00027088
960 | 0.99968 | 0.00026827
960.2 | 0.99969 | 0.00026596
960.4 | 0.99969 | 0.00026394
960.6 | 0.99969 | 0.0002622
960.8 | 0.99969 | 0.00026075
961 | 0.99969 | 0.00025958
961.2 | 0.9997 | 0.0002587
961.4 | 0.9997 | 0.00025811
961.6 | 0.9997 | 0.0002578
961.8 | 0.9997 | 0.00025778
962 | 0.9997 | 0.00025805
962.2 | 0.9997 | 0.0002586
962.4 | 0.9997 | 0.00025944
962.6 | 0.99969 | 0.00026057
962.8 | 0.99969 | 0.00026198
963 | 0.99969 | 0.00026369
963.2 | 0.99969 | 0.00026569
963.4 | 0.99968 | 0.00026798
963.6 | 0.99968 | 0.00027056
963.8 | 0.99968 | 0.00027344
964 | 0.99967 | 0.0002766
964.2 | 0.99967 | 0.00028005
964.4 | 0.99967 | 0.00028378
964.6 | 0.99966 | 0.00028779
964.8 | 0.99966 | 0.00029206
965 | 0.99965 | 0.00029659
965.2 | 0.99964 | 0.00030136
965.4 | 0.99964 | 0.00030636
965.6 | 0.99963 | 0.00031156
965.8 | 0.99962 | 0.00031693
966 | 0.99962 | 0.00032244
966.2 | 0.99961 | 0.00032805
966.4 | 0.9996 | 0.00033372
966.6 | 0.99959 | 0.00033941
966.8 | 0.99959 | 0.00034504
967 | 0.99958 | 0.00035057
967.2 | 0.99957 | 0.00035591
967.4 | 0.99957 | 0.00036101
967.6 | 0.99956 | 0.00036578
967.8 | 0.99955 | 0.00037014
968 | 0.99955 | 0.00037401
968.2 | 0.99954 | 0.00037732
968.4 | 0.99954 | 0.00037999
968.6 | 0.99954 | 0.00038194
968.8 | 0.99953 | 0.00038313
969 | 0.99953 | 0.00038349
969.2 | 0.99953 | 0.00038299
969.4 | 0.99953 | 0.0003816
969.6 | 0.99953 | 0.00037931
969.8 | 0.99954 | 0.00037613
970 | 0.99954 | 0.00037208
970.2 | 0.99955 | 0.00036718
970.4 | 0.99955 | 0.0003615
970.6 | 0.99956 | 0.0003551
970.8 | 0.99957 | 0.00034804
971 | 0.99958 | 0.0003404
971.2 | 0.99958 | 0.00033227
971.4 | 0.99959 | 0.00032375
971.6 | 0.9996 | 0.00031491
971.8 | 0.99961 | 0.00030585
972 | 0.99963 | 0.00029666
972.2 | 0.99964 | 0.0002874
972.4 | 0.99965 | 0.00027815
972.6 | 0.99966 | 0.00026899
972.8 | 0.99967 | 0.00025995
973 | 0.99968 | 0.0002511
973.2 | 0.99969 | 0.00024247
973.4 | 0.9997 | 0.0002341
973.6 | 0.99971 | 0.00022601
973.8 | 0.99972 | 0.00021823
974 | 0.99973 | 0.00021075
974.2 | 0.99974 | 0.00020361
974.4 | 0.99974 | 0.00019678
974.6 | 0.99975 | 0.00019029
974.8 | 0.99976 | 0.00018412
975 | 0.99977 | 0.00017827
975.2 | 0.99977 | 0.00017274
975.4 | 0.99978 | 0.00016751
975.6 | 0.99979 | 0.00016258
975.8 | 0.99979 | 0.00015793
976 | 0.9998 | 0.00015356
976.2 | 0.9998 | 0.00014945
976.4 | 0.99981 | 0.0001456
976.6 | 0.99981 | 0.000142
976.8 | 0.99982 | 0.00013863
977 | 0.99982 | 0.00013548
977.2 | 0.99982 | 0.00013254
977.4 | 0.99983 | 0.00012981
977.6 | 0.99983 | 0.00012728
977.8 | 0.99983 | 0.00012494
978 | 0.99984 | 0.00012277
978.2 | 0.99984 | 0.00012078
978.4 | 0.99984 | 0.00011896
978.6 | 0.99984 | 0.00011729
978.8 | 0.99985 | 0.00011579
979 | 0.99985 | 0.00011443
979.2 | 0.99985 | 0.00011322
979.4 | 0.99985 | 0.00011214
979.6 | 0.99985 | 0.00011121
979.8 | 0.99985 | 0.00011041
980 | 0.99985 | 0.00010974
980.2 | 0.99985 | 0.0001092
980.4 | 0.99985 | 0.00010879
980.6 | 0.99985 | 0.00010851
980.8 | 0.99985 | 0.00010835
981 | 0.99985 | 0.00010831
981.2 | 0.99985 | 0.0001084
981.4 | 0.99985 | 0.00010861
981.6 | 0.99985 | 0.00010894
981.8 | 0.99985 | 0.0001094
982 | 0.99985 | 0.00010999
982.2 | 0.99985 | 0.00011071
982.4 | 0.99985 | 0.00011155
982.6 | 0.99985 | 0.00011253
982.8 | 0.99985 | 0.00011365
983 | 0.99984 | 0.0001149
983.2 | 0.99984 | 0.00011629
983.4 | 0.99984 | 0.00011784
983.6 | 0.99984 | 0.00011953
983.8 | 0.99984 | 0.00012137
984 | 0.99983 | 0.00012337
984.2 | 0.99983 | 0.00012554
984.4 | 0.99983 | 0.00012787
984.6 | 0.99982 | 0.00013038
984.8 | 0.99982 | 0.00013306
985 | 0.99982 | 0.00013592
985.2 | 0.99981 | 0.00013897
985.4 | 0.99981 | 0.00014221
985.6 | 0.9998 | 0.00014564
985.8 | 0.9998 | 0.00014926
986 | 0.99979 | 0.00015308
986.2 | 0.99979 | 0.00015709
986.4 | 0.99978 | 0.0001613
986.6 | 0.99978 | 0.00016569
986.8 | 0.99977 | 0.00017025
987 | 0.99976 | 0.00017499
987.2 | 0.99976 | 0.00017987
987.4 | 0.99975 | 0.00018489
987.6 | 0.99974 | 0.00019001
987.8 | 0.99974 | 0.00019521
988 | 0.99973 | 0.00020046
988.2 | 0.99972 | 0.00020571
988.4 | 0.99972 | 0.00021091
988.6 | 0.99971 | 0.00021602
988.8 | 0.9997 | 0.00022098
989 | 0.9997 | 0.00022572
989.2 | 0.99969 | 0.00023018
989.4 | 0.99968 | 0.0002343
989.6 | 0.99968 | 0.000238
989.8 | 0.99967 | 0.00024123
990 | 0.99967 | 0.00024392
990.2 | 0.99967 | 0.00024602
990.4 | 0.99966 | 0.00024748
990.6 | 0.99966 | 0.00024827
990.8 | 0.99966 | 0.00024836
991 | 0.99966 | 0.00024775
991.2 | 0.99966 | 0.00024643
991.4 | 0.99967 | 0.00024442
991.6 | 0.99967 | 0.00024174
991.8 | 0.99968 | 0.00023844
992 | 0.99968 | 0.00023457
992.2 | 0.99969 | 0.00023019
992.4 | 0.99969 | 0.00022535
992.6 | 0.9997 | 0.00022013
992.8 | 0.99971 | 0.0002146
993 | 0.99971 | 0.00020882
993.2 | 0.99972 | 0.00020287
993.4 | 0.99973 | 0.00019681
993.6 | 0.99974 | 0.0001907
993.8 | 0.99975 | 0.00018459
994 | 0.99975 | 0.00017852
994.2 | 0.99976 | 0.00017255
994.4 | 0.99977 | 0.0001667
994.6 | 0.99978 | 0.00016101
994.8 | 0.99979 | 0.00015548
995 | 0.99979 | 0.00015015
995.2 | 0.9998 | 0.00014502
995.4 | 0.99981 | 0.00014011
995.6 | 0.99981 | 0.00013541
995.8 | 0.99982 | 0.00013093
996 | 0.99982 | 0.00012667
996.2 | 0.99983 | 0.00012262
996.4 | 0.99984 | 0.00011879
996.6 | 0.99984 | 0.00011516
996.8 | 0.99984 | 0.00011174
997 | 0.99985 | 0.00010851
997.2 | 0.99985 | 0.00010547
997.4 | 0.99986 | 0.0001026
997.6 | 0.99986 | 9.9915e-05
997.8 | 0.99986 | 9.7392e-05
998 | 0.99987 | 9.5028e-05
998.2 | 0.99987 | 9.2817e-05
998.4 | 0.99987 | 9.0751e-05
998.6 | 0.99988 | 8.8824e-05
998.8 | 0.99988 | 8.703e-05
999 | 0.99988 | 8.5363e-05
999.2 | 0.99988 | 8.3819e-05
999.4 | 0.99988 | 8.2391e-05
999.6 | 0.99989 | 8.1075e-05
999.8 | 0.99989 | 7.9867e-05
1000 | 0.99989 | 7.8763e-05
1000.2 | 0.99989 | 7.776e-05
1000.4 | 0.99989 | 7.6853e-05
1000.6 | 0.99989 | 7.604e-05
1000.8 | 0.99989 | 7.5318e-05
1001 | 0.99989 | 7.4684e-05
1001.2 | 0.9999 | 7.4137e-05
1001.4 | 0.9999 | 7.3675e-05
1001.6 | 0.9999 | 7.3296e-05
1001.8 | 0.9999 | 7.2999e-05
1002 | 0.9999 | 7.2782e-05
1002.2 | 0.9999 | 7.2646e-05
1002.4 | 0.9999 | 7.2589e-05
1002.6 | 0.9999 | 7.2612e-05
1002.8 | 0.9999 | 7.2714e-05
1003 | 0.9999 | 7.2896e-05
1003.2 | 0.9999 | 7.3158e-05
1003.4 | 0.9999 | 7.3502e-05
1003.6 | 0.9999 | 7.3928e-05
1003.8 | 0.9999 | 7.4437e-05
1004 | 0.99989 | 7.5032e-05
1004.2 | 0.99989 | 7.5714e-05
1004.4 | 0.99989 | 7.6486e-05
1004.6 | 0.99989 | 7.7349e-05
1004.8 | 0.99989 | 7.8307e-05
1005 | 0.99989 | 7.9363e-05
1005.2 | 0.99989 | 8.052e-05
1005.4 | 0.99988 | 8.1782e-05
1005.6 | 0.99988 | 8.3153e-05
1005.8 | 0.99988 | 8.4637e-05
1006 | 0.99988 | 8.6238e-05
1006.2 | 0.99988 | 8.7962e-05
1006.4 | 0.99987 | 8.9814e-05
1006.6 | 0.99987 | 9.1799e-05
1006.8 | 0.99987 | 9.3923e-05
1007 | 0.99987 | 9.6192e-05
1007.2 | 0.99986 | 9.8611e-05
1007.4 | 0.99986 | 0.00010119
1007.6 | 0.99985 | 0.00010393
1007.8 | 0.99985 | 0.00010684
1008 | 0.99985 | 0.00010992
1008.2 | 0.99984 | 0.00011318
1008.4 | 0.99984 | 0.00011663
1008.6 | 0.99983 | 0.00012027
1008.8 | 0.99983 | 0.0001241
1009 | 0.99982 | 0.00012813
1009.2 | 0.99982 | 0.00013235
1009.4 | 0.99981 | 0.00013677
1009.6 | 0.9998 | 0.00014137
1009.8 | 0.9998 | 0.00014616
1010 | 0.99979 | 0.00015113
1010.2 | 0.99978 | 0.00015625
1010.4 | 0.99978 | 0.00016151
1010.6 | 0.99977 | 0.00016688
1010.8 | 0.99976 | 0.00017235
1011 | 0.99975 | 0.00017787
1011.2 | 0.99975 | 0.00018339
1011.4 | 0.99974 | 0.00018889
1011.6 | 0.99973 | 0.0001943
1011.8 | 0.99972 | 0.00019957
1012 | 0.99972 | 0.00020464
1012.2 | 0.99971 | 0.00020945
1012.4 | 0.99971 | 0.00021392
1012.6 | 0.9997 | 0.00021801
1012.8 | 0.99969 | 0.00022165
1013 | 0.99969 | 0.00022478
1013.2 | 0.99969 | 0.00022735
1013.4 | 0.99968 | 0.00022932
1013.6 | 0.99968 | 0.00023067
1013.8 | 0.99968 | 0.00023136
1014 | 0.99968 | 0.0002314
1014.2 | 0.99968 | 0.00023078
1014.4 | 0.99969 | 0.00022953
1014.6 | 0.99969 | 0.00022766
1014.8 | 0.99969 | 0.00022522
1015 | 0.9997 | 0.00022226
1015.2 | 0.9997 | 0.00021882
1015.4 | 0.99971 | 0.00021497
1015.6 | 0.99971 | 0.00021077
1015.8 | 0.99972 | 0.00020627
1016 | 0.99973 | 0.00020155
1016.2 | 0.99973 | 0.00019666
1016.4 | 0.99974 | 0.00019165
1016.6 | 0.99975 | 0.00018659
1016.8 | 0.99975 | 0.0001815
1017 | 0.99976 | 0.00017644
1017.2 | 0.99977 | 0.00017144
1017.4 | 0.99977 | 0.00016653
1017.6 | 0.99978 | 0.00016173
1017.8 | 0.99979 | 0.00015706
1018 | 0.99979 | 0.00015255
1018.2 | 0.9998 | 0.00014819
1018.4 | 0.99981 | 0.00014401
1018.6 | 0.99981 | 0.00014
1018.8 | 0.99982 | 0.00013617
1019 | 0.99982 | 0.00013251
1019.2 | 0.99983 | 0.00012904
1019.4 | 0.99983 | 0.00012574
1019.6 | 0.99984 | 0.00012262
1019.8 | 0.99984 | 0.00011966
1020 | 0.99984 | 0.00011688
1020.2 | 0.99985 | 0.00011425
1020.4 | 0.99985 | 0.00011178
1020.6 | 0.99985 | 0.00010945
1020.8 | 0.99986 | 0.00010728
1021 | 0.99986 | 0.00010524
1021.2 | 0.99986 | 0.00010334
1021.4 | 0.99987 | 0.00010157
1021.6 | 0.99987 | 9.9926e-05
1021.8 | 0.99987 | 9.8401e-05
1022 | 0.99987 | 9.6993e-05
1022.2 | 0.99987 | 9.5698e-05
1022.4 | 0.99988 | 9.4512e-05
1022.6 | 0.99988 | 9.3431e-05
1022.8 | 0.99988 | 9.2451e-05
1023 | 0.99988 | 9.1572e-05
1023.2 | 0.99988 | 9.0788e-05
1023.4 | 0.99988 | 9.0099e-05
1023.6 | 0.99988 | 8.9501e-05
1023.8 | 0.99988 | 8.8994e-05
1024 | 0.99988 | 8.8575e-05
1024.2 | 0.99989 | 8.8243e-05
1024.4 | 0.99989 | 8.7997e-05
1024.6 | 0.99989 | 8.7836e-05
1024.8 | 0.99989 | 8.776e-05
1025 | 0.99989 | 8.7768e-05
1025.2 | 0.99989 | 8.7861e-05
1025.4 | 0.99989 | 8.8039e-05
1025.6 | 0.99989 | 8.8301e-05
1025.8 | 0.99989 | 8.8649e-05
1026 | 0.99989 | 8.9084e-05
1026.2 | 0.99989 | 8.9608e-05
1026.4 | 0.99988 | 9.0222e-05
1026.6 | 0.99988 | 9.0927e-05
1026.8 | 0.99988 | 9.1726e-05
1027 | 0.99988 | 9.2622e-05
1027.2 | 0.99988 | 9.3617e-05
1027.4 | 0.99988 | 9.4716e-05
1027.6 | 0.99988 | 9.592e-05
1027.8 | 0.99988 | 9.7234e-05
1028 | 0.99987 | 9.8664e-05
1028.2 | 0.99987 | 0.00010021
1028.4 | 0.99987 | 0.00010188
1028.6 | 0.99987 | 0.00010369
1028.8 | 0.99987 | 0.00010563
1029 | 0.99986 | 0.00010771
1029.2 | 0.99986 | 0.00010994
1029.4 | 0.99986 | 0.00011233
1029.6 | 0.99986 | 0.00011489
1029.8 | 0.99985 | 0.00011762
1030 | 0.99985 | 0.00012053
1030.2 | 0.99984 | 0.00012364
1030.4 | 0.99984 | 0.00012695
1030.6 | 0.99984 | 0.00013048
1030.8 | 0.99983 | 0.00013423
1031 | 0.99983 | 0.00013821
1031.2 | 0.99982 | 0.00014245
1031.4 | 0.99982 | 0.00014694
1031.6 | 0.99981 | 0.00015169
1031.8 | 0.9998 | 0.00015673
1032 | 0.9998 | 0.00016205
1032.2 | 0.99979 | 0.00016767
1032.4 | 0.99978 | 0.00017359
1032.6 | 0.99978 | 0.00017981
1032.8 | 0.99977 | 0.00018632
1033 | 0.99976 | 0.00019313
1033.2 | 0.99975 | 0.00020023
1033.4 | 0.99974 | 0.00020758
1033.6 | 0.99973 | 0.00021517
1033.8 | 0.99972 | 0.00022296
1034 | 0.99971 | 0.0002309
1034.2 | 0.9997 | 0.00023893
1034.4 | 0.99969 | 0.00024697
1034.6 | 0.99968 | 0.00025495
1034.8 | 0.99967 | 0.00026275
1035 | 0.99966 | 0.00027027
1035.2 | 0.99965 | 0.00027738
1035.4 | 0.99964 | 0.00028395
1035.6 | 0.99964 | 0.00028986
1035.8 | 0.99963 | 0.00029497
1036 | 0.99962 | 0.00029915
1036.2 | 0.99962 | 0.00030229
1036.4 | 0.99962 | 0.0003043
1036.6 | 0.99961 | 0.00030511
1036.8 | 0.99961 | 0.00030467
1037 | 0.99962 | 0.00030298
1037.2 | 0.99962 | 0.00030003
1037.4 | 0.99962 | 0.00029589
1037.6 | 0.99963 | 0.00029061
1037.8 | 0.99964 | 0.00028431
1038 | 0.99965 | 0.00027709
1038.2 | 0.99965 | 0.00026908
1038.4 | 0.99967 | 0.00026043
1038.6 | 0.99968 | 0.00025129
1038.8 | 0.99969 | 0.00024178
1039 | 0.9997 | 0.00023205
1039.2 | 0.99971 | 0.00022223
1039.4 | 0.99972 | 0.00021243
1039.6 | 0.99974 | 0.00020274
1039.8 | 0.99975 | 0.00019324
1040 | 0.99976 | 0.00018401
1040.2 | 0.99977 | 0.00017509
1040.4 | 0.99978 | 0.00016652
1040.6 | 0.99979 | 0.00015833
1040.8 | 0.9998 | 0.00015053
1041 | 0.99981 | 0.00014313
1041.2 | 0.99982 | 0.00013614
1041.4 | 0.99983 | 0.00012953
1041.6 | 0.99984 | 0.00012332
1041.8 | 0.99984 | 0.00011747
1042 | 0.99985 | 0.00011198
1042.2 | 0.99986 | 0.00010684
1042.4 | 0.99986 | 0.00010202
1042.6 | 0.99987 | 9.7499e-05
1042.8 | 0.99987 | 9.3272e-05
1043 | 0.99988 | 8.9317e-05
1043.2 | 0.99988 | 8.5617e-05
1043.4 | 0.99989 | 8.2156e-05
1043.6 | 0.99989 | 7.8918e-05
1043.8 | 0.99989 | 7.589e-05
1044 | 0.9999 | 7.3058e-05
1044.2 | 0.9999 | 7.0409e-05
1044.4 | 0.99991 | 6.793e-05
1044.6 | 0.99991 | 6.5612e-05
1044.8 | 0.99991 | 6.3443e-05
1045 | 0.99991 | 6.1413e-05
1045.2 | 0.99992 | 5.9514e-05
1045.4 | 0.99992 | 5.7738e-05
1045.6 | 0.99992 | 5.6077e-05
1045.8 | 0.99992 | 5.4524e-05
1046 | 0.99992 | 5.3072e-05
1046.2 | 0.99993 | 5.1715e-05
1046.4 | 0.99993 | 5.0448e-05
1046.6 | 0.99993 | 4.9265e-05
1046.8 | 0.99993 | 4.8163e-05
1047 | 0.99993 | 4.7136e-05
1047.2 | 0.99993 | 4.6181e-05
1047.4 | 0.99993 | 4.5293e-05
1047.6 | 0.99993 | 4.4471e-05
1047.8 | 0.99994 | 4.371e-05
1048 | 0.99994 | 4.3008e-05
1048.2 | 0.99994 | 4.2363e-05
1048.4 | 0.99994 | 4.1771e-05
1048.6 | 0.99994 | 4.1232e-05
1048.8 | 0.99994 | 4.0743e-05
1049 | 0.99994 | 4.0302e-05
1049.2 | 0.99994 | 3.9909e-05
1049.4 | 0.99994 | 3.9562e-05
1049.6 | 0.99994 | 3.9259e-05
1049.8 | 0.99994 | 3.9e-05
1050 | 0.99994 | 3.8785e-05
1050.2 | 0.99994 | 3.8611e-05
1050.4 | 0.99994 | 3.848e-05
1050.6 | 0.99994 | 3.839e-05
1050.8 | 0.99994 | 3.8342e-05
1051 | 0.99994 | 3.8336e-05
1051.2 | 0.99994 | 3.837e-05
1051.4 | 0.99994 | 3.8447e-05
1051.6 | 0.99994 | 3.8566e-05
1051.8 | 0.99994 | 3.8728e-05
1052 | 0.99994 | 3.8934e-05
1052.2 | 0.99994 | 3.9183e-05
1052.4 | 0.99994 | 3.9478e-05
1052.6 | 0.99994 | 3.982e-05
1052.8 | 0.99994 | 4.021e-05
1053 | 0.99993 | 4.0649e-05
1053.2 | 0.99993 | 4.1138e-05
1053.4 | 0.99993 | 4.1681e-05
1053.6 | 0.99993 | 4.2279e-05
1053.8 | 0.99993 | 4.2933e-05
1054 | 0.99993 | 4.3647e-05
1054.2 | 0.99993 | 4.4422e-05
1054.4 | 0.99993 | 4.5262e-05
1054.6 | 0.99992 | 4.6169e-05
1054.8 | 0.99992 | 4.7147e-05
1055 | 0.99992 | 4.8197e-05
1055.2 | 0.99992 | 4.9324e-05
1055.4 | 0.99992 | 5.0531e-05
1055.6 | 0.99991 | 5.1821e-05
1055.8 | 0.99991 | 5.3196e-05
1056 | 0.99991 | 5.466e-05
1056.2 | 0.99991 | 5.6217e-05
1056.4 | 0.9999 | 5.7867e-05
1056.6 | 0.9999 | 5.9613e-05
1056.8 | 0.9999 | 6.1456e-05
1057 | 0.99989 | 6.3396e-05
1057.2 | 0.99989 | 6.5432e-05
1057.4 | 0.99989 | 6.7562e-05
1057.6 | 0.99988 | 6.9781e-05
1057.8 | 0.99988 | 7.2082e-05
1058 | 0.99987 | 7.4456e-05
1058.2 | 0.99987 | 7.6892e-05
1058.4 | 0.99986 | 7.9374e-05
1058.6 | 0.99986 | 8.1882e-05
1058.8 | 0.99986 | 8.4394e-05
1059 | 0.99985 | 8.6883e-05
1059.2 | 0.99985 | 8.9318e-05
1059.4 | 0.99984 | 9.1665e-05
1059.6 | 0.99984 | 9.3887e-05
1059.8 | 0.99983 | 9.5945e-05
1060 | 0.99983 | 9.78e-05
1060.2 | 0.99983 | 9.9411e-05
1060.4 | 0.99982 | 0.00010074
1060.6 | 0.99982 | 0.00010176
1060.8 | 0.99982 | 0.00010244
1061 | 0.99982 | 0.00010275
1061.2 | 0.99982 | 0.00010268
1061.4 | 0.99982 | 0.00010224
1061.6 | 0.99982 | 0.00010141
1061.8 | 0.99982 | 0.00010022
1062 | 0.99982 | 9.8673e-05
1062.2 | 0.99983 | 9.681e-05
1062.4 | 0.99983 | 9.4658e-05
1062.6 | 0.99983 | 9.2256e-05
1062.8 | 0.99984 | 8.9644e-05
1063 | 0.99984 | 8.6864e-05
1063.2 | 0.99985 | 8.3959e-05
1063.4 | 0.99985 | 8.0967e-05
1063.6 | 0.99986 | 7.7929e-05
1063.8 | 0.99986 | 7.4878e-05
1064 | 0.99987 | 7.1846e-05
1064.2 | 0.99987 | 6.8857e-05
1064.4 | 0.99988 | 6.5935e-05
1064.6 | 0.99988 | 6.3096e-05
1064.8 | 0.99989 | 6.0355e-05
1065 | 0.99989 | 5.7721e-05
1065.2 | 0.9999 | 5.5202e-05
1065.4 | 0.9999 | 5.28e-05
1065.6 | 0.99991 | 5.0519e-05
1065.8 | 0.99991 | 4.8357e-05
1066 | 0.99991 | 4.6314e-05
1066.2 | 0.99992 | 4.4386e-05
1066.4 | 0.99992 | 4.2569e-05
1066.6 | 0.99992 | 4.0861e-05
1066.8 | 0.99993 | 3.9256e-05
1067 | 0.99993 | 3.775e-05
1067.2 | 0.99993 | 3.6336e-05
1067.4 | 0.99993 | 3.5012e-05
1067.6 | 0.99994 | 3.3771e-05
1067.8 | 0.99994 | 3.2609e-05
1068 | 0.99994 | 3.1522e-05
1068.2 | 0.99994 | 3.0505e-05
1068.4 | 0.99994 | 2.9553e-05
1068.6 | 0.99995 | 2.8664e-05
1068.8 | 0.99995 | 2.7833e-05
1069 | 0.99995 | 2.7057e-05
1069.2 | 0.99995 | 2.6333e-05
1069.4 | 0.99995 | 2.5658e-05
1069.6 | 0.99995 | 2.5028e-05
1069.8 | 0.99995 | 2.4442e-05
1070 | 0.99995 | 2.3897e-05
1070.2 | 0.99996 | 2.339e-05
1070.4 | 0.99996 | 2.2921e-05
1070.6 | 0.99996 | 2.2486e-05
1070.8 | 0.99996 | 2.2084e-05
1071 | 0.99996 | 2.1713e-05
1071.2 | 0.99996 | 2.1373e-05
1071.4 | 0.99996 | 2.1062e-05
1071.6 | 0.99996 | 2.0778e-05
1071.8 | 0.99996 | 2.0521e-05
1072 | 0.99996 | 2.0289e-05
1072.2 | 0.99996 | 2.0082e-05
1072.4 | 0.99996 | 1.9899e-05
1072.6 | 0.99996 | 1.974e-05
1072.8 | 0.99996 | 1.9603e-05
1073 | 0.99996 | 1.9488e-05
1073.2 | 0.99996 | 1.9395e-05
1073.4 | 0.99996 | 1.9324e-05
1073.6 | 0.99996 | 1.9274e-05
1073.8 | 0.99996 | 1.9245e-05
1074 | 0.99996 | 1.9238e-05
1074.2 | 0.99996 | 1.9251e-05
1074.4 | 0.99996 | 1.9286e-05
1074.6 | 0.99996 | 1.9341e-05
1074.8 | 0.99996 | 1.9419e-05
1075 | 0.99996 | 1.9518e-05
1075.2 | 0.99996 | 1.9639e-05
1075.4 | 0.99996 | 1.9783e-05
1075.6 | 0.99996 | 1.995e-05
1075.8 | 0.99996 | 2.0142e-05
1076 | 0.99996 | 2.0358e-05
1076.2 | 0.99996 | 2.0599e-05
1076.4 | 0.99996 | 2.0867e-05
1076.6 | 0.99996 | 2.1163e-05
1076.8 | 0.99996 | 2.1487e-05
1077 | 0.99996 | 2.1842e-05
1077.2 | 0.99996 | 2.2228e-05
1077.4 | 0.99996 | 2.2647e-05
1077.6 | 0.99996 | 2.3101e-05
1077.8 | 0.99996 | 2.3592e-05
1078 | 0.99995 | 2.4122e-05
1078.2 | 0.99995 | 2.4692e-05
1078.4 | 0.99995 | 2.5306e-05
1078.6 | 0.99995 | 2.5966e-05
1078.8 | 0.99995 | 2.6675e-05
1079 | 0.99995 | 2.7435e-05
1079.2 | 0.99995 | 2.8251e-05
1079.4 | 0.99994 | 2.9124e-05
1079.6 | 0.99994 | 3.0059e-05
1079.8 | 0.99994 | 3.1059e-05
1080 | 0.99994 | 3.2129e-05
1080.2 | 0.99994 | 3.3272e-05
1080.4 | 0.99994 | 3.4492e-05
1080.6 | 0.99993 | 3.5793e-05
1080.8 | 0.99993 | 3.718e-05
1081 | 0.99993 | 3.8655e-05
1081.2 | 0.99992 | 4.0223e-05
1081.4 | 0.99992 | 4.1886e-05
1081.6 | 0.99992 | 4.3646e-05
1081.8 | 0.99992 | 4.5505e-05
1082 | 0.99991 | 4.7463e-05
1082.2 | 0.99991 | 4.9517e-05
1082.4 | 0.9999 | 5.1665e-05
1082.6 | 0.9999 | 5.3899e-05
1082.8 | 0.9999 | 5.6212e-05
1083 | 0.99989 | 5.8589e-05
1083.2 | 0.99989 | 6.1016e-05
1083.4 | 0.99988 | 6.3473e-05
1083.6 | 0.99988 | 6.5933e-05
1083.8 | 0.99987 | 6.8371e-05
1084 | 0.99987 | 7.0752e-05
1084.2 | 0.99987 | 7.3042e-05
1084.4 | 0.99986 | 7.5202e-05
1084.6 | 0.99986 | 7.7195e-05
1084.8 | 0.99985 | 7.8981e-05
1085 | 0.99985 | 8.0524e-05
1085.2 | 0.99985 | 8.1792e-05
1085.4 | 0.99985 | 8.2757e-05
1085.6 | 0.99985 | 8.3396e-05
1085.8 | 0.99985 | 8.3698e-05
1086 | 0.99985 | 8.3655e-05
1086.2 | 0.99985 | 8.3269e-05
1086.4 | 0.99985 | 8.2552e-05
1086.6 | 0.99985 | 8.152e-05
1086.8 | 0.99985 | 8.0198e-05
1087 | 0.99985 | 7.8616e-05
1087.2 | 0.99986 | 7.6807e-05
1087.4 | 0.99986 | 7.4808e-05
1087.6 | 0.99987 | 7.2657e-05
1087.8 | 0.99987 | 7.0392e-05
1088 | 0.99987 | 6.8048e-05
1088.2 | 0.99988 | 6.566e-05
1088.4 | 0.99988 | 6.3257e-05
1088.6 | 0.99989 | 6.0867e-05
1088.8 | 0.99989 | 5.8512e-05
1089 | 0.9999 | 5.621e-05
1089.2 | 0.9999 | 5.3975e-05
1089.4 | 0.9999 | 5.182e-05
1089.6 | 0.99991 | 4.9752e-05
1089.8 | 0.99991 | 4.7776e-05
1090 | 0.99992 | 4.5895e-05
1090.2 | 0.99992 | 4.4111e-05
1090.4 | 0.99992 | 4.2424e-05
1090.6 | 0.99992 | 4.0831e-05
1090.8 | 0.99993 | 3.9331e-05
1091 | 0.99993 | 3.7921e-05
1091.2 | 0.99993 | 3.6596e-05
1091.4 | 0.99994 | 3.5355e-05
1091.6 | 0.99994 | 3.4191e-05
1091.8 | 0.99994 | 3.3103e-05
1092 | 0.99994 | 3.2086e-05
1092.2 | 0.99994 | 3.1136e-05
1092.4 | 0.99994 | 3.0249e-05
1092.6 | 0.99995 | 2.9423e-05
1092.8 | 0.99995 | 2.8653e-05
1093 | 0.99995 | 2.7937e-05
1093.2 | 0.99995 | 2.7271e-05
1093.4 | 0.99995 | 2.6654e-05
1093.6 | 0.99995 | 2.6081e-05
1093.8 | 0.99995 | 2.5552e-05
1094 | 0.99995 | 2.5064e-05
1094.2 | 0.99996 | 2.4614e-05
1094.4 | 0.99996 | 2.4201e-05
1094.6 | 0.99996 | 2.3823e-05
1094.8 | 0.99996 | 2.3479e-05
1095 | 0.99996 | 2.3167e-05
1095.2 | 0.99996 | 2.2886e-05
1095.4 | 0.99996 | 2.2634e-05
1095.6 | 0.99996 | 2.2412e-05
1095.8 | 0.99996 | 2.2217e-05
1096 | 0.99996 | 2.2049e-05
1096.2 | 0.99996 | 2.1908e-05
1096.4 | 0.99996 | 2.1793e-05
1096.6 | 0.99996 | 2.1703e-05
1096.8 | 0.99996 | 2.1638e-05
1097 | 0.99996 | 2.1599e-05
1097.2 | 0.99996 | 2.1583e-05
1097.4 | 0.99996 | 2.1592e-05
1097.6 | 0.99996 | 2.1626e-05
1097.8 | 0.99996 | 2.1685e-05
1098 | 0.99996 | 2.1768e-05
1098.2 | 0.99996 | 2.1877e-05
1098.4 | 0.99996 | 2.2011e-05
1098.6 | 0.99996 | 2.2172e-05
1098.8 | 0.99996 | 2.236e-05
1099 | 0.99996 | 2.2575e-05
1099.2 | 0.99996 | 2.2819e-05
1099.4 | 0.99996 | 2.3092e-05
1099.6 | 0.99996 | 2.3396e-05
1099.8 | 0.99996 | 2.3732e-05
1100 | 0.99996 | 2.4101e-05
1100.2 | 0.99996 | 2.4506e-05
1100.4 | 0.99996 | 2.4947e-05
1100.6 | 0.99996 | 2.5427e-05
1100.8 | 0.99996 | 2.5948e-05
1101 | 0.99996 | 2.6513e-05
1101.2 | 0.99996 | 2.7124e-05
1101.4 | 0.99995 | 2.7783e-05
1101.6 | 0.99995 | 2.8495e-05
1101.8 | 0.99995 | 2.9263e-05
1102 | 0.99995 | 3.0091e-05
1102.2 | 0.99995 | 3.0983e-05
1102.4 | 0.99995 | 3.1943e-05
1102.6 | 0.99995 | 3.2977e-05
1102.8 | 0.99995 | 3.4091e-05
1103 | 0.99994 | 3.529e-05
1103.2 | 0.99994 | 3.6582e-05
1103.4 | 0.99994 | 3.7974e-05
1103.6 | 0.99994 | 3.9473e-05
1103.8 | 0.99994 | 4.1089e-05
1104 | 0.99993 | 4.2831e-05
1104.2 | 0.99993 | 4.4709e-05
1104.4 | 0.99993 | 4.6735e-05
1104.6 | 0.99993 | 4.8919e-05
1104.8 | 0.99992 | 5.1276e-05
1105 | 0.99992 | 5.3817e-05
1105.2 | 0.99992 | 5.6557e-05
1105.4 | 0.99991 | 5.9511e-05
1105.6 | 0.99991 | 6.2692e-05
1105.8 | 0.9999 | 6.6115e-05
1106 | 0.9999 | 6.9794e-05
1106.2 | 0.99989 | 7.374e-05
1106.4 | 0.99989 | 7.7964e-05
1106.6 | 0.99988 | 8.2473e-05
1106.8 | 0.99987 | 8.7268e-05
1107 | 0.99987 | 9.2345e-05
1107.2 | 0.99986 | 9.7693e-05
1107.4 | 0.99985 | 0.00010329
1107.6 | 0.99985 | 0.0001091
1107.8 | 0.99984 | 0.00011508
1108 | 0.99983 | 0.00012117
1108.2 | 0.99982 | 0.00012729
1108.4 | 0.99981 | 0.00013335
1108.6 | 0.99981 | 0.00013924
1108.8 | 0.9998 | 0.00014487
1109 | 0.99979 | 0.00015012
1109.2 | 0.99979 | 0.00015486
1109.4 | 0.99978 | 0.00015901
1109.6 | 0.99978 | 0.00016247
1109.8 | 0.99977 | 0.00016518
1110 | 0.99977 | 0.00016707
1110.2 | 0.99977 | 0.00016811
1110.4 | 0.99977 | 0.00016831
1110.6 | 0.99977 | 0.00016766
1110.8 | 0.99977 | 0.00016621
1111 | 0.99978 | 0.00016399
1111.2 | 0.99978 | 0.00016106
1111.4 | 0.99979 | 0.00015752
1111.6 | 0.99979 | 0.00015344
1111.8 | 0.9998 | 0.00014893
1112 | 0.99981 | 0.00014409
1112.2 | 0.99981 | 0.00013902
1112.4 | 0.99982 | 0.0001338
1112.6 | 0.99983 | 0.00012854
1112.8 | 0.99984 | 0.0001233
1113 | 0.99984 | 0.00011814
1113.2 | 0.99985 | 0.00011312
1113.4 | 0.99986 | 0.00010828
1113.6 | 0.99986 | 0.00010364
1113.8 | 0.99987 | 9.9226e-05
1114 | 0.99987 | 9.5043e-05
1114.2 | 0.99988 | 9.1098e-05
1114.4 | 0.99988 | 8.739e-05
1114.6 | 0.99989 | 8.3916e-05
1114.8 | 0.99989 | 8.0669e-05
1115 | 0.9999 | 7.764e-05
1115.2 | 0.9999 | 7.482e-05
1115.4 | 0.99991 | 7.2199e-05
1115.6 | 0.99991 | 6.9764e-05
1115.8 | 0.99991 | 6.7507e-05
1116 | 0.99991 | 6.5416e-05
1116.2 | 0.99992 | 6.3482e-05
1116.4 | 0.99992 | 6.1695e-05
1116.6 | 0.99992 | 6.0046e-05
1116.8 | 0.99992 | 5.8526e-05
1117 | 0.99993 | 5.7127e-05
1117.2 | 0.99993 | 5.5844e-05
1117.4 | 0.99993 | 5.4667e-05
1117.6 | 0.99993 | 5.3593e-05
1117.8 | 0.99993 | 5.2614e-05
1118 | 0.99993 | 5.1727e-05
1118.2 | 0.99993 | 5.0926e-05
1118.4 | 0.99993 | 5.0207e-05
1118.6 | 0.99993 | 4.9566e-05
1118.8 | 0.99994 | 4.9001e-05
1119 | 0.99994 | 4.8508e-05
1119.2 | 0.99994 | 4.8085e-05
1119.4 | 0.99994 | 4.773e-05
1119.6 | 0.99994 | 4.744e-05
1119.8 | 0.99994 | 4.7214e-05
1120 | 0.99994 | 4.7051e-05
1120.2 | 0.99994 | 4.695e-05
1120.4 | 0.99994 | 4.6911e-05
1120.6 | 0.99994 | 4.6931e-05
1120.8 | 0.99994 | 4.7013e-05
1121 | 0.99994 | 4.7155e-05
1121.2 | 0.99994 | 4.7358e-05
1121.4 | 0.99994 | 4.7622e-05
1121.6 | 0.99994 | 4.795e-05
1121.8 | 0.99994 | 4.834e-05
1122 | 0.99994 | 4.8796e-05
1122.2 | 0.99993 | 4.9319e-05
1122.4 | 0.99993 | 4.9911e-05
1122.6 | 0.99993 | 5.0573e-05
1122.8 | 0.99993 | 5.131e-05
1123 | 0.99993 | 5.2123e-05
1123.2 | 0.99993 | 5.3017e-05
1123.4 | 0.99993 | 5.3994e-05
1123.6 | 0.99993 | 5.5059e-05
1123.8 | 0.99993 | 5.6216e-05
1124 | 0.99992 | 5.747e-05
1124.2 | 0.99992 | 5.8826e-05
1124.4 | 0.99992 | 6.029e-05
1124.6 | 0.99992 | 6.1869e-05
1124.8 | 0.99992 | 6.3568e-05
1125 | 0.99991 | 6.5395e-05
1125.2 | 0.99991 | 6.7358e-05
1125.4 | 0.99991 | 6.9463e-05
1125.6 | 0.99991 | 7.1721e-05
1125.8 | 0.9999 | 7.414e-05
1126 | 0.9999 | 7.6728e-05
1126.2 | 0.9999 | 7.9495e-05
1126.4 | 0.99989 | 8.2451e-05
1126.6 | 0.99989 | 8.5605e-05
1126.8 | 0.99988 | 8.8965e-05
1127 | 0.99988 | 9.2541e-05
1127.2 | 0.99987 | 9.6339e-05
1127.4 | 0.99987 | 0.00010037
1127.6 | 0.99986 | 0.00010463
1127.8 | 0.99986 | 0.00010912
1128 | 0.99985 | 0.00011384
1128.2 | 0.99985 | 0.00011879
1128.4 | 0.99984 | 0.00012396
1128.6 | 0.99983 | 0.00012932
1128.8 | 0.99983 | 0.00013485
1129 | 0.99982 | 0.00014052
1129.2 | 0.99981 | 0.00014628
1129.4 | 0.9998 | 0.0001521
1129.6 | 0.9998 | 0.00015789
1129.8 | 0.99979 | 0.00016361
1130 | 0.99978 | 0.00016917
1130.2 | 0.99978 | 0.0001745
1130.4 | 0.99977 | 0.0001795
1130.6 | 0.99976 | 0.00018409
1130.8 | 0.99976 | 0.00018819
1131 | 0.99975 | 0.00019172
1131.2 | 0.99975 | 0.00019462
1131.4 | 0.99975 | 0.00019682
1131.6 | 0.99975 | 0.00019829
1131.8 | 0.99975 | 0.000199
1132 | 0.99975 | 0.00019893
1132.2 | 0.99975 | 0.0001981
1132.4 | 0.99975 | 0.00019652
1132.6 | 0.99975 | 0.00019422
1132.8 | 0.99976 | 0.00019128
1133 | 0.99976 | 0.00018773
1133.2 | 0.99977 | 0.00018366
1133.4 | 0.99977 | 0.00017915
1133.6 | 0.99978 | 0.00017428
1133.8 | 0.99979 | 0.00016913
1134 | 0.99979 | 0.00016379
1134.2 | 0.9998 | 0.00015833
1134.4 | 0.99981 | 0.00015282
1134.6 | 0.99981 | 0.00014733
1134.8 | 0.99982 | 0.00014189
1135 | 0.99983 | 0.00013657
1135.2 | 0.99983 | 0.00013139
1135.4 | 0.99984 | 0.00012638
1135.6 | 0.99985 | 0.00012157
1135.8 | 0.99985 | 0.00011696
1136 | 0.99986 | 0.00011257
1136.2 | 0.99986 | 0.00010839
1136.4 | 0.99987 | 0.00010443
1136.6 | 0.99987 | 0.00010069
1136.8 | 0.99988 | 9.7166e-05
1137 | 0.99988 | 9.3846e-05
1137.2 | 0.99989 | 9.0725e-05
1137.4 | 0.99989 | 8.7795e-05
1137.6 | 0.99989 | 8.5049e-05
1137.8 | 0.9999 | 8.2478e-05
1138 | 0.9999 | 8.0072e-05
1138.2 | 0.9999 | 7.7825e-05
1138.4 | 0.99991 | 7.5727e-05
1138.6 | 0.99991 | 7.3771e-05
1138.8 | 0.99991 | 7.1948e-05
1139 | 0.99991 | 7.0253e-05
1139.2 | 0.99991 | 6.8677e-05
1139.4 | 0.99992 | 6.7216e-05
1139.6 | 0.99992 | 6.5862e-05
1139.8 | 0.99992 | 6.4611e-05
1140 | 0.99992 | 6.3456e-05
1140.2 | 0.99992 | 6.2395e-05
1140.4 | 0.99992 | 6.1421e-05
1140.6 | 0.99992 | 6.0531e-05
1140.8 | 0.99993 | 5.9722e-05
1141 | 0.99993 | 5.899e-05
1141.2 | 0.99993 | 5.8332e-05
1141.4 | 0.99993 | 5.7746e-05
1141.6 | 0.99993 | 5.7228e-05
1141.8 | 0.99993 | 5.6778e-05
1142 | 0.99993 | 5.6391e-05
1142.2 | 0.99993 | 5.6068e-05
1142.4 | 0.99993 | 5.5807e-05
1142.6 | 0.99993 | 5.5606e-05
1142.8 | 0.99993 | 5.5464e-05
1143 | 0.99993 | 5.5381e-05
1143.2 | 0.99993 | 5.5355e-05
1143.4 | 0.99993 | 5.5386e-05
1143.6 | 0.99993 | 5.5475e-05
1143.8 | 0.99993 | 5.562e-05
1144 | 0.99993 | 5.5823e-05
1144.2 | 0.99993 | 5.6082e-05
1144.4 | 0.99993 | 5.6399e-05
1144.6 | 0.99993 | 5.6773e-05
1144.8 | 0.99993 | 5.7207e-05
1145 | 0.99993 | 5.77e-05
1145.2 | 0.99993 | 5.8253e-05
1145.4 | 0.99992 | 5.8868e-05
1145.6 | 0.99992 | 5.9546e-05
1145.8 | 0.99992 | 6.0289e-05
1146 | 0.99992 | 6.1097e-05
1146.2 | 0.99992 | 6.1972e-05
1146.4 | 0.99992 | 6.2916e-05
1146.6 | 0.99992 | 6.3932e-05
1146.8 | 0.99992 | 6.502e-05
1147 | 0.99991 | 6.6183e-05
1147.2 | 0.99991 | 6.7423e-05
1147.4 | 0.99991 | 6.8741e-05
1147.6 | 0.99991 | 7.014e-05
1147.8 | 0.99991 | 7.1622e-05
1148 | 0.9999 | 7.3187e-05
1148.2 | 0.9999 | 7.4838e-05
1148.4 | 0.9999 | 7.6577e-05
1148.6 | 0.9999 | 7.8403e-05
1148.8 | 0.99989 | 8.0317e-05
1149 | 0.99989 | 8.2321e-05
1149.2 | 0.99989 | 8.4412e-05
1149.4 | 0.99989 | 8.6589e-05
1149.6 | 0.99988 | 8.8852e-05
1149.8 | 0.99988 | 9.1195e-05
1150 | 0.99988 | 9.3616e-05
1150.2 | 0.99987 | 9.6107e-05
1150.4 | 0.99987 | 9.8662e-05
1150.6 | 0.99987 | 0.00010127
1150.8 | 0.99986 | 0.00010393
1151 | 0.99986 | 0.00010662
1151.2 | 0.99985 | 0.00010932
1151.4 | 0.99985 | 0.00011203
1151.6 | 0.99985 | 0.00011472
1151.8 | 0.99984 | 0.00011737
1152 | 0.99984 | 0.00011996
1152.2 | 0.99984 | 0.00012247
1152.4 | 0.99983 | 0.00012487
1152.6 | 0.99983 | 0.00012714
1152.8 | 0.99983 | 0.00012925
1153 | 0.99982 | 0.00013117
1153.2 | 0.99982 | 0.00013289
1153.4 | 0.99982 | 0.00013437
1153.6 | 0.99982 | 0.0001356
1153.8 | 0.99982 | 0.00013656
1154 | 0.99981 | 0.00013723
1154.2 | 0.99981 | 0.00013761
1154.4 | 0.99981 | 0.00013768
1154.6 | 0.99981 | 0.00013745
1154.8 | 0.99981 | 0.00013691
1155 | 0.99982 | 0.00013608
1155.2 | 0.99982 | 0.00013497
1155.4 | 0.99982 | 0.00013359
1155.6 | 0.99982 | 0.00013196
1155.8 | 0.99982 | 0.0001301
1156 | 0.99983 | 0.00012803
1156.2 | 0.99983 | 0.00012579
1156.4 | 0.99983 | 0.00012339
1156.6 | 0.99984 | 0.00012087
1156.8 | 0.99984 | 0.00011824
1157 | 0.99984 | 0.00011554
1157.2 | 0.99985 | 0.00011278
1157.4 | 0.99985 | 0.00011
1157.6 | 0.99985 | 0.0001072
1157.8 | 0.99986 | 0.00010442
1158 | 0.99986 | 0.00010165
1158.2 | 0.99987 | 9.8926e-05
1158.4 | 0.99987 | 9.6248e-05
1158.6 | 0.99987 | 9.363e-05
1158.8 | 0.99988 | 9.108e-05
1159 | 0.99988 | 8.8603e-05
1159.2 | 0.99988 | 8.6206e-05
1159.4 | 0.99989 | 8.3891e-05
1159.6 | 0.99989 | 8.1661e-05
1159.8 | 0.99989 | 7.9518e-05
1160 | 0.9999 | 7.7463e-05
1160.2 | 0.9999 | 7.5496e-05
1160.4 | 0.9999 | 7.3615e-05
1160.6 | 0.9999 | 7.1821e-05
1160.8 | 0.99991 | 7.0111e-05
1161 | 0.99991 | 6.8483e-05
1161.2 | 0.99991 | 6.6937e-05
1161.4 | 0.99991 | 6.5469e-05
1161.6 | 0.99991 | 6.4078e-05
1161.8 | 0.99992 | 6.2761e-05
1162 | 0.99992 | 6.1515e-05
1162.2 | 0.99992 | 6.0339e-05
1162.4 | 0.99992 | 5.923e-05
1162.6 | 0.99992 | 5.8186e-05
1162.8 | 0.99992 | 5.7205e-05
1163 | 0.99993 | 5.6284e-05
1163.2 | 0.99993 | 5.5421e-05
1163.4 | 0.99993 | 5.4615e-05
1163.6 | 0.99993 | 5.3863e-05
1163.8 | 0.99993 | 5.3164e-05
1164 | 0.99993 | 5.2516e-05
1164.2 | 0.99993 | 5.1918e-05
1164.4 | 0.99993 | 5.1368e-05
1164.6 | 0.99993 | 5.0865e-05
1164.8 | 0.99993 | 5.0408e-05
1165 | 0.99993 | 4.9995e-05
1165.2 | 0.99994 | 4.9626e-05
1165.4 | 0.99994 | 4.9299e-05
1165.6 | 0.99994 | 4.9014e-05
1165.8 | 0.99994 | 4.8771e-05
1166 | 0.99994 | 4.8567e-05
1166.2 | 0.99994 | 4.8404e-05
1166.4 | 0.99994 | 4.828e-05
1166.6 | 0.99994 | 4.8196e-05
1166.8 | 0.99994 | 4.815e-05
1167 | 0.99994 | 4.8143e-05
1167.2 | 0.99994 | 4.8176e-05
1167.4 | 0.99994 | 4.8247e-05
1167.6 | 0.99994 | 4.8358e-05
1167.8 | 0.99994 | 4.8509e-05
1168 | 0.99994 | 4.8699e-05
1168.2 | 0.99994 | 4.8931e-05
1168.4 | 0.99994 | 4.9203e-05
1168.6 | 0.99994 | 4.9518e-05
1168.8 | 0.99994 | 4.9875e-05
1169 | 0.99994 | 5.0276e-05
1169.2 | 0.99994 | 5.0723e-05
1169.4 | 0.99994 | 5.1215e-05
1169.6 | 0.99994 | 5.1756e-05
1169.8 | 0.99994 | 5.2346e-05
1170 | 0.99993 | 5.2986e-05
1170.2 | 0.99993 | 5.368e-05
1170.4 | 0.99993 | 5.4429e-05
1170.6 | 0.99993 | 5.5235e-05
1170.8 | 0.99993 | 5.6101e-05
1171 | 0.99993 | 5.7029e-05
1171.2 | 0.99993 | 5.8023e-05
1171.4 | 0.99993 | 5.9086e-05
1171.6 | 0.99993 | 6.0221e-05
1171.8 | 0.99993 | 6.1433e-05
1172 | 0.99992 | 6.2724e-05
1172.2 | 0.99992 | 6.4101e-05
1172.4 | 0.99992 | 6.5567e-05
1172.6 | 0.99992 | 6.7128e-05
1172.8 | 0.99992 | 6.8789e-05
1173 | 0.99992 | 7.0558e-05
1173.2 | 0.99991 | 7.244e-05
1173.4 | 0.99991 | 7.4443e-05
1173.6 | 0.99991 | 7.6574e-05
1173.8 | 0.99991 | 7.8843e-05
1174 | 0.9999 | 8.1258e-05
1174.2 | 0.9999 | 8.3829e-05
1174.4 | 0.9999 | 8.6568e-05
1174.6 | 0.9999 | 8.9487e-05
1174.8 | 0.99989 | 9.2597e-05
1175 | 0.99989 | 9.5913e-05
1175.2 | 0.99988 | 9.945e-05
1175.4 | 0.99988 | 0.00010322
1175.6 | 0.99988 | 0.00010725
1175.8 | 0.99987 | 0.00011156
1176 | 0.99987 | 0.00011616
1176.2 | 0.99986 | 0.00012107
1176.4 | 0.99986 | 0.00012633
1176.6 | 0.99985 | 0.00013195
1176.8 | 0.99984 | 0.00013797
1177 | 0.99984 | 0.00014441
1177.2 | 0.99983 | 0.00015131
1177.4 | 0.99982 | 0.0001587
1177.6 | 0.99981 | 0.00016662
1177.8 | 0.9998 | 0.00017509
1178 | 0.99979 | 0.00018417
1178.2 | 0.99978 | 0.0001939
1178.4 | 0.99977 | 0.0002043
1178.6 | 0.99976 | 0.00021543
1178.8 | 0.99975 | 0.00022732
1179 | 0.99973 | 0.00024
1179.2 | 0.99972 | 0.00025351
1179.4 | 0.9997 | 0.00026788
1179.6 | 0.99969 | 0.00028313
1179.8 | 0.99967 | 0.00029925
1180 | 0.99965 | 0.00031624
1180.2 | 0.99963 | 0.00033408
1180.4 | 0.99961 | 0.00035271
1180.6 | 0.99959 | 0.00037207
1180.8 | 0.99957 | 0.00039205
1181 | 0.99955 | 0.00041252
1181.2 | 0.99953 | 0.00043332
1181.4 | 0.99951 | 0.00045425
1181.6 | 0.99948 | 0.00047509
1181.8 | 0.99946 | 0.00049557
1182 | 0.99944 | 0.00051544
1182.2 | 0.99942 | 0.00053441
1182.4 | 0.9994 | 0.0005522
1182.6 | 0.99939 | 0.00056853
1182.8 | 0.99937 | 0.00058315
1183 | 0.99936 | 0.00059585
1183.2 | 0.99935 | 0.00060643
1183.4 | 0.99934 | 0.00061477
1183.6 | 0.99933 | 0.00062076
1183.8 | 0.99933 | 0.00062437
1184 | 0.99933 | 0.0006256
1184.2 | 0.99933 | 0.00062452
1184.4 | 0.99934 | 0.00062122
1184.6 | 0.99934 | 0.00061585
1184.8 | 0.99935 | 0.0006086
1185 | 0.99936 | 0.00059967
1185.2 | 0.99937 | 0.0005893
1185.4 | 0.99939 | 0.00057772
1185.6 | 0.9994 | 0.00056519
1185.8 | 0.99941 | 0.00055194
1186 | 0.99943 | 0.00053819
1186.2 | 0.99944 | 0.00052414
1186.4 | 0.99946 | 0.00050999
1186.6 | 0.99948 | 0.00049589
1186.8 | 0.99949 | 0.00048197
1187 | 0.99951 | 0.00046836
1187.2 | 0.99952 | 0.00045513
1187.4 | 0.99953 | 0.00044236
1187.6 | 0.99955 | 0.0004301
1187.8 | 0.99956 | 0.00041838
1188 | 0.99957 | 0.00040724
1188.2 | 0.99958 | 0.00039668
1188.4 | 0.99959 | 0.00038672
1188.6 | 0.9996 | 0.00037734
1188.8 | 0.99961 | 0.00036854
1189 | 0.99962 | 0.00036032
1189.2 | 0.99963 | 0.00035266
1189.4 | 0.99964 | 0.00034554
1189.6 | 0.99965 | 0.00033895
1189.8 | 0.99965 | 0.00033287
1190 | 0.99966 | 0.00032729
1190.2 | 0.99966 | 0.00032218
1190.4 | 0.99967 | 0.00031754
1190.6 | 0.99967 | 0.00031335
1190.8 | 0.99968 | 0.0003096
1191 | 0.99968 | 0.00030626
1191.2 | 0.99968 | 0.00030334
1191.4 | 0.99969 | 0.00030082
1191.6 | 0.99969 | 0.00029868
1191.8 | 0.99969 | 0.00029693
1192 | 0.99969 | 0.00029556
1192.2 | 0.99969 | 0.00029456
1192.4 | 0.99969 | 0.00029392
1192.6 | 0.99969 | 0.00029365
1192.8 | 0.99969 | 0.00029375
1193 | 0.99969 | 0.00029421
1193.2 | 0.99969 | 0.00029503
1193.4 | 0.99969 | 0.00029622
1193.6 | 0.99969 | 0.00029779
1193.8 | 0.99969 | 0.00029974
1194 | 0.99969 | 0.00030207
1194.2 | 0.99968 | 0.00030481
1194.4 | 0.99968 | 0.00030795
1194.6 | 0.99968 | 0.00031151
1194.8 | 0.99967 | 0.00031552
1195 | 0.99967 | 0.00031997
1195.2 | 0.99966 | 0.00032491
1195.4 | 0.99966 | 0.00033033
1195.6 | 0.99965 | 0.00033628
1195.8 | 0.99965 | 0.00034278
1196 | 0.99964 | 0.00034985
1196.2 | 0.99963 | 0.00035753
1196.4 | 0.99962 | 0.00036586
1196.6 | 0.99961 | 0.00037487
1196.8 | 0.9996 | 0.00038461
1197 | 0.99959 | 0.00039513
1197.2 | 0.99958 | 0.00040647
1197.4 | 0.99957 | 0.0004187
1197.6 | 0.99955 | 0.00043188
1197.8 | 0.99954 | 0.00044607
1198 | 0.99952 | 0.00046134
1198.2 | 0.99951 | 0.00047778
1198.4 | 0.99949 | 0.00049546
1198.6 | 0.99947 | 0.00051448
1198.8 | 0.99945 | 0.00053492
1199 | 0.99943 | 0.00055689
1199.2 | 0.9994 | 0.00058049
1199.4 | 0.99938 | 0.00060582
1199.6 | 0.99935 | 0.000633
1199.8 | 0.99932 | 0.00066213
1200 | 0.99929 | 0.0006933
1200.2 | 0.99906 | 0.00093681
1200.4 | 0.99907 | 0.00092862
1200.6 | 0.99908 | 0.00092133
1200.8 | 0.99909 | 0.00091493
1201 | 0.99909 | 0.00090939
1201.2 | 0.9991 | 0.0009047
1201.4 | 0.9991 | 0.00090085
1201.6 | 0.9991 | 0.0008978
1201.8 | 0.9991 | 0.00089556
1202 | 0.99911 | 0.00089412
1202.2 | 0.99911 | 0.00089346
1202.4 | 0.99911 | 0.00089358
1202.6 | 0.99911 | 0.00089447
1202.8 | 0.9991 | 0.00089614
1203 | 0.9991 | 0.00089858
1203.2 | 0.9991 | 0.0009018
1203.4 | 0.99909 | 0.0009058
1203.6 | 0.99909 | 0.00091059
1203.8 | 0.99908 | 0.00091617
1204 | 0.99908 | 0.00092255
1204.2 | 0.99907 | 0.00092975
1204.4 | 0.99906 | 0.00093779
1204.6 | 0.99905 | 0.00094666
1204.8 | 0.99904 | 0.00095641
1205 | 0.99903 | 0.00096704
1205.2 | 0.99902 | 0.00097857
1205.4 | 0.99901 | 0.00099104
1205.6 | 0.999 | 0.0010045
1205.8 | 0.99898 | 0.0010189
1206 | 0.99897 | 0.0010343
1206.2 | 0.99895 | 0.0010508
1206.4 | 0.99893 | 0.0010683
1206.6 | 0.99891 | 0.001087
1206.8 | 0.99889 | 0.0011068
1207 | 0.99887 | 0.0011278
1207.2 | 0.99885 | 0.00115
1207.4 | 0.99883 | 0.0011735
1207.6 | 0.9988 | 0.0011984
1207.8 | 0.99878 | 0.0012246
1208 | 0.99875 | 0.0012522
1208.2 | 0.99872 | 0.0012813
1208.4 | 0.99869 | 0.0013119
1208.6 | 0.99866 | 0.001344
1208.8 | 0.99862 | 0.0013777
1209 | 0.99859 | 0.0014131
1209.2 | 0.99855 | 0.0014501
1209.4 | 0.99851 | 0.0014889
1209.6 | 0.99847 | 0.0015293
1209.8 | 0.99843 | 0.0015716
1210 | 0.99838 | 0.0016156
1210.2 | 0.99834 | 0.0016614
1210.4 | 0.99829 | 0.0017091
1210.6 | 0.99824 | 0.0017584
1210.8 | 0.99819 | 0.0018096
1211 | 0.99814 | 0.0018624
1211.2 | 0.99808 | 0.0019168
1211.4 | 0.99803 | 0.0019728
1211.6 | 0.99797 | 0.0020302
1211.8 | 0.99791 | 0.0020889
1212 | 0.99785 | 0.0021487
1212.2 | 0.99779 | 0.0022095
1212.4 | 0.99773 | 0.0022709
1212.6 | 0.99767 | 0.0023328
1212.8 | 0.99761 | 0.0023948
1213 | 0.99754 | 0.0024567
1213.2 | 0.99748 | 0.002518
1213.4 | 0.99742 | 0.0025784
1213.6 | 0.99736 | 0.0026375
1213.8 | 0.99731 | 0.0026949
1214 | 0.99725 | 0.0027502
1214.2 | 0.9972 | 0.0028029
1214.4 | 0.99715 | 0.0028527
1214.6 | 0.9971 | 0.0028992
1214.8 | 0.99706 | 0.002942
1215 | 0.99702 | 0.0029809
1215.2 | 0.99698 | 0.0030154
1215.4 | 0.99695 | 0.0030456
1215.6 | 0.99693 | 0.0030711
1215.8 | 0.99691 | 0.0030921
1216 | 0.99689 | 0.0031084
1216.2 | 0.99688 | 0.0031201
1216.4 | 0.99687 | 0.0031276
1216.6 | 0.99687 | 0.0031309
1216.8 | 0.99687 | 0.0031304
1217 | 0.99687 | 0.0031264
1217.2 | 0.99688 | 0.0031193
1217.4 | 0.99689 | 0.0031095
1217.6 | 0.9969 | 0.0030973
1217.8 | 0.99692 | 0.0030828
1218 | 0.99693 | 0.0030663
1218.2 | 0.99695 | 0.0030475
1218.4 | 0.99697 | 0.0030263
1218.6 | 0.997 | 0.0030021
1218.8 | 0.99703 | 0.0029744
1219 | 0.99706 | 0.0029428
1219.2 | 0.99709 | 0.0029069
1219.4 | 0.99713 | 0.0028669
1219.6 | 0.99718 | 0.0028233
1219.8 | 0.99722 | 0.002777
1220 | 0.99727 | 0.0027291
1220.2 | 0.99732 | 0.0026804
1220.4 | 0.99737 | 0.002632
1220.6 | 0.99742 | 0.0025845
1220.8 | 0.99746 | 0.0025385
1221 | 0.99751 | 0.0024943
1221.2 | 0.99755 | 0.0024521
1221.4 | 0.99759 | 0.0024121
1221.6 | 0.99763 | 0.0023743
1221.8 | 0.99766 | 0.0023386
1222 | 0.99769 | 0.0023052
1222.2 | 0.99773 | 0.002274
1222.4 | 0.99776 | 0.0022448
1222.6 | 0.99778 | 0.0022178
1222.8 | 0.99781 | 0.0021928
1223 | 0.99783 | 0.0021697
1223.2 | 0.99785 | 0.0021486
1223.4 | 0.99787 | 0.0021294
1223.6 | 0.99789 | 0.002112
1223.8 | 0.9979 | 0.0020965
1224 | 0.99792 | 0.0020827
1224.2 | 0.99793 | 0.0020708
1224.4 | 0.99794 | 0.0020606
1224.6 | 0.99795 | 0.0020522
1224.8 | 0.99795 | 0.0020455
1225 | 0.99796 | 0.0020405
1225.2 | 0.99796 | 0.0020373
1225.4 | 0.99796 | 0.0020357
1225.6 | 0.99796 | 0.002036
1225.8 | 0.99796 | 0.0020379
1226 | 0.99796 | 0.0020416
1226.2 | 0.99795 | 0.0020471
1226.4 | 0.99795 | 0.0020543
1226.6 | 0.99794 | 0.0020634
1226.8 | 0.99793 | 0.0020744
1227 | 0.99791 | 0.0020872
1227.2 | 0.9979 | 0.002102
1227.4 | 0.99788 | 0.0021188
1227.6 | 0.99786 | 0.0021376
1227.8 | 0.99784 | 0.0021585
1228 | 0.99782 | 0.0021816
1228.2 | 0.99779 | 0.002207
1228.4 | 0.99777 | 0.0022347
1228.6 | 0.99774 | 0.0022648
1228.8 | 0.9977 | 0.0022975
1229 | 0.99767 | 0.0023328
1229.2 | 0.99763 | 0.0023709
1229.4 | 0.99759 | 0.0024119
1229.6 | 0.99754 | 0.0024559
1229.8 | 0.9975 | 0.0025031
1230 | 0.99745 | 0.0025537
1230.2 | 0.99739 | 0.0026079
1230.4 | 0.99733 | 0.0026658
1230.6 | 0.99727 | 0.0027276
1230.8 | 0.99721 | 0.0027937
1231 | 0.99714 | 0.0028642
1231.2 | 0.99706 | 0.0029394
1231.4 | 0.99698 | 0.0030196
1231.6 | 0.99689 | 0.0031052
1231.8 | 0.9968 | 0.0031965
1232 | 0.99671 | 0.0032938
1232.2 | 0.9966 | 0.0033976
1232.4 | 0.99649 | 0.0035083
1232.6 | 0.99637 | 0.0036263
1232.8 | 0.99625 | 0.0037521
1233 | 0.99611 | 0.0038863
1233.2 | 0.99597 | 0.0040294
1233.4 | 0.99582 | 0.0041821
1233.6 | 0.99566 | 0.0043449
1233.8 | 0.99548 | 0.0045184
1234 | 0.9953 | 0.0047035
1234.2 | 0.9951 | 0.0049008
1234.4 | 0.99489 | 0.005111
1234.6 | 0.99467 | 0.005335
1234.8 | 0.99443 | 0.0055734
1235 | 0.99417 | 0.0058271
1235.2 | 0.9939 | 0.0060968
1235.4 | 0.99362 | 0.0063831
1235.6 | 0.99331 | 0.0066866
1235.8 | 0.99299 | 0.0070079
1236 | 0.99265 | 0.0073471
1236.2 | 0.9923 | 0.0077045
1236.4 | 0.99192 | 0.0080797
1236.6 | 0.99153 | 0.0084723
1236.8 | 0.99112 | 0.0088812
1237 | 0.9907 | 0.0093049
1237.2 | 0.99026 | 0.0097413
1237.4 | 0.98981 | 0.010188
1237.6 | 0.98936 | 0.01064
1237.8 | 0.98891 | 0.011095
1238 | 0.98845 | 0.011546
1238.2 | 0.98801 | 0.011988
1238.4 | 0.98759 | 0.012414
1238.6 | 0.98718 | 0.012817
1238.8 | 0.98681 | 0.013189
1239 | 0.98648 | 0.013523
1239.2 | 0.98619 | 0.013813
1239.4 | 0.98595 | 0.014051
1239.6 | 0.98577 | 0.014233
1239.8 | 0.98565 | 0.014355
1240 | 0.98559 | 0.014414
1240.2 | 0.98559 | 0.014412
1240.4 | 0.98565 | 0.014349
1240.6 | 0.98577 | 0.014228
1240.8 | 0.98595 | 0.014054
1241 | 0.98617 | 0.013832
1241.2 | 0.98643 | 0.013569
1241.4 | 0.98673 | 0.013271
1241.6 | 0.98705 | 0.012946
1241.8 | 0.9874 | 0.0126
1242 | 0.98776 | 0.012238
1242.2 | 0.98813 | 0.011868
1242.4 | 0.98851 | 0.011493
1242.6 | 0.98888 | 0.011118
1242.8 | 0.98925 | 0.010747
1243 | 0.98962 | 0.010382
1243.2 | 0.98997 | 0.010026
1243.4 | 0.99032 | 0.0096816
1243.6 | 0.99065 | 0.0093491
1243.8 | 0.99097 | 0.0090297
1244 | 0.99128 | 0.008724
1244.2 | 0.99157 | 0.0084325
1244.4 | 0.99184 | 0.0081551
1244.6 | 0.99211 | 0.0078919
1244.8 | 0.99236 | 0.0076426
1245 | 0.99259 | 0.0074069
1245.2 | 0.99282 | 0.0071843
1245.4 | 0.99303 | 0.0069745
1245.6 | 0.99322 | 0.006777
1245.8 | 0.99341 | 0.0065911
1246 | 0.99358 | 0.0064166
1246.2 | 0.99375 | 0.0062527
1246.4 | 0.9939 | 0.0060991
1246.6 | 0.99404 | 0.0059552
1246.8 | 0.99418 | 0.0058205
1247 | 0.99431 | 0.0056947
1247.2 | 0.99442 | 0.0055773
1247.4 | 0.99453 | 0.0054679
1247.6 | 0.99463 | 0.0053661
1247.8 | 0.99473 | 0.0052715
1248 | 0.99482 | 0.0051838
1248.2 | 0.9949 | 0.0051028
1248.4 | 0.99497 | 0.005028
1248.6 | 0.99504 | 0.0049594
1248.8 | 0.9951 | 0.0048965
1249 | 0.99516 | 0.0048392
1249.2 | 0.99521 | 0.0047873
1249.4 | 0.99526 | 0.0047406
1249.6 | 0.9953 | 0.0046989
1249.8 | 0.99534 | 0.004662
1250 | 0.99537 | 0.0046299
1250.2 | 0.9954 | 0.0046025
1250.4 | 0.99542 | 0.0045795
1250.6 | 0.99544 | 0.0045609
1250.8 | 0.99545 | 0.0045467
1251 | 0.99546 | 0.0045368
1251.2 | 0.99547 | 0.0045311
1251.4 | 0.99547 | 0.0045296
1251.6 | 0.99547 | 0.0045323
1251.8 | 0.99546 | 0.0045392
1252 | 0.99545 | 0.0045502
1252.2 | 0.99543 | 0.0045655
1252.4 | 0.99542 | 0.004585
1252.6 | 0.99539 | 0.0046087
1252.8 | 0.99536 | 0.0046368
1253 | 0.99533 | 0.0046693
1253.2 | 0.99529 | 0.0047063
1253.4 | 0.99525 | 0.0047479
1253.6 | 0.99521 | 0.0047941
1253.8 | 0.99515 | 0.0048453
1254 | 0.9951 | 0.0049014
1254.2 | 0.99504 | 0.0049627
1254.4 | 0.99497 | 0.0050293
1254.6 | 0.9949 | 0.0051014
1254.8 | 0.99482 | 0.0051794
1255 | 0.99474 | 0.0052633
1255.2 | 0.99465 | 0.0053535
1255.4 | 0.99455 | 0.0054503
1255.6 | 0.99445 | 0.005554
1255.8 | 0.99434 | 0.0056649
1256 | 0.99422 | 0.0057834
1256.2 | 0.99409 | 0.0059099
1256.4 | 0.99396 | 0.0060449
1256.6 | 0.99381 | 0.0061888
1256.8 | 0.99366 | 0.0063422
1257 | 0.99349 | 0.0065056
1257.2 | 0.99332 | 0.0066796
1257.4 | 0.99314 | 0.0068648
1257.6 | 0.99294 | 0.007062
1257.8 | 0.99273 | 0.007272
1258 | 0.9925 | 0.0074954
1258.2 | 0.99227 | 0.0077333
1258.4 | 0.99201 | 0.0079866
1258.6 | 0.99174 | 0.0082562
1258.8 | 0.99146 | 0.0085434
1259 | 0.99115 | 0.0088493
1259.2 | 0.99082 | 0.0091751
1259.4 | 0.99048 | 0.0095222
1259.6 | 0.99011 | 0.0098921
1259.8 | 0.98971 | 0.010286
1260 | 0.98929 | 0.010707
1260.2 | 0.98885 | 0.011154
1260.4 | 0.98837 | 0.011632
1260.6 | 0.98786 | 0.012141
1260.8 | 0.98732 | 0.012683
1261 | 0.98674 | 0.013261
1261.2 | 0.98612 | 0.013876
1261.4 | 0.98547 | 0.014531
1261.6 | 0.98477 | 0.015227
1261.8 | 0.98403 | 0.015967
1262 | 0.98325 | 0.016752
1262.2 | 0.98242 | 0.017584
1262.4 | 0.98154 | 0.018465
1262.6 | 0.98061 | 0.019394
1262.8 | 0.97963 | 0.020373
1263 | 0.9786 | 0.0214
1263.2 | 0.97752 | 0.022475
1263.4 | 0.9764 | 0.023595
1263.6 | 0.97524 | 0.024756
1263.8 | 0.97405 | 0.025953
1264 | 0.97282 | 0.027179
1264.2 | 0.97158 | 0.028425
1264.4 | 0.97032 | 0.02968
1264.6 | 0.96907 | 0.030932
1264.8 | 0.96783 | 0.032166
1265 | 0.96663 | 0.033366
1265.2 | 0.96549 | 0.034515
1265.4 | 0.96441 | 0.035595
1265.6 | 0.96341 | 0.036588
1265.8 | 0.96252 | 0.037477
1266 | 0.96175 | 0.038247
1266.2 | 0.96111 | 0.038886
1266.4 | 0.96062 | 0.039383
1266.6 | 0.96027 | 0.039733
1266.8 | 0.96007 | 0.039931
1267 | 0.96002 | 0.039978
1267.2 | 0.96013 | 0.039875
1267.4 | 0.96037 | 0.039626
1267.6 | 0.96076 | 0.039237
1267.8 | 0.96129 | 0.038714
1268 | 0.96193 | 0.03807
1268.2 | 0.96268 | 0.037318
1268.4 | 0.96352 | 0.036477
1268.6 | 0.96443 | 0.035566
1268.8 | 0.96539 | 0.034607
1269 | 0.96638 | 0.033619
1269.2 | 0.96738 | 0.03262
1269.4 | 0.96838 | 0.031622
1269.6 | 0.96936 | 0.030638
1269.8 | 0.97032 | 0.029676
1270 | 0.97126 | 0.02874
1270.2 | 0.97216 | 0.027837
1270.4 | 0.97303 | 0.026968
1270.6 | 0.97386 | 0.026135
1270.8 | 0.97466 | 0.02534
1271 | 0.97542 | 0.024584
1271.2 | 0.97614 | 0.023864
1271.4 | 0.97682 | 0.023183
1271.6 | 0.97746 | 0.022537
1271.8 | 0.97807 | 0.021928
1272 | 0.97865 | 0.021353
1272.2 | 0.97919 | 0.020811
1272.4 | 0.9797 | 0.020302
1272.6 | 0.98018 | 0.019823
1272.8 | 0.98063 | 0.019374
1273 | 0.98105 | 0.018953
1273.2 | 0.98144 | 0.018559
1273.4 | 0.98181 | 0.018191
1273.6 | 0.98215 | 0.017848
1273.8 | 0.98247 | 0.017528
1274 | 0.98277 | 0.017231
1274.2 | 0.98304 | 0.016956
1274.4 | 0.9833 | 0.016701
1274.6 | 0.98353 | 0.016466
1274.8 | 0.98375 | 0.01625
1275 | 0.98395 | 0.016053
1275.2 | 0.98413 | 0.015873
1275.4 | 0.98429 | 0.01571
1275.6 | 0.98444 | 0.015564
1275.8 | 0.98457 | 0.015433
1276 | 0.98468 | 0.015318
1276.2 | 0.98478 | 0.015218
1276.4 | 0.98487 | 0.015133
1276.6 | 0.98494 | 0.015062
1276.8 | 0.98499 | 0.015005
1277 | 0.98504 | 0.014962
1277.2 | 0.98507 | 0.014932
1277.4 | 0.98508 | 0.014916
1277.6 | 0.98509 | 0.014913
1277.8 | 0.98508 | 0.014923
1278 | 0.98505 | 0.014946
1278.2 | 0.98502 | 0.014983
1278.4 | 0.98497 | 0.015032
1278.6 | 0.98491 | 0.015095
1278.8 | 0.98483 | 0.015171
1279 | 0.98474 | 0.015261
1279.2 | 0.98464 | 0.015364
1279.4 | 0.98452 | 0.015481
1279.6 | 0.98439 | 0.015612
1279.8 | 0.98424 | 0.015757
1280 | 0.98408 | 0.015918
1280.2 | 0.98391 | 0.016093
1280.4 | 0.98372 | 0.016284
1280.6 | 0.98351 | 0.016491
1280.8 | 0.98329 | 0.016714
1281 | 0.98305 | 0.016955
1281.2 | 0.98279 | 0.017213
1281.4 | 0.98251 | 0.01749
1281.6 | 0.98221 | 0.017786
1281.8 | 0.9819 | 0.018101
1282 | 0.98156 | 0.018438
1282.2 | 0.9812 | 0.018796
1282.4 | 0.98082 | 0.019176
1282.6 | 0.98042 | 0.019581
1282.8 | 0.97999 | 0.02001
1283 | 0.97953 | 0.020465
1283.2 | 0.97905 | 0.020948
1283.4 | 0.97854 | 0.02146
1283.6 | 0.978 | 0.022002
1283.8 | 0.97742 | 0.022576
1284 | 0.97682 | 0.023183
1284.2 | 0.97617 | 0.023826
1284.4 | 0.97549 | 0.024506
1284.6 | 0.97477 | 0.025225
1284.8 | 0.97401 | 0.025986
1285 | 0.97321 | 0.026791
1285.2 | 0.97236 | 0.027642
1285.4 | 0.97146 | 0.028542
1285.6 | 0.97051 | 0.029493
1285.8 | 0.9695 | 0.030498
1286 | 0.96844 | 0.03156
1286.2 | 0.96732 | 0.032682
1286.4 | 0.96613 | 0.033867
1286.6 | 0.96488 | 0.035119
1286.8 | 0.96356 | 0.03644
1287 | 0.96217 | 0.037834
1287.2 | 0.9607 | 0.039305
1287.4 | 0.95915 | 0.040855
1287.6 | 0.95751 | 0.042488
1287.8 | 0.95579 | 0.044207
1288 | 0.95398 | 0.046016
1288.2 | 0.95208 | 0.047916
1288.4 | 0.95009 | 0.04991
1288.6 | 0.948 | 0.052
1288.8 | 0.94581 | 0.054187
1289 | 0.94353 | 0.056472
1289.2 | 0.94115 | 0.058855
1289.4 | 0.93867 | 0.061333
1289.6 | 0.9361 | 0.063904
1289.8 | 0.93344 | 0.066565
1290 | 0.93069 | 0.069309
1290.2 | 0.92787 | 0.072129
1290.4 | 0.92498 | 0.075016
1290.6 | 0.92204 | 0.077959
1290.8 | 0.91906 | 0.080944
1291 | 0.91604 | 0.083956
1291.2 | 0.91302 | 0.086978
1291.4 | 0.91001 | 0.089991
1291.6 | 0.90703 | 0.092973
1291.8 | 0.9041 | 0.095904
1292 | 0.90124 | 0.09876
1292.2 | 0.89848 | 0.10152
1292.4 | 0.89584 | 0.10416
1292.6 | 0.89335 | 0.10665
1292.8 | 0.89102 | 0.10898
1293 | 0.88887 | 0.11113
1293.2 | 0.88692 | 0.11308
1293.4 | 0.88518 | 0.11482
1293.6 | 0.88367 | 0.11633
1293.8 | 0.8824 | 0.1176
1294 | 0.88138 | 0.11862
1294.2 | 0.88062 | 0.11938
1294.4 | 0.88012 | 0.11988
1294.6 | 0.87989 | 0.12011
1294.8 | 0.87993 | 0.12007
1295 | 0.88023 | 0.11977
1295.2 | 0.88077 | 0.11923
1295.4 | 0.88153 | 0.11847
1295.6 | 0.88248 | 0.11752
1295.8 | 0.88359 | 0.11641
1296 | 0.88483 | 0.11517
1296.2 | 0.88616 | 0.11384
1296.4 | 0.88756 | 0.11244
1296.6 | 0.88901 | 0.11099
1296.8 | 0.89048 | 0.10952
1297 | 0.89196 | 0.10804
1297.2 | 0.89343 | 0.10657
1297.4 | 0.89489 | 0.10511
1297.6 | 0.89632 | 0.10368
1297.8 | 0.89771 | 0.10229
1298 | 0.89905 | 0.10095
1298.2 | 0.90035 | 0.099646
1298.4 | 0.9016 | 0.098401
1298.6 | 0.90279 | 0.097213
1298.8 | 0.90391 | 0.096085
1299 | 0.90498 | 0.095021
1299.2 | 0.90598 | 0.094022
1299.4 | 0.90691 | 0.09309
1299.6 | 0.90777 | 0.092225
1299.8 | 0.90857 | 0.09143
1300 | 0.9093 | 0.090704

## UV Filters (hidden)
Wavelength (nm) | ND 0.5 % Transmission | ND 1 % Transmission | ND 2 % Transmission | ND 3 % Transmission | ND 4 % Transmission
200 | 35.0713 | 9.934 | 0.52524 | 0.033132 | 0.00224 | First | Transmission | 1
210 | 33.6708 | 9.069 | 0.47236 | 0.027696 | 0.00175 | Second | Transmission | 1
220 | 31.8873 | 8.1726 | 0.39637 | 0.021731 | 0.00134 | Third | Transmission | 1
230 | 30.1734 | 7.4937 | 0.34958 | 0.01885 | 0.001153
240 | 28.8073 | 7.1144 | 0.34126 | 0.018963 | 0.001195
250 | 28.0013 | 7.0744 | 0.37294 | 0.022256 | 0.001449
260 | 27.7911 | 7.3524 | 0.44033 | 0.029366 | 0.002055
270 | 28.1205 | 7.9003 | 0.54975 | 0.041474 | 0.003222
280 | 28.8781 | 8.643 | 0.70135 | 0.059896 | 0.005208
290 | 29.8996 | 9.4892 | 0.88471 | 0.084746 | 0.008218
300 | 31.0417 | 10.3583 | 1.088 | 0.11495 | 0.012269
310 | 32.1877 | 11.1762 | 1.2942 | 0.14781 | 0.017172
320 | 33.2294 | 11.8692 | 1.4737 | 0.17908 | 0.022168
330 | 34.1041 | 12.4116 | 1.6145 | 0.20534 | 0.026605
340 | 34.7864 | 12.7975 | 1.7148 | 0.22488 | 0.030042
350 | 35.2728 | 13.0334 | 1.7735 | 0.23785 | 0.032401
360 | 35.589 | 13.1547 | 1.8051 | 0.24557 | 0.033922
370 | 35.7887 | 13.1874 | 1.8171 | 0.24853 | 0.034513
380 | 35.8477 | 13.1407 | 1.8087 | 0.24774 | 0.034417
390 | 35.8042 | 13.0428 | 1.7919 | 0.24482 | 0.034021
400 | 35.709 | 12.9106 | 1.7675 | 0.24106 | 0.033476
410 | 35.5456 | 12.7462 | 1.7336 | 0.23679 | 0.032849
420 | 35.3526 | 12.5697 | 1.7014 | 0.23181 | 0.032182
430 | 35.1104 | 12.3827 | 1.6706 | 0.22728 | 0.031588
440 | 34.8234 | 12.1835 | 1.6382 | 0.22287 | 0.03104
450 | 34.5407 | 11.9891 | 1.6109 | 0.21835 | 0.030511
460 | 34.2544 | 11.8017 | 1.5855 | 0.21413 | 0.029994
470 | 33.9339 | 11.6071 | 1.5541 | 0.21043 | 0.029547
480 | 33.6298 | 11.4253 | 1.5294 | 0.2083 | 0.029412
490 | 33.3371 | 11.2582 | 1.514 | 0.20661 | 0.029341
500 | 33.0108 | 11.0849 | 1.495 | 0.20423 | 0.029137
510 | 32.7144 | 10.925 | 1.476 | 0.20215 | 0.029079
520 | 32.4342 | 10.7781 | 1.4595 | 0.20084 | 0.029098
530 | 32.1291 | 10.6267 | 1.4414 | 0.19959 | 0.029056
540 | 31.8466 | 10.4836 | 1.4273 | 0.1988 | 0.02912
550 | 31.5677 | 10.3441 | 1.417 | 0.19785 | 0.029211
560 | 31.2826 | 10.2054 | 1.4039 | 0.19678 | 0.029291
570 | 31.0073 | 10.0715 | 1.3887 | 0.19617 | 0.029388
580 | 30.7377 | 9.943 | 1.3771 | 0.19527 | 0.02948
590 | 30.4789 | 9.8194 | 1.3659 | 0.1946 | 0.029586
600 | 30.2113 | 9.694 | 1.3521 | 0.194 | 0.029667
610 | 29.956 | 9.5764 | 1.3416 | 0.19343 | 0.029832
620 | 29.7179 | 9.466 | 1.3334 | 0.19323 | 0.030045
630 | 29.4619 | 9.3535 | 1.3221 | 0.19265 | 0.03019
640 | 29.2278 | 9.2492 | 1.3132 | 0.1925 | 0.030402
650 | 29.0117 | 9.1522 | 1.3081 | 0.19286 | 0.030733
660 | 28.7853 | 9.0557 | 1.3016 | 0.19308 | 0.030999
670 | 28.5815 | 8.9694 | 1.2974 | 0.19394 | 0.031373
680 | 28.3742 | 8.8833 | 1.2955 | 0.1947 | 0.031736
690 | 28.1714 | 8.8035 | 1.2923 | 0.19563 | 0.032117
700 | 28.0041 | 8.7354 | 1.2916 | 0.19693 | 0.032619
710 | 27.8339 | 8.6684 | 1.2926 | 0.19854 | 0.03322
720 | 27.6566 | 8.6009 | 1.2908 | 0.20017 | 0.033866
730 | 27.5106 | 8.5448 | 1.292 | 0.20257 | 0.034612
740 | 27.3735 | 8.4983 | 1.296 | 0.20519 | 0.03543
750 | 27.232 | 8.4506 | 1.2982 | 0.20731 | 0.036181
760 | 27.1108 | 8.4119 | 1.3042 | 0.21024 | 0.037028
770 | 27.0041 | 8.3831 | 1.3124 | 0.21356 | 0.037905
780 | 26.8959 | 8.3541 | 1.3174 | 0.21616 | 0.038827
790 | 26.8079 | 8.3307 | 1.3264 | 0.22006 | 0.039905
800 | 26.7183 | 8.3103 | 1.3365 | 0.22435 | 0.040986
810 | 26.6393 | 8.2929 | 1.3463 | 0.22805 | 0.042119
820 | 26.5796 | 8.2845 | 1.3556 | 0.23174 | 0.043194
830 | 26.5028 | 8.2692 | 1.3622 | 0.2362 | 0.044642
840 | 26.4455 | 8.2603 | 1.3733 | 0.24091 | 0.046458
850 | 26.489 | 8.2923 | 1.3718 | 0.23626 | 0.045582
860 | 26.587 | 8.3977 | 1.4535 | 0.25455 | 0.049763
870 | 26.652 | 8.5008 | 1.5974 | 0.29246 | 0.058745
880 | 26.6876 | 8.5566 | 1.699 | 0.3146 | 0.063908
890 | 26.6583 | 8.5808 | 1.7179 | 0.32001 | 0.064957
900 | 26.5876 | 8.5527 | 1.6852 | 0.31601 | 0.064366
910 | 26.5089 | 8.4987 | 1.6534 | 0.30997 | 0.063747
920 | 26.4201 | 8.4286 | 1.6088 | 0.29656 | 0.061335
930 | 26.3775 | 8.4124 | 1.6129 | 0.29731 | 0.062003
940 | 26.338 | 8.4179 | 1.6333 | 0.30575 | 0.064309
950 | 26.2584 | 8.4008 | 1.6216 | 0.30819 | 0.065191
960 | 26.1765 | 8.3712 | 1.5962 | 0.30831 | 0.065555
970 | 26.138 | 8.3321 | 1.572 | 0.30222 | 0.064782
980 | 26.1003 | 8.305 | 1.5748 | 0.30128 | 0.065006
990 | 26.0381 | 8.2738 | 1.5663 | 0.30063 | 0.065227
1000 | 25.9706 | 8.2345 | 1.5462 | 0.29752 | 0.064962
1010 | 25.9113 | 8.2313 | 1.5424 | 0.29655 | 0.065101
1020 | 25.8591 | 8.239 | 1.5478 | 0.2979 | 0.065863
1030 | 25.8068 | 8.2047 | 1.5516 | 0.29986 | 0.066753
1040 | 25.7331 | 8.1463 | 1.5541 | 0.29446 | 0.065938
1050 | 25.6525 | 8.0986 | 1.5319 | 0.29284 | 0.065985
1060 | 25.6181 | 8.0938 | 1.5204 | 0.30717 | 0.069598
1070 | 25.6368 | 8.1459 | 1.555 | 0.33061 | 0.075296
1080 | 25.6399 | 8.1848 | 1.5931 | 0.33705 | 0.077169
1090 | 25.5532 | 8.1228 | 1.574 | 0.31499 | 0.072533
1100 | 25.4249 | 8.017 | 1.5058 | 0.28913 | 0.066862
1110 | 25.3252 | 7.9429 | 1.456 | 0.27278 | 0.063409
1120 | 25.2931 | 7.9502 | 1.4796 | 0.29013 | 0.067782
1130 | 25.3236 | 8.0325 | 1.5836 | 0.34463 | 0.080919
1140 | 25.2866 | 8.0389 | 1.6037 | 0.36411 | 0.085885
1150 | 25.1991 | 7.9922 | 1.5723 | 0.345 | 0.08171
1160 | 25.1591 | 7.9537 | 1.5495 | 0.3393 | 0.080606
1170 | 25.1106 | 7.9118 | 1.5178 | 0.33694 | 0.080339
1180 | 25.0433 | 7.8947 | 1.5432 | 0.32267 | 0.077301
1190 | 25.0698 | 7.9696 | 1.6352 | 0.34649 | 0.083374
1200 | 25.0055 | 7.9378 | 1.5994 | 0.34678 | 0.083837
1210 | 24.8926 | 7.8454 | 1.522 | 0.32275 | 0.078352
1220 | 24.8694 | 7.8599 | 1.5547 | 0.33648 | 0.0819
1230 | 24.8572 | 7.8932 | 1.6065 | 0.35774 | 0.087359
1240 | 24.7453 | 7.7992 | 1.5381 | 0.32829 | 0.080464
1250 | 24.6174 | 7.6719 | 1.4326 | 0.27955 | 0.068721
1260 | 24.5917 | 7.6826 | 1.4544 | 0.28356 | 0.069908
1270 | 24.6117 | 7.7318 | 1.5574 | 0.32205 | 0.079605
1280 | 24.5615 | 7.7373 | 1.5724 | 0.34376 | 0.08525
1290 | 24.5042 | 7.7245 | 1.5616 | 0.34843 | 0.086677
1300 | 24.588 | 7.8386 | 1.7019 | 0.39785 | 0.099235
1310 | 24.7002 | 8.0022 | 1.877 | 0.46517 | 0.1165
1320 | 24.5944 | 7.939 | 1.813 | 0.4361 | 0.10949
1330 | 24.1925 | 7.5484 | 1.4634 | 0.29581 | 0.074507
1340 | 23.8011 | 7.1429 | 1.1111 | 0.17817 | 0.045009
1350 | 23.8645 | 7.1886 | 1.2156 | 0.21191 | 0.053687
1360 | 24.0238 | 7.3875 | 1.4483 | 0.30188 | 0.076759
1370 | 23.7538 | 7.1742 | 1.2156 | 0.21917 | 0.055818
1380 | 23.9047 | 7.4581 | 1.5058 | 0.29567 | 0.075405
1390 | 23.85 | 7.4543 | 1.5347 | 0.33474 | 0.085593
1400 | 23.7717 | 7.3321 | 1.4808 | 0.32075 | 0.082281
1410 | 23.8372 | 7.3814 | 1.5976 | 0.36856 | 0.095036
1420 | 23.7736 | 7.3433 | 1.5928 | 0.37574 | 0.097279
1430 | 23.623 | 7.2863 | 1.4994 | 0.3443 | 0.089339
1440 | 23.4935 | 7.2247 | 1.4132 | 0.31776 | 0.08262
1450 | 23.4124 | 7.1843 | 1.3814 | 0.30299 | 0.078837
1460 | 23.4223 | 7.2417 | 1.458 | 0.32477 | 0.084583
1470 | 23.4263 | 7.2973 | 1.5418 | 0.35319 | 0.092165
1480 | 23.3388 | 7.2479 | 1.4921 | 0.33949 | 0.088858
1490 | 23.2198 | 7.1252 | 1.418 | 0.31819 | 0.08348
1500 | 23.1175 | 7.0376 | 1.3709 | 0.30098 | 0.079046
1510 | 23.073 | 7.0344 | 1.3818 | 0.30399 | 0.080214
1520 | 23.0699 | 7.0686 | 1.4499 | 0.32686 | 0.086383
1530 | 23.0309 | 7.0667 | 1.4677 | 0.33691 | 0.089191
1540 | 22.9585 | 7.0445 | 1.4411 | 0.33702 | 0.089413
1550 | 22.8404 | 6.9755 | 1.3805 | 0.319 | 0.084822
1560 | 22.7155 | 6.8758 | 1.3305 | 0.29703 | 0.079133
1570 | 22.6317 | 6.8227 | 1.3245 | 0.28704 | 0.076602
1580 | 22.6164 | 6.8554 | 1.3435 | 0.30887 | 0.082715
1590 | 22.6289 | 6.9036 | 1.4054 | 0.33996 | 0.09121
1600 | 22.6122 | 6.9355 | 1.4502 | 0.35919 | 0.096643
1610 | 22.508 | 6.8811 | 1.4068 | 0.35313 | 0.095292
1620 | 22.3573 | 6.727 | 1.2892 | 0.32452 | 0.087745
1630 | 22.2231 | 6.6216 | 1.2074 | 0.30122 | 0.081648
1640 | 22.2171 | 6.6583 | 1.2679 | 0.31446 | 0.085142
1650 | 22.2219 | 6.709 | 1.3358 | 0.34467 | 0.093465
1660 | 22.2005 | 6.714 | 1.3433 | 0.32685 | 0.088789
1670 | 22.1454 | 6.6592 | 1.3265 | 0.31207 | 0.085221
1680 | 22.0678 | 6.5998 | 1.3025 | 0.30611 | 0.083614
1690 | 22.0085 | 6.624 | 1.3492 | 0.31985 | 0.087436
1700 | 22.004 | 6.6801 | 1.4236 | 0.32896 | 0.089849
1710 | 21.9544 | 6.641 | 1.3978 | 0.32563 | 0.089136
1720 | 21.8575 | 6.5507 | 1.2881 | 0.30737 | 0.084717
1730 | 21.7222 | 6.4351 | 1.165 | 0.27736 | 0.076614
1740 | 21.5983 | 6.3368 | 1.1188 | 0.2482 | 0.068584
1750 | 21.626 | 6.4056 | 1.2314 | 0.28233 | 0.078448
1760 | 21.8339 | 6.7064 | 1.5557 | 0.37687 | 0.10496
1770 | 22.1081 | 7.0653 | 1.9354 | 0.50476 | 0.14046
1780 | 21.9467 | 6.931 | 1.8157 | 0.46265 | 0.1289
1790 | 21.5724 | 6.5447 | 1.4101 | 0.33675 | 0.093876
1800 | 21.3298 | 6.292 | 1.1577 | 0.26537 | 0.074137
1810 | 21.2372 | 6.1854 | 1.0944 | 0.2451 | 0.068653
1820 | 21.1925 | 6.1918 | 1.1366 | 0.24997 | 0.070149
1830 | 21.2213 | 6.2314 | 1.2201 | 0.27646 | 0.077727
1840 | 21.2991 | 6.4258 | 1.3828 | 0.3281 | 0.092307
1850 | 21.1699 | 6.2699 | 1.2264 | 0.29057 | 0.082282
1860 | 21.1086 | 6.2554 | 1.2129 | 0.28977 | 0.082051
1870 | 21.7559 | 6.9944 | 1.9722 | 0.5079 | 0.14385
1880 | 20.8312 | 6.0026 | 1.0016 | 0.14477 | 0.041065
1890 | 20.391 | 5.5483 | 0.56689 | 0.054382 | 0.015469
1900 | 20.4317 | 5.6342 | 0.68304 | 0.091052 | 0.025913
1910 | 20.5974 | 5.833 | 0.9209 | 0.1745 | 0.049784
1920 | 21.0771 | 6.4457 | 1.5314 | 0.35939 | 0.10282
1930 | 21.2243 | 6.622 | 1.7504 | 0.46566 | 0.13328
1940 | 21.1719 | 6.6031 | 1.7316 | 0.46063 | 0.13282
1950 | 21.0889 | 6.5304 | 1.6595 | 0.44444 | 0.1279
1960 | 20.998 | 6.4522 | 1.5788 | 0.42397 | 0.12189
1970 | 20.8961 | 6.3656 | 1.485 | 0.39499 | 0.11361
1980 | 20.7417 | 6.2274 | 1.3669 | 0.35488 | 0.10197
1990 | 20.6305 | 6.1318 | 1.2849 | 0.32186 | 0.092762
2000 | 20.584 | 6.1055 | 1.2631 | 0.30987 | 0.089543
2010 | 20.5053 | 6.0794 | 1.2441 | 0.30636 | 0.088357
2020 | 20.4532 | 6.0213 | 1.2305 | 0.3058 | 0.088345
2030 | 20.4104 | 6.0013 | 1.2289 | 0.30533 | 0.088414
2040 | 20.365 | 5.9858 | 1.2288 | 0.30504 | 0.088384
2050 | 20.3365 | 5.978 | 1.2526 | 0.31341 | 0.09098
2060 | 20.3101 | 5.9979 | 1.2565 | 0.31837 | 0.092548
2070 | 20.2889 | 5.9929 | 1.2645 | 0.31788 | 0.092592
2080 | 20.2473 | 5.9763 | 1.2595 | 0.31421 | 0.091538
2090 | 20.1919 | 5.9361 | 1.2249 | 0.31064 | 0.090133
2100 | 20.138 | 5.8906 | 1.2048 | 0.30694 | 0.089707
2110 | 20.0796 | 5.8659 | 1.184 | 0.30164 | 0.088329
2120 | 20.0433 | 5.8629 | 1.1871 | 0.30203 | 0.088632
2130 | 20.0088 | 5.8778 | 1.2184 | 0.31589 | 0.092613
2140 | 20.0169 | 5.9058 | 1.2619 | 0.33187 | 0.097392
2150 | 20.0517 | 5.9457 | 1.354 | 0.34883 | 0.10237
2160 | 20.0132 | 5.9073 | 1.3573 | 0.33274 | 0.098026
2170 | 19.8629 | 5.7241 | 1.2051 | 0.30238 | 0.088958
2180 | 19.6725 | 5.5585 | 1.0337 | 0.25007 | 0.073859
2190 | 19.5073 | 5.4304 | 0.90183 | 0.22328 | 0.065828
2200 | 19.5709 | 5.4296 | 1.0057 | 0.25531 | 0.075108
2210 | 19.6824 | 5.4674 | 1.1568 | 0.2969 | 0.087445
2220 | 19.652 | 5.4524 | 1.1607 | 0.29704 | 0.088582
2230 | 19.5749 | 5.4085 | 1.1381 | 0.28882 | 0.086503
2240 | 19.4795 | 5.4075 | 1.1339 | 0.2918 | 0.08717
2250 | 19.4688 | 5.4781 | 1.1654 | 0.29683 | 0.088941
2260 | 19.4923 | 5.5027 | 1.1811 | 0.30046 | 0.088333
2270 | 19.4702 | 5.5307 | 1.1799 | 0.29981 | 0.089268
2280 | 19.4243 | 5.535 | 1.1609 | 0.29651 | 0.088434
2290 | 19.3796 | 5.5435 | 1.1274 | 0.29731 | 0.088759
2300 | 19.3423 | 5.5441 | 1.1539 | 0.30704 | 0.092939
2310 | 19.3015 | 5.5316 | 1.1805 | 0.30793 | 0.093079
2320 | 19.25 | 5.5122 | 1.1747 | 0.30432 | 0.091714
2330 | 19.197 | 5.51 | 1.1541 | 0.29765 | 0.090536
2340 | 19.1605 | 5.5192 | 1.1357 | 0.29883 | 0.089862
2350 | 19.16 | 5.5246 | 1.161 | 0.30425 | 0.092986
2360 | 19.1209 | 5.5015 | 1.1521 | 0.30538 | 0.093439
2370 | 19.0704 | 5.4832 | 1.1649 | 0.31016 | 0.095358
2380 | 19.0436 | 5.4722 | 1.1558 | 0.32268 | 0.098073
2390 | 19.0213 | 5.467 | 1.1522 | 0.3293 | 0.1007
2400 | 18.9847 | 5.455 | 1.1496 | 0.33459 | 0.10122
2410 | 18.9537 | 5.4402 | 1.1605 | 0.33013 | 0.099676
2420 | 18.9256 | 5.4208 | 1.1591 | 0.32433 | 0.097321
2430 | 18.8801 | 5.3788 | 1.1291 | 0.31244 | 0.095742
2440 | 18.8458 | 5.3662 | 1.1285 | 0.33056 | 0.099773
2450 | 18.8648 | 5.4224 | 1.2105 | 0.39311 | 0.1228
2460 | 18.8159 | 5.3974 | 1.2025 | 0.45671 | 0.14029
2470 | 18.8553 | 5.4557 | 1.2755 | 0.62464 | 0.19334
2480 | 18.8569 | 5.4359 | 1.3073 | 0.74217 | 0.22806
2490 | 18.8792 | 5.4776 | 1.3401 | 0.78429 | 0.24022
2500 | 18.8016 | 5.3931 | 1.2675 | 0.67738 | 0.20781
2510 | 18.7757 | 5.3568 | 1.2509 | 0.51914 | 0.16044
2520 | 18.7693 | 5.3452 | 1.2674 | 0.38454 | 0.11839
2530 | 18.9141 | 5.5067 | 1.4528 | 0.17972 | 0.05406
2540 | 18.8511 | 5.4416 | 1.4272 | 0.5047 | 0.7086
2550 | 18.6148 | 5.2558 | 1.31 | 0.1263 | 0.0925
2560 | 18.2985 | 4.9309 | 1.0198 | 0.5572 | 0.3879
2570 | 17.516 | 4.0027 | 0.11044 | 0.6647 | 0.9786
2580 | 17.6344 | 4.1506 | 0.30021 | 0.0368 | 0.01124
2590 | 17.9291 | 4.4992 | 0.66942 | 0.1599 | 0.050154
2600 | 18.1281 | 4.7334 | 0.92419 | 0.24733 | 0.077168

## IR Filters (hidden)
|  |  | ND0.3 |  | ND0.6 |  | ND1 |  | ND2 |  | ND3
First | Transmission | 1 | Wavelength (nm) | Tx. (%) | Wavelength (nm) | Tx. (%) | Wavelength (nm) | Tx. (%) | Wavelength (nm) | Tx. (%) | Wavelength (nm) | Tx. (%)
Second | Transmission | 1 | 2000 | 46.23 | 2000 | 24.992 | 2000 | 9.9436 | 2000 | 0.9835 | 2000 | 0.0643
Third | Transmission | 1 | 2050 | 46.235 | 2005 | 25.0214 | 2005.64 | 9.9378 | 2005 | 0.98724 | 2028.06 | 0.0616
|  |  | 2100 | 46.22 | 2010 | 25.0248 | 2011.28 | 9.9498 | 2010 | 0.98778 | 2056.12 | 0.0769
|  |  | 2150 | 46.2001 | 2015 | 25.025 | 2016.92 | 9.9567 | 2015 | 0.98834 | 2084.18 | 0.0691
|  |  | 2200 | 46.245 | 2020 | 25.032 | 2022.56 | 9.9599 | 2020 | 0.98909 | 2112.24 | 0.072006
|  |  | 2250 | 46.24 | 2025 | 25.0441 | 2028.2 | 9.966 | 2025 | 0.98964 | 2140.3 | 0.072059
|  |  | 2300 | 46.25 | 2030 | 25.0568 | 2033.84 | 9.9696 | 2030 | 0.99017 | 2168.36 | 0.070005
| 2 | #N/A | 2350 | 46.265 | 2035 | 25.0665 | 2039.48 | 9.9732 | 2035 | 0.99103 | 2196.42 | 0.070163
| 3 | #N/A | 2400 | 46.26 | 2040 | 25.0698 | 2045.12 | 9.9768 | 2040 | 0.99198 | 2224.48 | 0.069258
| 4 | #N/A | 2450 | 46.225 | 2045 | 25.0753 | 2050.76 | 9.9799 | 2045 | 0.99306 | 2252.54 | 0.068894
| 5 | #N/A | 2500 | 46.245 | 2050 | 25.0618 | 2056.4 | 9.983 | 2050 | 0.99423 | 2280.6 | 0.069022
| 6 | #N/A | 2550 | 46.235 | 2055 | 25.0828 | 2062.04 | 9.9861 | 2055 | 0.99533 | 2308.66 | 0.069225
|  |  | 2600 | 46.1802 | 2060 | 25.0841 | 2067.68 | 9.9891 | 2060 | 0.99633 | 2336.72 | 0.069799
|  |  | 2650 | 46.2449 | 2065 | 25.0814 | 2073.32 | 9.9922 | 2065 | 0.99721 | 2364.78 | 0.070493
|  |  | 2700 | 46.2449 | 2070 | 25.0852 | 2078.96 | 9.9954 | 2070 | 0.99799 | 2392.84 | 0.071224
|  |  | 2750 | 46.2152 | 2075 | 25.0936 | 2084.6 | 9.9986 | 2075 | 0.9987 | 2420.9 | 0.071937
|  |  | 2800 | 46.2152 | 2080 | 25.0978 | 2090.24 | 10.0018 | 2080 | 0.99938 | 2448.96 | 0.072522
|  |  | 2850 | 46.2496 | 2085 | 25.0904 | 2095.88 | 10.0051 | 2085 | 1.0001 | 2477.02 | 0.072948
|  |  | 2900 | 46.2398 | 2090 | 25.1004 | 2101.52 | 10.0084 | 2090 | 1.0008 | 2505.08 | 0.073194
|  |  | 2950 | 46.2251 | 2095 | 25.115 | 2107.16 | 10.0118 | 2095 | 1.0016 | 2533.14 | 0.073263
|  |  | 3000 | 46.2202 | 2100 | 25.1189 | 2112.8 | 10.0152 | 2100 | 1.0024 | 2561.2 | 0.073186
|  |  | 3050 | 46.2349 | 2105 | 25.1181 | 2118.44 | 10.0186 | 2105 | 1.0032 | 2589.26 | 0.073001
|  |  | 3100 | 46.2251 | 2110 | 25.1239 | 2124.08 | 10.022 | 2110 | 1.004 | 2617.32 | 0.072752
|  |  | 3150 | 46.2106 | 2115 | 25.1327 | 2129.72 | 10.0254 | 2115 | 1.0046 | 2645.38 | 0.07248
|  |  | 3200 | 46.2445 | 2120 | 25.139 | 2135.36 | 10.0287 | 2120 | 1.0051 | 2673.44 | 0.072219
|  |  | 3250 | 46.2445 | 2125 | 25.1414 | 2141 | 10.0319 | 2125 | 1.0054 | 2701.5 | 0.07199
|  |  | 3300 | 46.2251 | 2130 | 25.1417 | 2146.64 | 10.0351 | 2130 | 1.0055 | 2729.56 | 0.071804
|  |  | 3350 | 46.2154 | 2135 | 25.1569 | 2152.28 | 10.0382 | 2135 | 1.0055 | 2757.62 | 0.071665
|  |  | 3400 | 46.2202 | 2140 | 25.1489 | 2157.92 | 10.0412 | 2140 | 1.0053 | 2785.68 | 0.071575
|  |  | 3450 | 46.2105 | 2145 | 25.1519 | 2163.56 | 10.0442 | 2145 | 1.0051 | 2813.74 | 0.071535
|  |  | 3500 | 46.2494 | 2150 | 25.1579 | 2169.2 | 10.0471 | 2150 | 1.005 | 2841.8 | 0.071554
|  |  | 3550 | 46.2203 | 2155 | 25.165 | 2174.84 | 10.05 | 2155 | 1.0051 | 2869.86 | 0.071646
|  |  | 3600 | 46.2637 | 2160 | 25.1714 | 2180.48 | 10.0529 | 2160 | 1.0054 | 2897.92 | 0.071829
|  |  | 3650 | 46.2252 | 2165 | 25.1735 | 2186.12 | 10.0559 | 2165 | 1.0058 | 2925.98 | 0.072119
|  |  | 3700 | 46.2396 | 2170 | 25.1748 | 2191.76 | 10.0588 | 2170 | 1.0064 | 2954.04 | 0.072526
|  |  | 3750 | 46.2686 | 2175 | 25.1832 | 2197.4 | 10.0618 | 2175 | 1.007 | 2982.1 | 0.07304
|  |  | 3800 | 46.2252 | 2180 | 25.176 | 2203.04 | 10.0648 | 2180 | 1.0076 | 3010.16 | 0.073637
|  |  | 3850 | 46.2641 | 2185 | 25.1852 | 2208.68 | 10.0678 | 2185 | 1.0079 | 3038.22 | 0.074265
|  |  | 3900 | 46.2398 | 2190 | 25.1841 | 2214.32 | 10.0709 | 2190 | 1.008 | 3066.28 | 0.074859
|  |  | 3950 | 46.2594 | 2195 | 25.1897 | 2219.96 | 10.0741 | 2195 | 1.008 | 3094.34 | 0.075339
|  |  | 4000 | 46.2643 | 2200 | 25.196 | 2225.6 | 10.0773 | 2200 | 1.008 | 3122.4 | 0.075627
|  |  | 4050 | 46.23 | 2205 | 25.2021 | 2231.24 | 10.0806 | 2205 | 1.0082 | 3150.46 | 0.075656
|  |  | 4100 | 46.23 | 2210 | 25.2061 | 2236.88 | 10.0838 | 2210 | 1.0085 | 3178.52 | 0.075382
|  |  | 4150 | 46.264 | 2215 | 25.2074 | 2242.52 | 10.0871 | 2215 | 1.009 | 3206.58 | 0.074792
|  |  | 4200 | 46.2636 | 2220 | 25.2132 | 2248.16 | 10.0903 | 2220 | 1.0096 | 3234.64 | 0.073907
|  |  | 4250 | 46.3424 | 2225 | 25.2277 | 2253.8 | 10.0935 | 2225 | 1.0104 | 3262.7 | 0.072786
|  |  | 4300 | 46.3178 | 2230 | 25.2176 | 2259.44 | 10.0966 | 2230 | 1.0112 | 3290.76 | 0.07152
|  |  | 4350 | 46.2734 | 2235 | 25.2158 | 2265.08 | 10.0996 | 2235 | 1.0118 | 3318.82 | 0.070226
|  |  | 4400 | 46.2252 | 2240 | 25.2245 | 2270.72 | 10.1024 | 2240 | 1.0122 | 3346.88 | 0.069034
|  |  | 4450 | 46.2398 | 2245 | 25.2295 | 2276.36 | 10.1052 | 2245 | 1.0124 | 3374.94 | 0.068078
|  |  | 4500 | 46.2694 | 2250 | 25.237 | 2282 | 10.1079 | 2250 | 1.0125 | 3403 | 0.067474
|  |  | 4550 | 46.2548 | 2255 | 25.2423 | 2287.64 | 10.1105 | 2255 | 1.0126 | 3431.06 | 0.067311
|  |  | 4600 | 46.245 | 2260 | 25.2392 | 2293.28 | 10.1131 | 2260 | 1.0128 | 3459.12 | 0.067632
|  |  | 4650 | 46.245 | 2265 | 25.2397 | 2298.92 | 10.1158 | 2265 | 1.0132 | 3487.18 | 0.068423
|  |  | 4700 | 46.2651 | 2270 | 25.2458 | 2304.56 | 10.1184 | 2270 | 1.0138 | 3515.24 | 0.069614
|  |  | 4750 | 46.255 | 2275 | 25.2597 | 2310.2 | 10.1212 | 2275 | 1.0148 | 3543.3 | 0.071084
|  |  | 4800 | 46.225 | 2280 | 25.257 | 2315.84 | 10.124 | 2280 | 1.016 | 3571.36 | 0.072676
|  |  | 4850 | 46.2201 | 2285 | 25.2539 | 2321.48 | 10.1268 | 2285 | 1.0172 | 3599.42 | 0.074225
|  |  | 4900 | 46.23 | 2290 | 25.2567 | 2327.12 | 10.1297 | 2290 | 1.0185 | 3627.48 | 0.075585
|  |  | 4950 | 46.2497 | 2295 | 25.2628 | 2332.76 | 10.1326 | 2295 | 1.0196 | 3655.54 | 0.076657
|  |  | 5000 | 46.2153 | 2300 | 25.2667 | 2338.4 | 10.1355 | 2300 | 1.0205 | 3683.6 | 0.077407
|  |  | 5050 | 46.2103 | 2305 | 25.2655 | 2344.04 | 10.1384 | 2305 | 1.0212 | 3711.66 | 0.077871
|  |  | 5100 | 46.2349 | 2310 | 25.2627 | 2349.68 | 10.1412 | 2310 | 1.0217 | 3739.72 | 0.078147
|  |  | 5150 | 46.23 | 2315 | 25.267 | 2355.32 | 10.1439 | 2315 | 1.0221 | 3767.78 | 0.078368
|  |  | 5200 | 46.225 | 2320 | 25.2691 | 2360.96 | 10.1466 | 2320 | 1.0225 | 3795.84 | 0.078668
|  |  | 5250 | 46.2499 | 2325 | 25.2867 | 2366.6 | 10.1492 | 2325 | 1.023 | 3823.9 | 0.07915
|  |  | 5300 | 46.22 | 2330 | 25.2845 | 2372.24 | 10.1518 | 2330 | 1.0235 | 3851.96 | 0.079856
|  |  | 5350 | 46.22 | 2335 | 25.2813 | 2377.88 | 10.1544 | 2335 | 1.0243 | 3880.02 | 0.08075
|  |  | 5400 | 46.23 | 2340 | 25.2817 | 2383.52 | 10.1569 | 2340 | 1.0251 | 3908.08 | 0.081727
|  |  | 5450 | 46.2498 | 2345 | 25.2845 | 2389.16 | 10.1595 | 2345 | 1.026 | 3936.14 | 0.08263
|  |  | 5500 | 46.2744 | 2350 | 25.2893 | 2394.8 | 10.162 | 2350 | 1.0268 | 3964.2 | 0.08328
|  |  | 5550 | 46.2448 | 2355 | 25.2959 | 2400.44 | 10.1645 | 2355 | 1.0275 | 3992.26 | 0.083516
|  |  | 5600 | 46.2644 | 2360 | 25.3041 | 2406.08 | 10.1669 | 2360 | 1.028 | 4020.32 | 0.083226
|  |  | 5650 | 46.2546 | 2365 | 25.3084 | 2411.72 | 10.1694 | 2365 | 1.028 | 4048.38 | 0.082371
|  |  | 5700 | 46.2644 | 2370 | 25.3136 | 2417.36 | 10.1718 | 2370 | 1.0277 | 4076.44 | 0.08099
|  |  | 5750 | 46.2251 | 2375 | 25.319 | 2423 | 10.1741 | 2375 | 1.0271 | 4104.5 | 0.079199
|  |  | 5800 | 46.2054 | 2380 | 25.3242 | 2428.64 | 10.1764 | 2380 | 1.0263 | 4132.56 | 0.077165
|  |  | 5850 | 46.2943 | 2385 | 25.324 | 2434.28 | 10.1787 | 2385 | 1.0256 | 4160.62 | 0.075089
|  |  | 5900 | 46.2697 | 2390 | 25.3213 | 2439.92 | 10.1809 | 2390 | 1.025 | 4188.68 | 0.07317
|  |  | 5950 | 46.2847 | 2395 | 25.323 | 2445.56 | 10.1831 | 2395 | 1.0246 | 4216.74 | 0.071592
|  |  | 6000 | 46.24 | 2400 | 25.3284 | 2451.2 | 10.1853 | 2400 | 1.0247 | 4244.8 | 0.070503
|  |  | 6050 | 46.24 | 2405 | 25.3283 | 2456.84 | 10.1874 | 2405 | 1.025 | 4272.86 | 0.07001
|  |  | 6100 | 46.245 | 2410 | 25.3272 | 2462.48 | 10.1897 | 2410 | 1.0257 | 4300.92 | 0.070176
|  |  | 6150 | 46.2499 | 2415 | 25.3317 | 2468.12 | 10.1919 | 2415 | 1.0265 | 4328.98 | 0.071022
|  |  | 6200 | 46.235 | 2420 | 25.3362 | 2473.76 | 10.1942 | 2420 | 1.0273 | 4357.04 | 0.072533
|  |  | 6250 | 46.2251 | 2425 | 25.3417 | 2479.4 | 10.1966 | 2425 | 1.028 | 4385.1 | 0.074655
|  |  | 6300 | 46.2103 | 2430 | 25.3469 | 2485.04 | 10.199 | 2430 | 1.0287 | 4413.16 | 0.077299
|  |  | 6350 | 46.2054 | 2435 | 25.3494 | 2490.68 | 10.2014 | 2435 | 1.0293 | 4441.22 | 0.080339
|  |  | 6400 | 46.1763 | 2440 | 25.35 | 2496.32 | 10.2039 | 2440 | 1.0298 | 4469.28 | 0.083615
|  |  | 6450 | 46.2105 | 2445 | 25.3509 | 2501.96 | 10.2064 | 2445 | 1.0304 | 4497.34 | 0.086938
|  |  | 6500 | 46.23 | 2450 | 25.3502 | 2507.6 | 10.2089 | 2450 | 1.0309 | 4525.4 | 0.090102
|  |  | 6550 | 46.2397 | 2455 | 25.3591 | 2513.24 | 10.2115 | 2455 | 1.0314 | 4553.46 | 0.092902
|  |  | 6600 | 46.225 | 2460 | 25.3569 | 2518.88 | 10.2141 | 2460 | 1.0319 | 4581.52 | 0.095158
|  |  | 6650 | 46.1949 | 2465 | 25.3607 | 2524.52 | 10.2168 | 2465 | 1.0324 | 4609.58 | 0.096741
|  |  | 6700 | 46.205 | 2470 | 25.3697 | 2530.16 | 10.2196 | 2470 | 1.0328 | 4637.64 | 0.09759
|  |  | 6750 | 46.2451 | 2475 | 25.3775 | 2535.8 | 10.2225 | 2475 | 1.0332 | 4665.7 | 0.097732
|  |  | 6800 | 46.2199 | 2480 | 25.3817 | 2541.44 | 10.2257 | 2480 | 1.0336 | 4693.76 | 0.097281
|  |  | 6850 | 46.2502 | 2485 | 25.3844 | 2547.08 | 10.2292 | 2485 | 1.0339 | 4721.82 | 0.096426
|  |  | 6900 | 46.2199 | 2490 | 25.3849 | 2552.72 | 10.233 | 2490 | 1.0342 | 4749.88 | 0.095407
|  |  | 6950 | 46.2048 | 2495 | 25.3776 | 2558.36 | 10.2371 | 2495 | 1.0345 | 4777.94 | 0.094477
|  |  | 7000 | 46.2099 | 2500 | 25.3851 | 2564 | 10.2415 | 2500 | 1.0347 | 4806 | 0.093862
|  |  | 7050 | 46.21 | 2505 | 25.3902 | 2569.64 | 10.2461 | 2505 | 1.0347 | 4834.06 | 0.093725
|  |  | 7100 | 46.22 | 2510 | 25.3942 | 2575.28 | 10.2507 | 2510 | 1.0344 | 4862.12 | 0.094136
|  |  | 7150 | 46.235 | 2515 | 25.3959 | 2580.92 | 10.2552 | 2515 | 1.034 | 4890.18 | 0.095058
|  |  | 7200 | 46.2052 | 2520 | 25.3936 | 2586.56 | 10.2594 | 2520 | 1.0333 | 4918.24 | 0.096361
|  |  | 7250 | 46.2102 | 2525 | 25.3933 | 2592.2 | 10.263 | 2525 | 1.0324 | 4946.3 | 0.097837
|  |  | 7300 | 46.1904 | 2530 | 25.4 | 2597.84 | 10.2659 | 2530 | 1.0314 | 4974.36 | 0.099244
|  |  | 7350 | 46.2152 | 2535 | 25.4096 | 2603.48 | 10.2682 | 2535 | 1.0304 | 5002.42 | 0.10035
|  |  | 7400 | 46.1955 | 2540 | 25.4139 | 2609.12 | 10.2697 | 2540 | 1.0295 | 5030.48 | 0.10096
|  |  | 7450 | 46.23 | 2545 | 25.408 | 2614.76 | 10.2708 | 2545 | 1.0288 | 5058.54 | 0.10099
|  |  | 7500 | 46.23 | 2550 | 25.4034 | 2620.4 | 10.2715 | 2550 | 1.0284 | 5086.6 | 0.10043
|  |  | 7550 | 46.2201 | 2555 | 25.4149 | 2626.04 | 10.2721 | 2555 | 1.0282 | 5114.66 | 0.099391
|  |  | 7600 | 46.23 | 2560 | 25.4212 | 2631.68 | 10.2727 | 2560 | 1.0284 | 5142.72 | 0.098045
|  |  | 7650 | 46.205 | 2565 | 25.4244 | 2637.32 | 10.2736 | 2565 | 1.0289 | 5170.78 | 0.096628
|  |  | 7700 | 46.225 | 2570 | 25.4266 | 2642.96 | 10.2749 | 2570 | 1.0297 | 5198.84 | 0.095381
|  |  | 7750 | 46.1949 | 2575 | 25.4283 | 2648.6 | 10.2766 | 2575 | 1.0307 | 5226.9 | 0.094512
|  |  | 7800 | 46.1999 | 2580 | 25.4313 | 2654.24 | 10.2785 | 2580 | 1.032 | 5254.96 | 0.094162
|  |  | 7850 | 46.1999 | 2585 | 25.4348 | 2659.88 | 10.2807 | 2585 | 1.0334 | 5283.02 | 0.094382
|  |  | 7900 | 46.225 | 2590 | 25.4406 | 2665.52 | 10.283 | 2590 | 1.035 | 5311.08 | 0.095133
|  |  | 7950 | 46.215 | 2595 | 25.4451 | 2671.16 | 10.2851 | 2595 | 1.0368 | 5339.14 | 0.096292
|  |  | 8000 | 46.235 | 2600 | 25.4452 | 2676.8 | 10.2872 | 2600 | 1.0385 | 5367.2 | 0.097679
|  |  | 8050 | 46.235 | 2605 | 25.4516 | 2682.44 | 10.2891 | 2605 | 1.0403 | 5395.26 | 0.099094
|  |  | 8100 | 46.1953 | 2610 | 25.4523 | 2688.08 | 10.2908 | 2610 | 1.042 | 5423.32 | 0.10035
|  |  | 8150 | 46.2151 | 2615 | 25.451 | 2693.72 | 10.2925 | 2615 | 1.0435 | 5451.38 | 0.10132
|  |  | 8200 | 46.235 | 2620 | 25.4535 | 2699.36 | 10.294 | 2620 | 1.0447 | 5479.44 | 0.10191
|  |  | 8250 | 46.2399 | 2625 | 25.4559 | 2705 | 10.2956 | 2625 | 1.0455 | 5507.5 | 0.10215
|  |  | 8300 | 46.2152 | 2630 | 25.4564 | 2710.64 | 10.2972 | 2630 | 1.0459 | 5535.56 | 0.10209
|  |  | 8350 | 46.2152 | 2635 | 25.459 | 2716.28 | 10.2989 | 2635 | 1.0457 | 5563.62 | 0.10183
|  |  | 8400 | 46.23 | 2640 | 25.4546 | 2721.92 | 10.3005 | 2640 | 1.0448 | 5591.68 | 0.10149
|  |  | 8450 | 46.23 | 2645 | 25.4479 | 2727.56 | 10.3021 | 2645 | 1.0435 | 5619.74 | 0.10116
|  |  | 8500 | 46.2349 | 2650 | 25.4563 | 2733.2 | 10.3036 | 2650 | 1.042 | 5647.8 | 0.10087
|  |  | 8550 | 46.23 | 2655 | 25.4671 | 2738.84 | 10.305 | 2655 | 1.0403 | 5675.86 | 0.10061
|  |  | 8600 | 46.2545 | 2660 | 25.4695 | 2744.48 | 10.3063 | 2660 | 1.0388 | 5703.92 | 0.10031
|  |  | 8650 | 46.2349 | 2665 | 25.4677 | 2750.12 | 10.3075 | 2665 | 1.0376 | 5731.98 | 0.099886
|  |  | 8700 | 46.2153 | 2670 | 25.4723 | 2755.76 | 10.3087 | 2670 | 1.0368 | 5760.04 | 0.099221
|  |  | 8750 | 46.2497 | 2675 | 25.4837 | 2761.4 | 10.3101 | 2675 | 1.0365 | 5788.1 | 0.09825
|  |  | 8800 | 46.2201 | 2680 | 25.486 | 2767.04 | 10.3116 | 2680 | 1.0366 | 5816.16 | 0.096958
|  |  | 8850 | 46.2201 | 2685 | 25.4808 | 2772.68 | 10.3134 | 2685 | 1.0371 | 5844.22 | 0.095406
|  |  | 8900 | 46.23 | 2690 | 25.474 | 2778.32 | 10.3155 | 2690 | 1.0378 | 5872.28 | 0.093728
|  |  | 8950 | 46.2399 | 2695 | 25.4741 | 2783.96 | 10.3177 | 2695 | 1.0385 | 5900.34 | 0.092114
|  |  | 9000 | 46.2152 | 2700 | 25.4803 | 2789.6 | 10.3201 | 2700 | 1.0393 | 5928.4 | 0.090787
|  |  | 9050 | 46.2349 | 2705 | 25.489 | 2795.24 | 10.3226 | 2705 | 1.0401 | 5956.46 | 0.089959
|  |  | 9100 | 46.2006 | 2710 | 25.4925 | 2800.88 | 10.325 | 2710 | 1.041 | 5984.52 | 0.089804
|  |  | 9150 | 46.2349 | 2715 | 25.491 | 2806.52 | 10.3272 | 2715 | 1.0419 | 6012.58 | 0.090423
|  |  | 9200 | 46.2058 | 2720 | 25.5013 | 2812.16 | 10.3292 | 2720 | 1.0428 | 6040.64 | 0.091828
|  |  | 9250 | 46.2252 | 2725 | 25.5048 | 2817.8 | 10.3308 | 2725 | 1.0437 | 6068.7 | 0.093938
|  |  | 9300 | 46.2348 | 2730 | 25.5034 | 2823.44 | 10.3321 | 2730 | 1.0445 | 6096.76 | 0.096586
|  |  | 9350 | 46.2348 | 2735 | 25.5027 | 2829.08 | 10.3331 | 2735 | 1.0451 | 6124.82 | 0.099542
|  |  | 9400 | 46.2111 | 2740 | 25.5063 | 2834.72 | 10.3338 | 2740 | 1.0455 | 6152.88 | 0.10254
|  |  | 9450 | 46.1969 | 2745 | 25.5069 | 2840.36 | 10.3343 | 2745 | 1.0455 | 6180.94 | 0.1053
|  |  | 9500 | 46.2206 | 2750 | 25.5053 | 2846 | 10.3348 | 2750 | 1.0451 | 6209 | 0.10758
|  |  | 9550 | 46.2394 | 2755 | 25.5091 | 2851.64 | 10.3352 | 2755 | 1.0446 | 6237.06 | 0.10918
|  |  | 9600 | 46.2063 | 2760 | 25.517 | 2857.28 | 10.3358 | 2760 | 1.044 | 6265.12 | 0.10998
|  |  | 9650 | 46.2205 | 2765 | 25.5129 | 2862.92 | 10.3365 | 2765 | 1.0433 | 6293.18 | 0.10994
|  |  | 9700 | 46.2062 | 2770 | 25.5035 | 2868.56 | 10.3373 | 2770 | 1.0429 | 6321.24 | 0.10913
|  |  | 9750 | 46.2013 | 2775 | 25.5072 | 2874.2 | 10.3383 | 2775 | 1.0426 | 6349.3 | 0.10768
|  |  | 9800 | 46.206 | 2780 | 25.5129 | 2879.84 | 10.3394 | 2780 | 1.0426 | 6377.36 | 0.1058
|  |  | 9850 | 46.2058 | 2785 | 25.5117 | 2885.48 | 10.3407 | 2785 | 1.0429 | 6405.42 | 0.10372
|  |  | 9900 | 46.2251 | 2790 | 25.5102 | 2891.12 | 10.3421 | 2790 | 1.0433 | 6433.48 | 0.10169
|  |  | 9950 | 46.2251 | 2795 | 25.5102 | 2896.76 | 10.3436 | 2795 | 1.0439 | 6461.54 | 0.099922
|  |  | 10000 | 46.2056 | 2800 | 25.5119 | 2902.4 | 10.3452 | 2800 | 1.0446 | 6489.6 | 0.098561
|  |  | 10050 | 46.2202 | 2805 | 25.5158 | 2908.04 | 10.3469 | 2805 | 1.0454 | 6517.66 | 0.097675
|  |  | 10100 | 46.2153 | 2810 | 25.5212 | 2913.68 | 10.3488 | 2810 | 1.0462 | 6545.72 | 0.097246
|  |  | 10150 | 46.23 | 2815 | 25.5267 | 2919.32 | 10.3507 | 2815 | 1.0471 | 6573.78 | 0.097182
|  |  | 10200 | 46.23 | 2820 | 25.5262 | 2924.96 | 10.3528 | 2820 | 1.048 | 6601.84 | 0.09734
|  |  | 10250 | 46.2104 | 2825 | 25.524 | 2930.6 | 10.3549 | 2825 | 1.049 | 6629.9 | 0.09756
|  |  | 10300 | 46.2103 | 2830 | 25.5214 | 2936.24 | 10.357 | 2830 | 1.05 | 6657.96 | 0.097704
|  |  | 10350 | 46.2103 | 2835 | 25.5197 | 2941.88 | 10.359 | 2835 | 1.051 | 6686.02 | 0.097681
|  |  | 10400 | 46.1858 | 2840 | 25.5215 | 2947.52 | 10.3609 | 2840 | 1.0522 | 6714.08 | 0.097468
|  |  | 10450 | 46.2152 | 2845 | 25.5251 | 2953.16 | 10.3627 | 2845 | 1.0534 | 6742.14 | 0.097112
|  |  | 10500 | 46.2103 | 2850 | 25.5258 | 2958.8 | 10.3643 | 2850 | 1.0546 | 6770.2 | 0.096717
|  |  | 10550 | 46.2251 | 2855 | 25.5237 | 2964.44 | 10.3657 | 2855 | 1.0558 | 6798.26 | 0.096421
|  |  | 10600 | 46.2004 | 2860 | 25.5272 | 2970.08 | 10.367 | 2860 | 1.057 | 6826.32 | 0.096362
|  |  | 10650 | 46.2053 | 2865 | 25.5306 | 2975.72 | 10.3682 | 2865 | 1.058 | 6854.38 | 0.096649
|  |  | 10700 | 46.2003 | 2870 | 25.5311 | 2981.36 | 10.3694 | 2870 | 1.059 | 6882.44 | 0.097336
|  |  | 10750 | 46.1853 | 2875 | 25.5374 | 2987 | 10.3706 | 2875 | 1.0598 | 6910.5 | 0.098412
|  |  | 10800 | 46.1902 | 2880 | 25.5357 | 2992.64 | 10.372 | 2880 | 1.0605 | 6938.56 | 0.099798
|  |  | 10850 | 46.1703 | 2885 | 25.5305 | 2998.28 | 10.3734 | 2885 | 1.0612 | 6966.62 | 0.10136
|  |  | 10900 | 46.1802 | 2890 | 25.5318 | 3003.92 | 10.375 | 2890 | 1.0619 | 6994.68 | 0.10294
|  |  | 10950 | 46.22 | 2895 | 25.5379 | 3009.56 | 10.3767 | 2895 | 1.0626 | 7022.74 | 0.10436
|  |  | 11000 | 46.2002 | 2900 | 25.5373 | 3015.2 | 10.3784 | 2900 | 1.0633 | 7050.8 | 0.10547
|  |  | 11050 | 46.1952 | 2905 | 25.5414 | 3020.84 | 10.3803 | 2905 | 1.0639 | 7078.86 | 0.10614
|  |  | 11100 | 46.1904 | 2910 | 25.5445 | 3026.48 | 10.3821 | 2910 | 1.0646 | 7106.92 | 0.10631
|  |  | 11150 | 46.2152 | 2915 | 25.5489 | 3032.12 | 10.384 | 2915 | 1.0651 | 7134.98 | 0.10594
|  |  | 11200 | 46.1906 | 2920 | 25.5515 | 3037.76 | 10.3859 | 2920 | 1.0654 | 7163.04 | 0.10507
|  |  | 11250 | 46.1809 | 2925 | 25.5518 | 3043.4 | 10.3878 | 2925 | 1.0655 | 7191.1 | 0.10376
|  |  | 11300 | 46.1859 | 2930 | 25.5519 | 3049.04 | 10.3896 | 2930 | 1.0655 | 7219.16 | 0.1021
|  |  | 11350 | 46.1761 | 2935 | 25.5525 | 3054.68 | 10.3915 | 2935 | 1.0652 | 7247.22 | 0.10019
|  |  | 11400 | 46.1957 | 2940 | 25.5528 | 3060.32 | 10.3933 | 2940 | 1.0647 | 7275.28 | 0.098157
|  |  | 11450 | 46.1809 | 2945 | 25.5521 | 3065.96 | 10.3951 | 2945 | 1.0641 | 7303.34 | 0.096114
|  |  | 11500 | 46.1759 | 2950 | 25.5527 | 3071.6 | 10.3969 | 2950 | 1.0634 | 7331.4 | 0.094175
|  |  | 11550 | 46.1807 | 2955 | 25.56 | 3077.24 | 10.3987 | 2955 | 1.0629 | 7359.46 | 0.092448
|  |  | 11600 | 46.1558 | 2960 | 25.5642 | 3082.88 | 10.4005 | 2960 | 1.0627 | 7387.52 | 0.091027
|  |  | 11650 | 46.1804 | 2965 | 25.5653 | 3088.52 | 10.4022 | 2965 | 1.0629 | 7415.58 | 0.089999
|  |  | 11700 | 46.1901 | 2970 | 25.5673 | 3094.16 | 10.4038 | 2970 | 1.0634 | 7443.64 | 0.089429
|  |  | 11750 | 46.1699 | 2975 | 25.5693 | 3099.8 | 10.4053 | 2975 | 1.0644 | 7471.7 | 0.089364
|  |  | 11800 | 46.1697 | 2980 | 25.5715 | 3105.44 | 10.4067 | 2980 | 1.0656 | 7499.76 | 0.089818
|  |  | 11850 | 46.1796 | 2985 | 25.5735 | 3111.08 | 10.4079 | 2985 | 1.0671 | 7527.82 | 0.090773
|  |  | 11900 | 46.1543 | 2990 | 25.5723 | 3116.72 | 10.4091 | 2990 | 1.0685 | 7555.88 | 0.092169
|  |  | 11950 | 46.1694 | 2995 | 25.57 | 3122.36 | 10.4101 | 2995 | 1.0697 | 7583.94 | 0.093906
|  |  | 12000 | 46.1695 | 3000 | 25.5713 | 3128 | 10.4111 | 3000 | 1.0707 | 7612 | 0.095846
|  |  | 12050 | 46.2048 | 3005 | 25.5802 | 3133.64 | 10.4121 | 3005 | 1.0713 | 7640.06 | 0.097819
|  |  | 12100 | 46.1848 | 3010 | 25.5845 | 3139.28 | 10.4131 | 3010 | 1.0715 | 7668.12 | 0.099641
|  |  | 12150 | 46.2 | 3015 | 25.5854 | 3144.92 | 10.4142 | 3015 | 1.0716 | 7696.18 | 0.10113
|  |  | 12200 | 46.1651 | 3020 | 25.5854 | 3150.56 | 10.4155 | 3020 | 1.0714 | 7724.24 | 0.10213
|  |  | 12250 | 46.1901 | 3025 | 25.5872 | 3156.2 | 10.417 | 3025 | 1.0713 | 7752.3 | 0.10253
|  |  | 12300 | 46.1804 | 3030 | 25.5894 | 3161.84 | 10.4187 | 3030 | 1.0712 | 7780.36 | 0.10226
|  |  | 12350 | 46.1606 | 3035 | 25.5882 | 3167.48 | 10.4206 | 3035 | 1.0711 | 7808.42 | 0.10133
|  |  | 12400 | 46.1657 | 3040 | 25.5857 | 3173.12 | 10.4226 | 3040 | 1.0712 | 7836.48 | 0.099841
|  |  | 12450 | 46.1707 | 3045 | 25.5904 | 3178.76 | 10.4247 | 3045 | 1.0714 | 7864.54 | 0.097919
|  |  | 12500 | 46.1708 | 3050 | 25.5925 | 3184.4 | 10.4269 | 3050 | 1.0717 | 7892.6 | 0.095754
|  |  | 12550 | 46.1758 | 3055 | 25.5954 | 3190.04 | 10.429 | 3055 | 1.072 | 7920.66 | 0.093554
|  |  | 12600 | 46.1759 | 3060 | 25.5995 | 3195.68 | 10.4311 | 3060 | 1.0723 | 7948.72 | 0.091523
|  |  | 12650 | 46.1858 | 3065 | 25.6061 | 3201.32 | 10.433 | 3065 | 1.0725 | 7976.78 | 0.089836
|  |  | 12700 | 46.1613 | 3070 | 25.6086 | 3206.96 | 10.4348 | 3070 | 1.0728 | 8004.84 | 0.088623
|  |  | 12750 | 46.1663 | 3075 | 25.6048 | 3212.6 | 10.4364 | 3075 | 1.0731 | 8032.9 | 0.087953
|  |  | 12800 | 46.1712 | 3080 | 25.6045 | 3218.24 | 10.4379 | 3080 | 1.0733 | 8060.96 | 0.087836
|  |  | 12850 | 46.1858 | 3085 | 25.6088 | 3223.88 | 10.4392 | 3085 | 1.0736 | 8089.02 | 0.088227
|  |  | 12900 | 46.1563 | 3090 | 25.6038 | 3229.52 | 10.4405 | 3090 | 1.0739 | 8117.08 | 0.089041
|  |  | 12950 | 46.2005 | 3095 | 25.6005 | 3235.16 | 10.4417 | 3095 | 1.0742 | 8145.14 | 0.090173
|  |  | 13000 | 46.1954 | 3100 | 25.6032 | 3240.8 | 10.4428 | 3100 | 1.0745 | 8173.2 | 0.091512
|  |  | 13050 | 46.1953 | 3105 | 25.6085 | 3246.44 | 10.444 | 3105 | 1.0748 | 8201.26 | 0.092961
|  |  | 13100 | 46.1804 | 3110 | 25.6159 | 3252.08 | 10.4452 | 3110 | 1.0752 | 8229.32 | 0.094443
|  |  | 13150 | 46.1952 | 3115 | 25.62 | 3257.72 | 10.4464 | 3115 | 1.0757 | 8257.38 | 0.095906
|  |  | 13200 | 46.1702 | 3120 | 25.6187 | 3263.36 | 10.4476 | 3120 | 1.0761 | 8285.44 | 0.097318
|  |  | 13250 | 46.185 | 3125 | 25.6189 | 3269 | 10.4489 | 3125 | 1.0767 | 8313.5 | 0.098661
|  |  | 13300 | 46.165 | 3130 | 25.6226 | 3274.64 | 10.4503 | 3130 | 1.0772 | 8341.56 | 0.099923
|  |  | 13350 | 46.21 | 3135 | 25.6272 | 3280.28 | 10.4516 | 3135 | 1.0777 | 8369.62 | 0.10109
|  |  | 13400 | 46.1651 | 3140 | 25.6286 | 3285.92 | 10.453 | 3140 | 1.0782 | 8397.68 | 0.10212
|  |  | 13450 | 46.1901 | 3145 | 25.6281 | 3291.56 | 10.4543 | 3145 | 1.0785 | 8425.74 | 0.10299
|  |  | 13500 | 46.1802 | 3150 | 25.6282 | 3297.2 | 10.4557 | 3150 | 1.0787 | 8453.8 | 0.10366
|  |  | 13550 | 46.1606 | 3155 | 25.6323 | 3302.84 | 10.457 | 3155 | 1.0787 | 8481.86 | 0.10409
|  |  | 13600 | 46.1806 | 3160 | 25.637 | 3308.48 | 10.4584 | 3160 | 1.0785 | 8509.92 | 0.10424
|  |  | 13650 | 46.1955 | 3165 | 25.6381 | 3314.12 | 10.4597 | 3165 | 1.0783 | 8537.98 | 0.10413
|  |  | 13700 | 46.1858 | 3170 | 25.6375 | 3319.76 | 10.4609 | 3170 | 1.0779 | 8566.04 | 0.10375
|  |  | 13750 | 46.2202 | 3175 | 25.6383 | 3325.4 | 10.4621 | 3175 | 1.0776 | 8594.1 | 0.10318
|  |  | 13800 | 46.1715 | 3180 | 25.6401 | 3331.04 | 10.4632 | 3180 | 1.0773 | 8622.16 | 0.10246
|  |  | 13850 | 46.1911 | 3185 | 25.642 | 3336.68 | 10.4643 | 3185 | 1.0772 | 8650.22 | 0.10167
|  |  | 13900 | 46.1718 | 3190 | 25.6416 | 3342.32 | 10.4653 | 3190 | 1.0772 | 8678.28 | 0.10089
|  |  | 13950 | 46.1865 | 3195 | 25.6436 | 3347.96 | 10.4663 | 3195 | 1.0774 | 8706.34 | 0.10019
|  |  | 14000 | 46.1962 | 3200 | 25.6451 | 3353.6 | 10.4673 | 3200 | 1.0778 | 8734.4 | 0.099591
[5 empty] | 3205 | 25.6459 | 3359.24 | 10.4683 | 3205 | 1.0783 | 8762.46 | 0.099107
[5 empty] | 3210 | 25.6492 | 3364.88 | 10.4693 | 3210 | 1.0788 | 8790.52 | 0.09872
[5 empty] | 3215 | 25.6556 | 3370.52 | 10.4703 | 3215 | 1.0793 | 8818.58 | 0.098387
[5 empty] | 3220 | 25.6595 | 3376.16 | 10.4714 | 3220 | 1.0798 | 8846.64 | 0.098057
[5 empty] | 3225 | 25.6581 | 3381.8 | 10.4726 | 3225 | 1.0802 | 8874.7 | 0.097678
[5 empty] | 3230 | 25.6594 | 3387.44 | 10.4739 | 3230 | 1.0803 | 8902.76 | 0.09721
[5 empty] | 3235 | 25.6639 | 3393.08 | 10.4753 | 3235 | 1.0803 | 8930.82 | 0.096634
[5 empty] | 3240 | 25.6641 | 3398.72 | 10.4767 | 3240 | 1.0802 | 8958.88 | 0.095954
[5 empty] | 3245 | 25.6627 | 3404.36 | 10.4783 | 3245 | 1.08 | 8986.94 | 0.095197
[5 empty] | 3250 | 25.6635 | 3410 | 10.4799 | 3250 | 1.0799 | 9015 | 0.094401
[5 empty] | 3255 | 25.6668 | 3415.64 | 10.4816 | 3255 | 1.0799 | 9043.06 | 0.093607
[5 empty] | 3260 | 25.6698 | 3421.28 | 10.4833 | 3260 | 1.0802 | 9071.12 | 0.092845
[5 empty] | 3265 | 25.6715 | 3426.92 | 10.485 | 3265 | 1.0806 | 9099.18 | 0.092124
[5 empty] | 3270 | 25.6749 | 3432.56 | 10.4867 | 3270 | 1.0811 | 9127.24 | 0.091427
[5 empty] | 3275 | 25.678 | 3438.2 | 10.4882 | 3275 | 1.0818 | 9155.3 | 0.090707
[5 empty] | 3280 | 25.6799 | 3443.84 | 10.4897 | 3280 | 1.0824 | 9183.36 | 0.089898
[5 empty] | 3285 | 25.6806 | 3449.48 | 10.491 | 3285 | 1.083 | 9211.42 | 0.088923
[5 empty] | 3290 | 25.6794 | 3455.12 | 10.4921 | 3290 | 1.0834 | 9239.48 | 0.087712
[5 empty] | 3295 | 25.6781 | 3460.76 | 10.493 | 3295 | 1.0837 | 9267.54 | 0.086221
[5 empty] | 3300 | 25.6794 | 3466.4 | 10.4937 | 3300 | 1.0838 | 9295.6 | 0.084446
[5 empty] | 3305 | 25.6833 | 3472.04 | 10.4942 | 3305 | 1.0838 | 9323.66 | 0.082431
[5 empty] | 3310 | 25.6858 | 3477.68 | 10.4947 | 3310 | 1.0838 | 9351.72 | 0.080276
[5 empty] | 3315 | 25.6895 | 3483.32 | 10.4952 | 3315 | 1.0838 | 9379.78 | 0.078127
[5 empty] | 3320 | 25.6925 | 3488.96 | 10.4958 | 3320 | 1.0838 | 9407.84 | 0.076166
[5 empty] | 3325 | 25.6916 | 3494.6 | 10.4965 | 3325 | 1.0839 | 9435.9 | 0.07459
[5 empty] | 3330 | 25.6943 | 3500.24 | 10.4974 | 3330 | 1.084 | 9463.96 | 0.073588
[5 empty] | 3335 | 25.6961 | 3505.88 | 10.4986 | 3335 | 1.0843 | 9492.02 | 0.073321
[5 empty] | 3340 | 25.6936 | 3511.52 | 10.5 | 3340 | 1.0847 | 9520.08 | 0.073898
[5 empty] | 3345 | 25.6908 | 3517.16 | 10.5017 | 3345 | 1.0852 | 9548.14 | 0.075369
[5 empty] | 3350 | 25.6925 | 3522.8 | 10.5036 | 3350 | 1.0858 | 9576.2 | 0.077716
[5 empty] | 3355 | 25.6955 | 3528.44 | 10.5056 | 3355 | 1.0864 | 9604.26 | 0.080857
[5 empty] | 3360 | 25.6959 | 3534.08 | 10.5077 | 3360 | 1.087 | 9632.32 | 0.084658
[5 empty] | 3365 | 25.7015 | 3539.72 | 10.5097 | 3365 | 1.0875 | 9660.38 | 0.088948
[5 empty] | 3370 | 25.7038 | 3545.36 | 10.5115 | 3370 | 1.0877 | 9688.44 | 0.093538
[5 empty] | 3375 | 25.7009 | 3551 | 10.5133 | 3375 | 1.0877 | 9716.5 | 0.098239
[5 empty] | 3380 | 25.6992 | 3556.64 | 10.5148 | 3380 | 1.0875 | 9744.56 | 0.10288
[5 empty] | 3385 | 25.7008 | 3562.28 | 10.5162 | 3385 | 1.0871 | 9772.62 | 0.10731
[5 empty] | 3390 | 25.7024 | 3567.92 | 10.5173 | 3390 | 1.0866 | 9800.68 | 0.11143
[5 empty] | 3395 | 25.7016 | 3573.56 | 10.5183 | 3395 | 1.0862 | 9828.74 | 0.11515
[5 empty] | 3400 | 25.7002 | 3579.2 | 10.5192 | 3400 | 1.0859 | 9856.8 | 0.11843
[5 empty] | 3405 | 25.6986 | 3584.84 | 10.52 | 3405 | 1.0858 | 9884.86 | 0.12124
[5 empty] | 3410 | 25.697 | 3590.48 | 10.5208 | 3410 | 1.086 | 9912.92 | 0.12356
[5 empty] | 3415 | 25.698 | 3596.12 | 10.5216 | 3415 | 1.0864 | 9940.98 | 0.12539
[5 empty] | 3420 | 25.6977 | 3601.76 | 10.5225 | 3420 | 1.087 | 9969.04 | 0.12672
[5 empty] | 3425 | 25.698 | 3607.4 | 10.5234 | 3425 | 1.0877 | 9997.1 | 0.12754
[5 empty] | 3430 | 25.703 | 3613.04 | 10.5244 | 3430 | 1.0884 | 10025.16 | 0.12784
[5 empty] | 3435 | 25.7115 | 3618.68 | 10.5256 | 3435 | 1.089 | 10053.22 | 0.12761
[5 empty] | 3440 | 25.718 | 3624.32 | 10.5268 | 3440 | 1.0895 | 10081.28 | 0.12685
[5 empty] | 3445 | 25.7209 | 3629.96 | 10.528 | 3445 | 1.0898 | 10109.34 | 0.12556
[5 empty] | 3450 | 25.7252 | 3635.6 | 10.5293 | 3450 | 1.0901 | 10137.4 | 0.12378
[5 empty] | 3455 | 25.7287 | 3641.24 | 10.5306 | 3455 | 1.0903 | 10165.46 | 0.12154
[5 empty] | 3460 | 25.7319 | 3646.88 | 10.5319 | 3460 | 1.0906 | 10193.52 | 0.11891
[5 empty] | 3465 | 25.7352 | 3652.52 | 10.533 | 3465 | 1.0909 | 10221.58 | 0.11597
[5 empty] | 3470 | 25.7363 | 3658.16 | 10.534 | 3470 | 1.0913 | 10249.64 | 0.11282
[5 empty] | 3475 | 25.7361 | 3663.8 | 10.5349 | 3475 | 1.0918 | 10277.7 | 0.10953
[5 empty] | 3480 | 25.7365 | 3669.44 | 10.5357 | 3480 | 1.0923 | 10305.76 | 0.10623
[5 empty] | 3485 | 25.7368 | 3675.08 | 10.5365 | 3485 | 1.0928 | 10333.82 | 0.10298
[5 empty] | 3490 | 25.7345 | 3680.72 | 10.5372 | 3490 | 1.0932 | 10361.88 | 0.099871
[5 empty] | 3495 | 25.7327 | 3686.36 | 10.5379 | 3495 | 1.0935 | 10389.94 | 0.096961
[5 empty] | 3500 | 25.7347 | 3692 | 10.5387 | 3500 | 1.0937 | 10418 | 0.094294
[5 empty] | 3505 | 25.738 | 3697.64 | 10.5397 | 3505 | 1.0939 | 10446.06 | 0.0919
[5 empty] | 3510 | 25.7417 | 3703.28 | 10.5407 | 3510 | 1.094 | 10474.12 | 0.089794
[5 empty] | 3515 | 25.744 | 3708.92 | 10.5419 | 3515 | 1.0941 | 10502.18 | 0.087982
[5 empty] | 3520 | 25.7455 | 3714.56 | 10.5433 | 3520 | 1.0943 | 10530.24 | 0.086468
[5 empty] | 3525 | 25.7484 | 3720.2 | 10.5447 | 3525 | 1.0945 | 10558.3 | 0.085251
[5 empty] | 3530 | 25.7536 | 3725.84 | 10.5461 | 3530 | 1.0946 | 10586.36 | 0.084335
[5 empty] | 3535 | 25.7587 | 3731.48 | 10.5475 | 3535 | 1.0948 | 10614.42 | 0.083727
[5 empty] | 3540 | 25.7618 | 3737.12 | 10.5489 | 3540 | 1.0949 | 10642.48 | 0.083437
[5 empty] | 3545 | 25.7647 | 3742.76 | 10.5502 | 3545 | 1.0948 | 10670.54 | 0.083477
[5 empty] | 3550 | 25.7668 | 3748.4 | 10.5514 | 3550 | 1.0947 | 10698.6 | 0.083862
[5 empty] | 3555 | 25.7676 | 3754.04 | 10.5525 | 3555 | 1.0945 | 10726.66 | 0.084599
[5 empty] | 3560 | 25.7678 | 3759.68 | 10.5536 | 3560 | 1.0942 | 10754.72 | 0.08569
[5 empty] | 3565 | 25.7677 | 3765.32 | 10.5546 | 3565 | 1.0939 | 10782.78 | 0.087127
[5 empty] | 3570 | 25.7689 | 3770.96 | 10.5557 | 3570 | 1.0936 | 10810.84 | 0.088888
[5 empty] | 3575 | 25.7718 | 3776.6 | 10.5568 | 3575 | 1.0932 | 10838.9 | 0.090936
[5 empty] | 3580 | 25.7735 | 3782.24 | 10.5579 | 3580 | 1.093 | 10866.96 | 0.093222
[5 empty] | 3585 | 25.7732 | 3787.88 | 10.559 | 3585 | 1.0929 | 10895.02 | 0.095681
[5 empty] | 3590 | 25.777 | 3793.52 | 10.56 | 3590 | 1.093 | 10923.08 | 0.09824
[5 empty] | 3595 | 25.7803 | 3799.16 | 10.561 | 3595 | 1.0934 | 10951.14 | 0.10082
[5 empty] | 3600 | 25.7802 | 3804.8 | 10.5619 | 3600 | 1.0941 | 10979.2 | 0.10334
[5 empty] | 3605 | 25.7794 | 3810.44 | 10.5627 | 3605 | 1.095 | 11007.26 | 0.10572
[5 empty] | 3610 | 25.7802 | 3816.08 | 10.5634 | 3610 | 1.0962 | 11035.32 | 0.1079
[5 empty] | 3615 | 25.7838 | 3821.72 | 10.564 | 3615 | 1.0973 | 11063.38 | 0.10981
[5 empty] | 3620 | 25.7891 | 3827.36 | 10.5645 | 3620 | 1.0984 | 11091.44 | 0.11143
[5 empty] | 3625 | 25.7924 | 3833 | 10.5651 | 3625 | 1.0993 | 11119.5 | 0.11272
[5 empty] | 3630 | 25.7914 | 3838.64 | 10.5657 | 3630 | 1.0998 | 11147.56 | 0.1137
[5 empty] | 3635 | 25.7903 | 3844.28 | 10.5664 | 3635 | 1.1 | 11175.62 | 0.11437
[5 empty] | 3640 | 25.7917 | 3849.92 | 10.5672 | 3640 | 1.0998 | 11203.68 | 0.11478
[5 empty] | 3645 | 25.7954 | 3855.56 | 10.5682 | 3645 | 1.0994 | 11231.74 | 0.11497
[5 empty] | 3650 | 25.7993 | 3861.2 | 10.5692 | 3650 | 1.0988 | 11259.8 | 0.11502
[5 empty] | 3655 | 25.8012 | 3866.84 | 10.5703 | 3655 | 1.0983 | 11287.86 | 0.11499
[5 empty] | 3660 | 25.8014 | 3872.48 | 10.5714 | 3660 | 1.0978 | 11315.92 | 0.11495
[5 empty] | 3665 | 25.8014 | 3878.12 | 10.5726 | 3665 | 1.0974 | 11343.98 | 0.11498
[5 empty] | 3670 | 25.8018 | 3883.76 | 10.5737 | 3670 | 1.0971 | 11372.04 | 0.11513
[5 empty] | 3675 | 25.8012 | 3889.4 | 10.5748 | 3675 | 1.097 | 11400.1 | 0.11545
[5 empty] | 3680 | 25.8014 | 3895.04 | 10.5758 | 3680 | 1.097 | 11428.16 | 0.11598
[5 empty] | 3685 | 25.8038 | 3900.68 | 10.5768 | 3685 | 1.0972 | 11456.22 | 0.11675
[5 empty] | 3690 | 25.8046 | 3906.32 | 10.5778 | 3690 | 1.0974 | 11484.28 | 0.11775
[5 empty] | 3695 | 25.8041 | 3911.96 | 10.5786 | 3695 | 1.0976 | 11512.34 | 0.11897
[5 empty] | 3700 | 25.8041 | 3917.6 | 10.5795 | 3700 | 1.0978 | 11540.4 | 0.12038
[5 empty] | 3705 | 25.8049 | 3923.24 | 10.5803 | 3705 | 1.098 | 11568.46 | 0.12195
[5 empty] | 3710 | 25.8054 | 3928.88 | 10.5811 | 3710 | 1.0982 | 11596.52 | 0.12362
[5 empty] | 3715 | 25.8047 | 3934.52 | 10.5819 | 3715 | 1.0982 | 11624.58 | 0.12534
[5 empty] | 3720 | 25.8031 | 3940.16 | 10.5827 | 3720 | 1.0983 | 11652.64 | 0.12704
[5 empty] | 3725 | 25.805 | 3945.8 | 10.5836 | 3725 | 1.0983 | 11680.7 | 0.12866
[5 empty] | 3730 | 25.81 | 3951.44 | 10.5844 | 3730 | 1.0983 | 11708.76 | 0.13015
[5 empty] | 3735 | 25.8137 | 3957.08 | 10.5851 | 3735 | 1.0985 | 11736.82 | 0.13147
[5 empty] | 3740 | 25.816 | 3962.72 | 10.5858 | 3740 | 1.0988 | 11764.88 | 0.13257
[5 empty] | 3745 | 25.8173 | 3968.36 | 10.5865 | 3745 | 1.0992 | 11792.94 | 0.13343
[5 empty] | 3750 | 25.8179 | 3974 | 10.5871 | 3750 | 1.0998 | 11821 | 0.13404
[5 empty] | 3755 | 25.8176 | 3979.64 | 10.5877 | 3755 | 1.1004 | 11849.06 | 0.13439
[5 empty] | 3760 | 25.817 | 3985.28 | 10.5883 | 3760 | 1.1011 | 11877.12 | 0.13449
[5 empty] | 3765 | 25.8163 | 3990.92 | 10.5889 | 3765 | 1.1018 | 11905.18 | 0.13436
[5 empty] | 3770 | 25.8179 | 3996.56 | 10.5895 | 3770 | 1.1024 | 11933.24 | 0.13404
[5 empty] | 3775 | 25.8223 | 4002.2 | 10.5902 | 3775 | 1.1028 | 11961.3 | 0.13354
[5 empty] | 3780 | 25.8264 | 4007.84 | 10.591 | 3780 | 1.1031 | 11989.36 | 0.1329
[5 empty] | 3785 | 25.8304 | 4013.48 | 10.5918 | 3785 | 1.1032 | 12017.42 | 0.13216
[5 empty] | 3790 | 25.8341 | 4019.12 | 10.5928 | 3790 | 1.1032 | 12045.48 | 0.13135
[5 empty] | 3795 | 25.8365 | 4024.76 | 10.5938 | 3795 | 1.103 | 12073.54 | 0.1305
[5 empty] | 3800 | 25.8374 | 4030.4 | 10.5948 | 3800 | 1.1026 | 12101.6 | 0.12963
[5 empty] | 3805 | 25.8375 | 4036.04 | 10.5957 | 3805 | 1.1022 | 12129.66 | 0.12875
[5 empty] | 3810 | 25.8369 | 4041.68 | 10.5966 | 3810 | 1.1018 | 12157.72 | 0.12789
[5 empty] | 3815 | 25.836 | 4047.32 | 10.5974 | 3815 | 1.1015 | 12185.78 | 0.12703
[5 empty] | 3820 | 25.8358 | 4052.96 | 10.5981 | 3820 | 1.1013 | 12213.84 | 0.12618
[5 empty] | 3825 | 25.8357 | 4058.6 | 10.5988 | 3825 | 1.1013 | 12241.9 | 0.12534
[5 empty] | 3830 | 25.8361 | 4064.24 | 10.5993 | 3830 | 1.1016 | 12269.96 | 0.1245
[5 empty] | 3835 | 25.8372 | 4069.88 | 10.5999 | 3835 | 1.102 | 12298.02 | 0.12365
[5 empty] | 3840 | 25.8386 | 4075.52 | 10.6004 | 3840 | 1.1027 | 12326.08 | 0.1228
[5 empty] | 3845 | 25.8402 | 4081.16 | 10.6011 | 3845 | 1.1035 | 12354.14 | 0.12193
[5 empty] | 3850 | 25.8419 | 4086.8 | 10.6019 | 3850 | 1.1044 | 12382.2 | 0.12105
[5 empty] | 3855 | 25.843 | 4092.44 | 10.6028 | 3855 | 1.1053 | 12410.26 | 0.12017
[5 empty] | 3860 | 25.8436 | 4098.08 | 10.6038 | 3860 | 1.106 | 12438.32 | 0.1193
[5 empty] | 3865 | 25.8445 | 4103.72 | 10.605 | 3865 | 1.1066 | 12466.38 | 0.11845
[5 empty] | 3870 | 25.8452 | 4109.36 | 10.6062 | 3870 | 1.107 | 12494.44 | 0.11765
[5 empty] | 3875 | 25.8453 | 4115 | 10.6076 | 3875 | 1.1073 | 12522.5 | 0.1169
[5 empty] | 3880 | 25.8451 | 4120.64 | 10.6089 | 3880 | 1.1075 | 12550.56 | 0.11624
[5 empty] | 3885 | 25.846 | 4126.28 | 10.6101 | 3885 | 1.1075 | 12578.62 | 0.11568
[5 empty] | 3890 | 25.8483 | 4131.92 | 10.6112 | 3890 | 1.1075 | 12606.68 | 0.11525
[5 empty] | 3895 | 25.8507 | 4137.56 | 10.612 | 3895 | 1.1075 | 12634.74 | 0.11494
[5 empty] | 3900 | 25.8507 | 4143.2 | 10.6127 | 3900 | 1.1074 | 12662.8 | 0.11477
[5 empty] | 3905 | 25.8497 | 4148.84 | 10.6131 | 3905 | 1.1073 | 12690.86 | 0.11474
[5 empty] | 3910 | 25.852 | 4154.48 | 10.6134 | 3910 | 1.1072 | 12718.92 | 0.11483
[5 empty] | 3915 | 25.8564 | 4160.12 | 10.6137 | 3915 | 1.1072 | 12746.98 | 0.11504
[5 empty] | 3920 | 25.8609 | 4165.76 | 10.6142 | 3920 | 1.1071 | 12775.04 | 0.11533
[5 empty] | 3925 | 25.8648 | 4171.4 | 10.6151 | 3925 | 1.1072 | 12803.1 | 0.11567
[5 empty] | 3930 | 25.8665 | 4177.04 | 10.6167 | 3930 | 1.1073 | 12831.16 | 0.11604
[5 empty] | 3935 | 25.8651 | 4182.68 | 10.619 | 3935 | 1.1074 | 12859.22 | 0.11638
[5 empty] | 3940 | 25.8622 | 4188.32 | 10.6222 | 3940 | 1.1076 | 12887.28 | 0.11666
[5 empty] | 3945 | 25.861 | 4193.96 | 10.626 | 3945 | 1.1079 | 12915.34 | 0.11683
[5 empty] | 3950 | 25.8629 | 4199.6 | 10.6304 | 3950 | 1.1081 | 12943.4 | 0.11685
[5 empty] | 3955 | 25.8651 | 4205.24 | 10.6348 | 3955 | 1.1083 | 12971.46 | 0.11669
[5 empty] | 3960 | 25.8659 | 4210.88 | 10.6388 | 3960 | 1.1085 | 12999.52 | 0.11631
[5 empty] | 3965 | 25.8668 | 4216.52 | 10.6419 | 3965 | 1.1087 | 13027.58 | 0.11569
[5 empty] | 3970 | 25.8686 | 4222.16 | 10.6438 | 3970 | 1.1089 | 13055.64 | 0.11481
[5 empty] | 3975 | 25.8707 | 4227.8 | 10.6441 | 3975 | 1.1091 | 13083.7 | 0.11368
[5 empty] | 3980 | 25.8717 | 4233.44 | 10.6427 | 3980 | 1.1093 | 13111.76 | 0.11229
[5 empty] | 3985 | 25.8713 | 4239.08 | 10.64 | 3985 | 1.1095 | 13139.82 | 0.11066
[5 empty] | 3990 | 25.8706 | 4244.72 | 10.6361 | 3990 | 1.1096 | 13167.88 | 0.1088
[5 empty] | 3995 | 25.8711 | 4250.36 | 10.6317 | 3995 | 1.1098 | 13195.94 | 0.10676
[5 empty] | 4000 | 25.8722 | 4256 | 10.6272 | 4000 | 1.1099 | 13224 | 0.10457
[5 empty] | 4005 | 25.8731 | 4261.64 | 10.6233 | 4005 | 1.1099 | 13252.06 | 0.10227
[5 empty] | 4010 | 25.874 | 4267.28 | 10.6202 | 4010 | 1.1099 | 13280.12 | 0.099911
[5 empty] | 4015 | 25.8742 | 4272.92 | 10.6182 | 4015 | 1.1098 | 13308.18 | 0.097539
[5 empty] | 4020 | 25.8742 | 4278.56 | 10.6173 | 4020 | 1.1098 | 13336.24 | 0.095202
[5 empty] | 4025 | 25.8752 | 4284.2 | 10.6174 | 4025 | 1.1098 | 13364.3 | 0.092946
[5 empty] | 4030 | 25.8787 | 4289.84 | 10.6182 | 4030 | 1.1098 | 13392.36 | 0.090815
[5 empty] | 4035 | 25.8837 | 4295.48 | 10.6193 | 4035 | 1.11 | 13420.42 | 0.088845
[5 empty] | 4040 | 25.8865 | 4301.12 | 10.6205 | 4040 | 1.1102 | 13448.48 | 0.08707
[5 empty] | 4045 | 25.8851 | 4306.76 | 10.6216 | 4045 | 1.1104 | 13476.54 | 0.085516
[5 empty] | 4050 | 25.8829 | 4312.4 | 10.6223 | 4050 | 1.1106 | 13504.6 | 0.084202
[5 empty] | 4055 | 25.8822 | 4318.04 | 10.6227 | 4055 | 1.1108 | 13532.66 | 0.083144
[5 empty] | 4060 | 25.883 | 4323.68 | 10.623 | 4060 | 1.111 | 13560.72 | 0.082347
[5 empty] | 4065 | 25.8845 | 4329.32 | 10.6232 | 4065 | 1.111 | 13588.78 | 0.081812
[5 empty] | 4070 | 25.8862 | 4334.96 | 10.6236 | 4070 | 1.1111 | 13616.84 | 0.081537
[5 empty] | 4075 | 25.8881 | 4340.6 | 10.6243 | 4075 | 1.1112 | 13644.9 | 0.08151
[5 empty] | 4080 | 25.8905 | 4346.24 | 10.6256 | 4080 | 1.1114 | 13672.96 | 0.081713
[5 empty] | 4085 | 25.8923 | 4351.88 | 10.6274 | 4085 | 1.1117 | 13701.02 | 0.082141
[5 empty] | 4090 | 25.8916 | 4357.52 | 10.6297 | 4090 | 1.1121 | 13729.08 | 0.082742
[5 empty] | 4095 | 25.891 | 4363.16 | 10.6324 | 4095 | 1.1126 | 13757.14 | 0.08352
[5 empty] | 4100 | 25.8926 | 4368.8 | 10.6353 | 4100 | 1.1132 | 13785.2 | 0.084459
[5 empty] | 4105 | 25.8954 | 4374.44 | 10.6382 | 4105 | 1.1136 | 13813.26 | 0.0854
[5 empty] | 4110 | 25.8973 | 4380.08 | 10.641 | 4110 | 1.1138 | 13841.32 | 0.0866
[5 empty] | 4115 | 25.8972 | 4385.72 | 10.6436 | 4115 | 1.1138 | 13869.38 | 0.0877
[5 empty] | 4120 | 25.8967 | 4391.36 | 10.6457 | 4120 | 1.1135 | 13897.44 | 0.0886
[5 empty] | 4125 | 25.898 | 4397 | 10.6474 | 4125 | 1.113
[5 empty] | 4130 | 25.8998 | 4402.64 | 10.6486 | 4130 | 1.1124
[5 empty] | 4135 | 25.8987 | 4408.28 | 10.6495 | 4135 | 1.1117
[5 empty] | 4140 | 25.897 | 4413.92 | 10.6501 | 4140 | 1.1111
[5 empty] | 4145 | 25.8974 | 4419.56 | 10.6505 | 4145 | 1.1107
[5 empty] | 4150 | 25.8999 | 4425.2 | 10.6507 | 4150 | 1.1105
[5 empty] | 4155 | 25.9028 | 4430.84 | 10.6509 | 4155 | 1.1106
[5 empty] | 4160 | 25.9047 | 4436.48 | 10.6511 | 4160 | 1.111
[5 empty] | 4165 | 25.9059 | 4442.12 | 10.6513 | 4165 | 1.1115
[5 empty] | 4170 | 25.9074 | 4447.76 | 10.6516 | 4170 | 1.112
[5 empty] | 4175 | 25.9097 | 4453.4 | 10.652 | 4175 | 1.1126
[5 empty] | 4180 | 25.9125 | 4459.04 | 10.6525 | 4180 | 1.113
[5 empty] | 4185 | 25.9151 | 4464.68 | 10.6531 | 4185 | 1.1132
[5 empty] | 4190 | 25.9169 | 4470.32 | 10.6538 | 4190 | 1.1133
[5 empty] | 4195 | 25.9165 | 4475.96 | 10.6546 | 4195 | 1.1132
[5 empty] | 4200 | 25.9147 | 4481.6 | 10.6555 | 4200 | 1.113
[5 empty] | 4205 | 25.9155 | 4487.24 | 10.6564 | 4205 | 1.1127
[5 empty] | 4210 | 25.9244 | 4492.88 | 10.6573 | 4210 | 1.1126
[5 empty] | 4215 | 25.9411 | 4498.52 | 10.6582 | 4215 | 1.1126
[5 empty] | 4220 | 25.9509 | 4504.16 | 10.659 | 4220 | 1.1127
[5 empty] | 4225 | 25.9368 | 4509.8 | 10.6598 | 4225 | 1.1131
[5 empty] | 4230 | 25.9236 | 4515.44 | 10.6605 | 4230 | 1.1136
[5 empty] | 4235 | 25.9262 | 4521.08 | 10.6612 | 4235 | 1.1141
[5 empty] | 4240 | 25.9288 | 4526.72 | 10.6618 | 4240 | 1.1146
[5 empty] | 4245 | 25.9236 | 4532.36 | 10.6623 | 4245 | 1.115
[5 empty] | 4250 | 25.9181 | 4538 | 10.6628 | 4250 | 1.1152
[5 empty] | 4255 | 25.9205 | 4543.64 | 10.6633 | 4255 | 1.1153
[5 empty] | 4260 | 25.9253 | 4549.28 | 10.6638 | 4260 | 1.1153
[5 empty] | 4265 | 25.9247 | 4554.92 | 10.6643 | 4265 | 1.1153
[5 empty] | 4270 | 25.9223 | 4560.56 | 10.6648 | 4270 | 1.1153
[5 empty] | 4275 | 25.925 | 4566.2 | 10.6653 | 4275 | 1.1153
[5 empty] | 4280 | 25.9292 | 4571.84 | 10.6657 | 4280 | 1.1154
[5 empty] | 4285 | 25.9293 | 4577.48 | 10.6661 | 4285 | 1.1155
[5 empty] | 4290 | 25.9246 | 4583.12 | 10.6664 | 4290 | 1.1156
[5 empty] | 4295 | 25.9196 | 4588.76 | 10.6666 | 4295 | 1.1157
[5 empty] | 4300 | 25.9187 | 4594.4 | 10.6667 | 4300 | 1.1158
[5 empty] | 4305 | 25.9218 | 4600.04 | 10.6669 | 4305 | 1.116
[5 empty] | 4310 | 25.9244 | 4605.68 | 10.6671 | 4310 | 1.1161
[5 empty] | 4315 | 25.9227 | 4611.32 | 10.6673 | 4315 | 1.1164
[5 empty] | 4320 | 25.9214 | 4616.96 | 10.6676 | 4320 | 1.1167
[5 empty] | 4325 | 25.9243 | 4622.6 | 10.6681 | 4325 | 1.1171
[5 empty] | 4330 | 25.9297 | 4628.24 | 10.6687 | 4330 | 1.1175
[5 empty] | 4335 | 25.9335 | 4633.88 | 10.6694 | 4335 | 1.1179
[5 empty] | 4340 | 25.9342 | 4639.52 | 10.6702 | 4340 | 1.1182
[5 empty] | 4345 | 25.9335 | 4645.16 | 10.671 | 4345 | 1.1184
[5 empty] | 4350 | 25.934 | 4650.8 | 10.672 | 4350 | 1.1186
[5 empty] | 4355 | 25.9356 | 4656.44 | 10.6729 | 4355 | 1.1186
[5 empty] | 4360 | 25.9369 | 4662.08 | 10.6738 | 4360 | 1.1186
[5 empty] | 4365 | 25.9381 | 4667.72 | 10.6747 | 4365 | 1.1185
[5 empty] | 4370 | 25.9389 | 4673.36 | 10.6756 | 4370 | 1.1184
[5 empty] | 4375 | 25.9388 | 4679 | 10.6765 | 4375 | 1.1184
[5 empty] | 4380 | 25.9385 | 4684.64 | 10.6774 | 4380 | 1.1184
[5 empty] | 4385 | 25.9389 | 4690.28 | 10.6783 | 4385 | 1.1185
[5 empty] | 4390 | 25.9407 | 4695.92 | 10.6792 | 4390 | 1.1186
[5 empty] | 4395 | 25.9434 | 4701.56 | 10.6801 | 4395 | 1.1187
[5 empty] | 4400 | 25.9462 | 4707.2 | 10.6811 | 4400 | 1.119
[5 empty] | 4405 | 25.9489 | 4712.84 | 10.6819 | 4405 | 1.1192
[5 empty] | 4410 | 25.9506 | 4718.48 | 10.6827 | 4410 | 1.1195
[5 empty] | 4415 | 25.9508 | 4724.12 | 10.6833 | 4415 | 1.1199
[5 empty] | 4420 | 25.9497 | 4729.76 | 10.6839 | 4420 | 1.1204
[5 empty] | 4425 | 25.9484 | 4735.4 | 10.6843 | 4425 | 1.1208
[5 empty] | 4430 | 25.9482 | 4741.04 | 10.6845 | 4430 | 1.1214
[5 empty] | 4435 | 25.9497 | 4746.68 | 10.6846 | 4435 | 1.1219
[5 empty] | 4440 | 25.9522 | 4752.32 | 10.6846 | 4440 | 1.1223
[5 empty] | 4445 | 25.9538 | 4757.96 | 10.6846 | 4445 | 1.1227
[5 empty] | 4450 | 25.9533 | 4763.6 | 10.6845 | 4450 | 1.1229
[5 empty] | 4455 | 25.9521 | 4769.24 | 10.6844 | 4455 | 1.1229
[5 empty] | 4460 | 25.9522 | 4774.88 | 10.6844 | 4460 | 1.1228
[5 empty] | 4465 | 25.9538 | 4780.52 | 10.6845 | 4465 | 1.1226
[5 empty] | 4470 | 25.956 | 4786.16 | 10.6847 | 4470 | 1.1222
[5 empty] | 4475 | 25.9577 | 4791.8 | 10.6851 | 4475 | 1.1219
[5 empty] | 4480 | 25.9586 | 4797.44 | 10.6855 | 4480 | 1.1216
[5 empty] | 4485 | 25.9592 | 4803.08 | 10.6861 | 4485 | 1.1214
[5 empty] | 4490 | 25.96 | 4808.72 | 10.6868 | 4490 | 1.1214
[5 empty] | 4495 | 25.9608 | 4814.36 | 10.6876 | 4495 | 1.1216
[5 empty] | 4500 | 25.9612 | 4820 | 10.6883 | 4500 | 1.122
[5 empty] | 4505 | 25.9615 | 4825.64 | 10.6891 | 4505 | 1.1226
[5 empty] | 4510 | 25.9621 | 4831.28 | 10.6898 | 4510 | 1.1233
[5 empty] | 4515 | 25.963 | 4836.92 | 10.6904 | 4515 | 1.1239
[5 empty] | 4520 | 25.9639 | 4842.56 | 10.6908 | 4520 | 1.1244
[5 empty] | 4525 | 25.9642 | 4848.2 | 10.6911 | 4525 | 1.1248
[5 empty] | 4530 | 25.9632 | 4853.84 | 10.6913 | 4530 | 1.125
[5 empty] | 4535 | 25.9615 | 4859.48 | 10.6914 | 4535 | 1.1251
[5 empty] | 4540 | 25.9605 | 4865.12 | 10.6913 | 4540 | 1.1249
[5 empty] | 4545 | 25.9616 | 4870.76 | 10.6913 | 4545 | 1.1245
[5 empty] | 4550 | 25.9637 | 4876.4 | 10.6912 | 4550 | 1.1241
[5 empty] | 4555 | 25.9653 | 4882.04 | 10.6913 | 4555 | 1.1235
[5 empty] | 4560 | 25.9654 | 4887.68 | 10.6915 | 4560 | 1.123
[5 empty] | 4565 | 25.9644 | 4893.32 | 10.6919 | 4565 | 1.1225
[5 empty] | 4570 | 25.9633 | 4898.96 | 10.6927 | 4570 | 1.1221
[5 empty] | 4575 | 25.9635 | 4904.6 | 10.6937 | 4575 | 1.1218
[5 empty] | 4580 | 25.9654 | 4910.24 | 10.6949 | 4580 | 1.1215
[5 empty] | 4585 | 25.9692 | 4915.88 | 10.6964 | 4585 | 1.1214
[5 empty] | 4590 | 25.9737 | 4921.52 | 10.698 | 4590 | 1.1215
[5 empty] | 4595 | 25.9766 | 4927.16 | 10.6997 | 4595 | 1.1216
[5 empty] | 4600 | 25.9771 | 4932.8 | 10.7013 | 4600 | 1.1218
[5 empty] | 4605 | 25.9756 | 4938.44 | 10.7029 | 4605 | 1.1221
[5 empty] | 4610 | 25.974 | 4944.08 | 10.7043 | 4610 | 1.1224
[5 empty] | 4615 | 25.974 | 4949.72 | 10.7055 | 4615 | 1.1227
[5 empty] | 4620 | 25.9757 | 4955.36 | 10.7063 | 4620 | 1.1229
[5 empty] | 4625 | 25.978 | 4961 | 10.7069 | 4625 | 1.1232
[5 empty] | 4630 | 25.9798 | 4966.64 | 10.7071 | 4630 | 1.1233
[5 empty] | 4635 | 25.981 | 4972.28 | 10.707 | 4635 | 1.1234
[5 empty] | 4640 | 25.9817 | 4977.92 | 10.7064 | 4640 | 1.1234
[5 empty] | 4645 | 25.9819 | 4983.56 | 10.7056 | 4645 | 1.1234
[5 empty] | 4650 | 25.9818 | 4989.2 | 10.7046 | 4650 | 1.1234
[5 empty] | 4655 | 25.9819 | 4994.84 | 10.7034 | 4655 | 1.1235
[5 empty] | 4660 | 25.9823 | 5000.48 | 10.7021 | 4660 | 1.1235
[5 empty] | 4665 | 25.9833 | 5006.12 | 10.701 | 4665 | 1.1237
[5 empty] | 4670 | 25.9847 | 5011.76 | 10.7001 | 4670 | 1.1239
[5 empty] | 4675 | 25.986 | 5017.4 | 10.6996 | 4675 | 1.1242
[5 empty] | 4680 | 25.9865 | 5023.04 | 10.6996 | 4680 | 1.1246
[5 empty] | 4685 | 25.9864 | 5028.68 | 10.7 | 4685 | 1.1251
[5 empty] | 4690 | 25.9861 | 5034.32 | 10.7009 | 4690 | 1.1256
[5 empty] | 4695 | 25.9863 | 5039.96 | 10.7023 | 4695 | 1.1261
[5 empty] | 4700 | 25.987 | 5045.6 | 10.7039 | 4700 | 1.1265
[5 empty] | 4705 | 25.9882 | 5051.24 | 10.7057 | 4705 | 1.1269
[5 empty] | 4710 | 25.9899 | 5056.88 | 10.7075 | 4710 | 1.1272
[5 empty] | 4715 | 25.9917 | 5062.52 | 10.7093 | 4715 | 1.1274
[5 empty] | 4720 | 25.993 | 5068.16 | 10.7109 | 4720 | 1.1274
[5 empty] | 4725 | 25.9929 | 5073.8 | 10.7124 | 4725 | 1.1273
[5 empty] | 4730 | 25.9923 | 5079.44 | 10.7137 | 4730 | 1.1271
[5 empty] | 4735 | 25.992 | 5085.08 | 10.715 | 4735 | 1.1268
[5 empty] | 4740 | 25.9924 | 5090.72 | 10.7161 | 4740 | 1.1265
[5 empty] | 4745 | 25.9934 | 5096.36 | 10.7173 | 4745 | 1.1262
[5 empty] | 4750 | 25.9946 | 5102 | 10.7187 | 4750 | 1.1261
[5 empty] | 4755 | 25.9959 | 5107.64 | 10.7202 | 4755 | 1.126
[5 empty] | 4760 | 25.9974 | 5113.28 | 10.7218 | 4760 | 1.1261
[5 empty] | 4765 | 25.9992 | 5118.92 | 10.7235 | 4765 | 1.1263
[5 empty] | 4770 | 26.0009 | 5124.56 | 10.7252 | 4770 | 1.1267
[5 empty] | 4775 | 26.0023 | 5130.2 | 10.7266 | 4775 | 1.1272
[5 empty] | 4780 | 26.0031 | 5135.84 | 10.7277 | 4780 | 1.1278
[5 empty] | 4785 | 26.0036 | 5141.48 | 10.7282 | 4785 | 1.1284
[5 empty] | 4790 | 26.0035 | 5147.12 | 10.7279 | 4790 | 1.1289
[5 empty] | 4795 | 26.0029 | 5152.76 | 10.7267 | 4795 | 1.1293
[5 empty] | 4800 | 26.002 | 5158.4 | 10.7245 | 4800 | 1.1295
[5 empty] | 4805 | 26.001 | 5164.04 | 10.7213 | 4805 | 1.1295
[5 empty] | 4810 | 26.0006 | 5169.68 | 10.7173 | 4810 | 1.1294
[5 empty] | 4815 | 26.001 | 5175.32 | 10.7127 | 4815 | 1.1291
[5 empty] | 4820 | 26.0018 | 5180.96 | 10.708 | 4820 | 1.1287
[5 empty] | 4825 | 26.0023 | 5186.6 | 10.7034 | 4825 | 1.1282
[5 empty] | 4830 | 26.0018 | 5192.24 | 10.6994 | 4830 | 1.1278
[5 empty] | 4835 | 26.0005 | 5197.88 | 10.6963 | 4835 | 1.1274
[5 empty] | 4840 | 25.9988 | 5203.52 | 10.6943 | 4840 | 1.127
[5 empty] | 4845 | 25.9981 | 5209.16 | 10.6937 | 4845 | 1.1267
[5 empty] | 4850 | 25.9992 | 5214.8 | 10.6943 | 4850 | 1.1265
[5 empty] | 4855 | 26.0021 | 5220.44 | 10.696 | 4855 | 1.1263
[5 empty] | 4860 | 26.0054 | 5226.08 | 10.6986 | 4860 | 1.1263
[5 empty] | 4865 | 26.0076 | 5231.72 | 10.7018 | 4865 | 1.1263
[5 empty] | 4870 | 26.0082 | 5237.36 | 10.7053 | 4870 | 1.1264
[5 empty] | 4875 | 26.0078 | 5243 | 10.7088 | 4875 | 1.1266
[5 empty] | 4880 | 26.0072 | 5248.64 | 10.7122 | 4880 | 1.1268
[5 empty] | 4885 | 26.0069 | 5254.28 | 10.7155 | 4885 | 1.1272
[5 empty] | 4890 | 26.0074 | 5259.92 | 10.7185 | 4890 | 1.1276
[5 empty] | 4895 | 26.0088 | 5265.56 | 10.7214 | 4895 | 1.128
[5 empty] | 4900 | 26.0107 | 5271.2 | 10.7242 | 4900 | 1.1286
[5 empty] | 4905 | 26.0132 | 5276.84 | 10.7269 | 4905 | 1.1291
[5 empty] | 4910 | 26.0159 | 5282.48 | 10.7296 | 4910 | 1.1296
[5 empty] | 4915 | 26.0182 | 5288.12 | 10.7322 | 4915 | 1.13
[5 empty] | 4920 | 26.0196 | 5293.76 | 10.7344 | 4920 | 1.1304
[5 empty] | 4925 | 26.0198 | 5299.4 | 10.7363 | 4925 | 1.1307
[5 empty] | 4930 | 26.0191 | 5305.04 | 10.7375 | 4930 | 1.1309
[5 empty] | 4935 | 26.0185 | 5310.68 | 10.738 | 4935 | 1.1311
[5 empty] | 4940 | 26.0186 | 5316.32 | 10.7377 | 4940 | 1.1313
[5 empty] | 4945 | 26.0194 | 5321.96 | 10.7366 | 4945 | 1.1315
[5 empty] | 4950 | 26.02 | 5327.6 | 10.7348 | 4950 | 1.1317
[5 empty] | 4955 | 26.0196 | 5333.24 | 10.7325 | 4955 | 1.132
[5 empty] | 4960 | 26.0188 | 5338.88 | 10.73 | 4960 | 1.1323
[5 empty] | 4965 | 26.0182 | 5344.52 | 10.7274 | 4965 | 1.1327
[5 empty] | 4970 | 26.0182 | 5350.16 | 10.7251 | 4970 | 1.1331
[5 empty] | 4975 | 26.0188 | 5355.8 | 10.7233 | 4975 | 1.1335
[5 empty] | 4980 | 26.0197 | 5361.44 | 10.7223 | 4980 | 1.1337
[5 empty] | 4985 | 26.0207 | 5367.08 | 10.722 | 4985 | 1.1339
[5 empty] | 4990 | 26.0222 | 5372.72 | 10.7225 | 4990 | 1.1338
[5 empty] | 4995 | 26.024 | 5378.36 | 10.7236 | 4995 | 1.1336
[5 empty] | 5000 | 26.0255 | 5384 | 10.7253 | 5000 | 1.1333
[5 empty] | 5005 | 26.0259 | 5389.64 | 10.7272 | 5005 | 1.1328
[5 empty] | 5010 | 26.0248 | 5395.28 | 10.7292 | 5010 | 1.1323
[5 empty] | 5015 | 26.023 | 5400.92 | 10.7311 | 5015 | 1.1317
[5 empty] | 5020 | 26.0215 | 5406.56 | 10.7328 | 5020 | 1.1312
[5 empty] | 5025 | 26.0215 | 5412.2 | 10.7343 | 5025 | 1.1308
[5 empty] | 5030 | 26.0231 | 5417.84 | 10.7356 | 5030 | 1.1305
[5 empty] | 5035 | 26.0256 | 5423.48 | 10.7369 | 5035 | 1.1304
[5 empty] | 5040 | 26.0276 | 5429.12 | 10.7382 | 5040 | 1.1305
[5 empty] | 5045 | 26.0292 | 5434.76 | 10.7398 | 5045 | 1.1308
[5 empty] | 5050 | 26.0306 | 5440.4 | 10.7417 | 5050 | 1.1313
[5 empty] | 5055 | 26.0322 | 5446.04 | 10.7438 | 5055 | 1.1319
[5 empty] | 5060 | 26.0339 | 5451.68 | 10.7461 | 5060 | 1.1325
[5 empty] | 5065 | 26.0357 | 5457.32 | 10.7482 | 5065 | 1.1332
[5 empty] | 5070 | 26.0376 | 5462.96 | 10.7499 | 5070 | 1.1339
[5 empty] | 5075 | 26.0398 | 5468.6 | 10.751 | 5075 | 1.1345
[5 empty] | 5080 | 26.0421 | 5474.24 | 10.7512 | 5080 | 1.135
[5 empty] | 5085 | 26.0435 | 5479.88 | 10.7505 | 5085 | 1.1354
[5 empty] | 5090 | 26.0436 | 5485.52 | 10.7488 | 5090 | 1.1356
[5 empty] | 5095 | 26.0427 | 5491.16 | 10.7463 | 5095 | 1.1358
[5 empty] | 5100 | 26.0415 | 5496.8 | 10.7433 | 5100 | 1.1358
[5 empty] | 5105 | 26.0407 | 5502.44 | 10.7399 | 5105 | 1.1358
[5 empty] | 5110 | 26.0403 | 5508.08 | 10.7367 | 5110 | 1.1358
[5 empty] | 5115 | 26.0403 | 5513.72 | 10.7336 | 5115 | 1.1358
[5 empty] | 5120 | 26.0404 | 5519.36 | 10.7311 | 5120 | 1.1357
[5 empty] | 5125 | 26.0401 | 5525 | 10.7291 | 5125 | 1.1357
[5 empty] | 5130 | 26.0393 | 5530.64 | 10.7277 | 5130 | 1.1358
[5 empty] | 5135 | 26.0384 | 5536.28 | 10.7269 | 5135 | 1.1358
[5 empty] | 5140 | 26.0376 | 5541.92 | 10.7266 | 5140 | 1.1358
[5 empty] | 5145 | 26.0375 | 5547.56 | 10.7266 | 5145 | 1.1357
[5 empty] | 5150 | 26.0382 | 5553.2 | 10.7269 | 5150 | 1.1356
[5 empty] | 5155 | 26.04 | 5558.84 | 10.7274 | 5155 | 1.1355
[5 empty] | 5160 | 26.0425 | 5564.48 | 10.7278 | 5160 | 1.1353
[5 empty] | 5165 | 26.0453 | 5570.12 | 10.7281 | 5165 | 1.1351
[5 empty] | 5170 | 26.048 | 5575.76 | 10.7282 | 5170 | 1.1348
[5 empty] | 5175 | 26.0501 | 5581.4 | 10.7278 | 5175 | 1.1346
[5 empty] | 5180 | 26.051 | 5587.04 | 10.727 | 5180 | 1.1345
[5 empty] | 5185 | 26.0506 | 5592.68 | 10.7257 | 5185 | 1.1344
[5 empty] | 5190 | 26.0489 | 5598.32 | 10.7239 | 5190 | 1.1345
[5 empty] | 5195 | 26.0462 | 5603.96 | 10.7217 | 5195 | 1.1346
[5 empty] | 5200 | 26.043 | 5609.6 | 10.7193 | 5200 | 1.1349
[5 empty] | 5205 | 26.0403 | 5615.24 | 10.7167 | 5205 | 1.1352
[5 empty] | 5210 | 26.039 | 5620.88 | 10.7139 | 5210 | 1.1356
[5 empty] | 5215 | 26.0394 | 5626.52 | 10.7109 | 5215 | 1.136
[5 empty] | 5220 | 26.0414 | 5632.16 | 10.7078 | 5220 | 1.1364
[5 empty] | 5225 | 26.0437 | 5637.8 | 10.7044 | 5225 | 1.1367
[5 empty] | 5230 | 26.0456 | 5643.44 | 10.7009 | 5230 | 1.137
[5 empty] | 5235 | 26.047 | 5649.08 | 10.6971 | 5235 | 1.1372
[5 empty] | 5240 | 26.0481 | 5654.72 | 10.6934 | 5240 | 1.1373
[5 empty] | 5245 | 26.049 | 5660.36 | 10.6901 | 5245 | 1.1374
[5 empty] | 5250 | 26.0498 | 5666 | 10.6875 | 5250 | 1.1374
[5 empty] | 5255 | 26.0502 | 5671.64 | 10.6864 | 5255 | 1.1374
[5 empty] | 5260 | 26.0502 | 5677.28 | 10.6871 | 5260 | 1.1374
[5 empty] | 5265 | 26.0496 | 5682.92 | 10.6901 | 5265 | 1.1374
[5 empty] | 5270 | 26.0493 | 5688.56 | 10.6957 | 5270 | 1.1374
[5 empty] | 5275 | 26.0496 | 5694.2 | 10.7039 | 5275 | 1.1375
[5 empty] | 5280 | 26.0508 | 5699.84 | 10.7145 | 5280 | 1.1375
[5 empty] | 5285 | 26.0523 | 5705.48 | 10.7268 | 5285 | 1.1376
[5 empty] | 5290 | 26.0536 | 5711.12 | 10.7401 | 5290 | 1.1376
[5 empty] | 5295 | 26.0541 | 5716.76 | 10.7536 | 5295 | 1.1377
[5 empty] | 5300 | 26.054 | 5722.4 | 10.7664 | 5300 | 1.1377
[5 empty] | 5305 | 26.0537 | 5728.04 | 10.7777 | 5305 | 1.1377
[5 empty] | 5310 | 26.0538 | 5733.68 | 10.7868 | 5310 | 1.1376
[5 empty] | 5315 | 26.0545 | 5739.32 | 10.7934 | 5315 | 1.1374
[5 empty] | 5320 | 26.0556 | 5744.96 | 10.7972 | 5320 | 1.1373
[5 empty] | 5325 | 26.0571 | 5750.6 | 10.7983 | 5325 | 1.1371
[5 empty] | 5330 | 26.0588 | 5756.24 | 10.7967 | 5330 | 1.1369
[5 empty] | 5335 | 26.0607 | 5761.88 | 10.7926 | 5335 | 1.1367
[5 empty] | 5340 | 26.0625 | 5767.52 | 10.7864 | 5340 | 1.1366
[5 empty] | 5345 | 26.0636 | 5773.16 | 10.7782 | 5345 | 1.1365
[5 empty] | 5350 | 26.0635 | 5778.8 | 10.7684 | 5350 | 1.1364
[5 empty] | 5355 | 26.0619 | 5784.44 | 10.7573 | 5355 | 1.1365
[5 empty] | 5360 | 26.0598 | 5790.08 | 10.7452 | 5360 | 1.1365
[5 empty] | 5365 | 26.0578 | 5795.72 | 10.7327 | 5365 | 1.1367
[5 empty] | 5370 | 26.0563 | 5801.36 | 10.7202 | 5370 | 1.1368
[5 empty] | 5375 | 26.0555 | 5807 | 10.7085 | 5375 | 1.137
[5 empty] | 5380 | 26.0556 | 5812.64 | 10.6981 | 5380 | 1.1371
[5 empty] | 5385 | 26.0568 | 5818.28 | 10.6898 | 5385 | 1.1372
[5 empty] | 5390 | 26.0596 | 5823.92 | 10.684 | 5390 | 1.1373
[5 empty] | 5395 | 26.0639 | 5829.56 | 10.6809 | 5395 | 1.1374
[5 empty] | 5400 | 26.0693 | 5835.2 | 10.6808 | 5400 | 1.1374
[5 empty] | 5405 | 26.0746 | 5840.84 | 10.6834 | 5405 | 1.1374
[5 empty] | 5410 | 26.0787 | 5846.48 | 10.6884 | 5410 | 1.1373
[5 empty] | 5415 | 26.0809 | 5852.12 | 10.6951 | 5415 | 1.1373
[5 empty] | 5420 | 26.0813 | 5857.76 | 10.703 | 5420 | 1.1373
[5 empty] | 5425 | 26.0806 | 5863.4 | 10.7114 | 5425 | 1.1373
[5 empty] | 5430 | 26.0795 | 5869.04 | 10.7198 | 5430 | 1.1374
[5 empty] | 5435 | 26.0785 | 5874.68 | 10.7278 | 5435 | 1.1376
[5 empty] | 5440 | 26.0778 | 5880.32 | 10.735 | 5440 | 1.1378
[5 empty] | 5445 | 26.0773 | 5885.96 | 10.7414 | 5445 | 1.1381
[5 empty] | 5450 | 26.077 | 5891.6 | 10.7469 | 5450 | 1.1385
[5 empty] | 5455 | 26.0768 | 5897.24 | 10.7517 | 5455 | 1.1389
[5 empty] | 5460 | 26.0768 | 5902.88 | 10.7559 | 5460 | 1.1393
[5 empty] | 5465 | 26.0765 | 5908.52 | 10.7596 | 5465 | 1.1397
[5 empty] | 5470 | 26.0758 | 5914.16 | 10.7631 | 5470 | 1.14
[5 empty] | 5475 | 26.0747 | 5919.8 | 10.7663 | 5475 | 1.1403
[5 empty] | 5480 | 26.0737 | 5925.44 | 10.7694 | 5480 | 1.1406
[5 empty] | 5485 | 26.0732 | 5931.08 | 10.7723 | 5485 | 1.1408
[5 empty] | 5490 | 26.0734 | 5936.72 | 10.7753 | 5490 | 1.1409
[5 empty] | 5495 | 26.074 | 5942.36 | 10.7781 | 5495 | 1.141
[5 empty] | 5500 | 26.0746 | 5948 | 10.7808 | 5500 | 1.1411
[5 empty] | 5505 | 26.0752 | 5953.64 | 10.7835 | 5505 | 1.1413
[5 empty] | 5510 | 26.0758 | 5959.28 | 10.786 | 5510 | 1.1415
[5 empty] | 5515 | 26.0766 | 5964.92 | 10.7885 | 5515 | 1.1418
[5 empty] | 5520 | 26.0775 | 5970.56 | 10.7909 | 5520 | 1.1421
[5 empty] | 5525 | 26.0782 | 5976.2 | 10.7932 | 5525 | 1.1425
[5 empty] | 5530 | 26.078 | 5981.84 | 10.7955 | 5530 | 1.143
[5 empty] | 5535 | 26.0768 | 5987.48 | 10.7976 | 5535 | 1.1435
[5 empty] | 5540 | 26.0754 | 5993.12 | 10.7996 | 5540 | 1.144
[5 empty] | 5545 | 26.0753 | 5998.76 | 10.8013 | 5545 | 1.1444
[5 empty] | 5550 | 26.0777 | 6004.4 | 10.8024 | 5550 | 1.1449
[5 empty] | 5555 | 26.0826 | 6010.04 | 10.803 | 5555 | 1.1453
[5 empty] | 5560 | 26.0883 | 6015.68 | 10.8027 | 5560 | 1.1456
[5 empty] | 5565 | 26.0929 | 6021.32 | 10.8017 | 5565 | 1.1458
[5 empty] | 5570 | 26.0941 | 6026.96 | 10.8001 | 5570 | 1.1459
[5 empty] | 5575 | 26.0911 | 6032.6 | 10.798 | 5575 | 1.146
[5 empty] | 5580 | 26.0854 | 6038.24 | 10.7959 | 5580 | 1.1459
[5 empty] | 5585 | 26.0801 | 6043.88 | 10.7942 | 5585 | 1.1458
[5 empty] | 5590 | 26.077 | 6049.52 | 10.7935 | 5590 | 1.1455
[5 empty] | 5595 | 26.0756 | 6055.16 | 10.794 | 5595 | 1.1452
[5 empty] | 5600 | 26.0751 | 6060.8 | 10.796 | 5600 | 1.1449
[5 empty] | 5605 | 26.0751 | 6066.44 | 10.7996 | 5605 | 1.1445
[5 empty] | 5610 | 26.0759 | 6072.08 | 10.8043 | 5610 | 1.1441
[5 empty] | 5615 | 26.0779 | 6077.72 | 10.8098 | 5615 | 1.1437
[5 empty] | 5620 | 26.0818 | 6083.36 | 10.8154 | 5620 | 1.1434
[5 empty] | 5625 | 26.087 | 6089 | 10.8204 | 5625 | 1.143
[5 empty] | 5630 | 26.0915 | 6094.64 | 10.8241 | 5630 | 1.1428
[5 empty] | 5635 | 26.0922 | 6100.28 | 10.826 | 5635 | 1.1426
[5 empty] | 5640 | 26.0885 | 6105.92 | 10.8261 | 5640 | 1.1426
[5 empty] | 5645 | 26.0821 | 6111.56 | 10.8243 | 5645 | 1.1426
[5 empty] | 5650 | 26.0765 | 6117.2 | 10.8212 | 5650 | 1.1427
[5 empty] | 5655 | 26.0742 | 6122.84 | 10.8172 | 5655 | 1.1429
[5 empty] | 5660 | 26.0763 | 6128.48 | 10.8131 | 5660 | 1.1431
[5 empty] | 5665 | 26.0818 | 6134.12 | 10.8096 | 5665 | 1.1434
[5 empty] | 5670 | 26.0885 | 6139.76 | 10.8071 | 5670 | 1.1438
[5 empty] | 5675 | 26.0936 | 6145.4 | 10.806 | 5675 | 1.1441
[5 empty] | 5680 | 26.0953 | 6151.04 | 10.8063 | 5680 | 1.1445
[5 empty] | 5685 | 26.0941 | 6156.68 | 10.8077 | 5685 | 1.1448
[5 empty] | 5690 | 26.0914 | 6162.32 | 10.81 | 5690 | 1.1451
[5 empty] | 5695 | 26.0886 | 6167.96 | 10.8125 | 5695 | 1.1453
[5 empty] | 5700 | 26.0871 | 6173.6 | 10.8149 | 5700 | 1.1454
[5 empty] | 5705 | 26.0876 | 6179.24 | 10.8165 | 5705 | 1.1454
[5 empty] | 5710 | 26.0907 | 6184.88 | 10.8173 | 5710 | 1.1453
[5 empty] | 5715 | 26.0962 | 6190.52 | 10.817 | 5715 | 1.1452
[5 empty] | 5720 | 26.1033 | 6196.16 | 10.8158 | 5720 | 1.1451
[5 empty] | 5725 | 26.1097 | 6201.8 | 10.8137 | 5725 | 1.145
[5 empty] | 5730 | 26.1142 | 6207.44 | 10.811 | 5730 | 1.1449
[5 empty] | 5735 | 26.1163 | 6213.08 | 10.808 | 5735 | 1.145
[5 empty] | 5740 | 26.1167 | 6218.72 | 10.805 | 5740 | 1.1452
[5 empty] | 5745 | 26.1165 | 6224.36 | 10.8022 | 5745 | 1.1455
[5 empty] | 5750 | 26.1169 | 6230 | 10.7996 | 5750 | 1.1459
[5 empty] | 5755 | 26.1192 | 6235.64 | 10.7974 | 5755 | 1.1465
[5 empty] | 5760 | 26.1237 | 6241.28 | 10.7957 | 5760 | 1.1471
[5 empty] | 5765 | 26.1289 | 6246.92 | 10.7944 | 5765 | 1.1478
[5 empty] | 5770 | 26.1316 | 6252.56 | 10.7934 | 5770 | 1.1485
[5 empty] | 5775 | 26.1308 | 6258.2 | 10.7929 | 5775 | 1.1491
[5 empty] | 5780 | 26.1266 | 6263.84 | 10.7926 | 5780 | 1.1496
[5 empty] | 5785 | 26.1203 | 6269.48 | 10.7926 | 5785 | 1.1499
[5 empty] | 5790 | 26.1132 | 6275.12 | 10.793 | 5790 | 1.1502
[5 empty] | 5795 | 26.1066 | 6280.76 | 10.7937 | 5795 | 1.1503
[5 empty] | 5800 | 26.1013 | 6286.4 | 10.7947 | 5800 | 1.1504
[5 empty] | 5805 | 26.0984 | 6292.04 | 10.796 | 5805 | 1.1503
[5 empty] | 5810 | 26.0996 | 6297.68 | 10.7975 | 5810 | 1.1502
[5 empty] | 5815 | 26.1059 | 6303.32 | 10.799 | 5815 | 1.1501
[5 empty] | 5820 | 26.1164 | 6308.96 | 10.8005 | 5820 | 1.15
[5 empty] | 5825 | 26.1284 | 6314.6 | 10.8016 | 5825 | 1.1499
[5 empty] | 5830 | 26.1376 | 6320.24 | 10.802 | 5830 | 1.1499
[5 empty] | 5835 | 26.1416 | 6325.88 | 10.8015 | 5835 | 1.1498
[5 empty] | 5840 | 26.1405 | 6331.52 | 10.7999 | 5840 | 1.1498
[5 empty] | 5845 | 26.1358 | 6337.16 | 10.7972 | 5845 | 1.1498
[5 empty] | 5850 | 26.1292 | 6342.8 | 10.7933 | 5850 | 1.1499
[5 empty] | 5855 | 26.1224 | 6348.44 | 10.7884 | 5855 | 1.1499
[5 empty] | 5860 | 26.117 | 6354.08 | 10.7829 | 5860 | 1.15
[5 empty] | 5865 | 26.1136 | 6359.72 | 10.777 | 5865 | 1.15
[5 empty] | 5870 | 26.1126 | 6365.36 | 10.7712 | 5870 | 1.1501
[5 empty] | 5875 | 26.114 | 6371 | 10.7659 | 5875 | 1.1501
[5 empty] | 5880 | 26.1172 | 6376.64 | 10.7613 | 5880 | 1.1502
[5 empty] | 5885 | 26.1213 | 6382.28 | 10.7578 | 5885 | 1.1502
[5 empty] | 5890 | 26.1252 | 6387.92 | 10.7553 | 5890 | 1.1502
[5 empty] | 5895 | 26.128 | 6393.56 | 10.7538 | 5895 | 1.1503
[5 empty] | 5900 | 26.1293 | 6399.2 | 10.7532 | 5900 | 1.1503
[5 empty] | 5905 | 26.1296 | 6404.84 | 10.7532 | 5905 | 1.1503
[5 empty] | 5910 | 26.1292 | 6410.48 | 10.7536 | 5910 | 1.1503
[5 empty] | 5915 | 26.1283 | 6416.12 | 10.7542 | 5915 | 1.1503
[5 empty] | 5920 | 26.1273 | 6421.76 | 10.7547 | 5920 | 1.1503
[5 empty] | 5925 | 26.1265 | 6427.4 | 10.755 | 5925 | 1.1503
[5 empty] | 5930 | 26.1265 | 6433.04 | 10.755 | 5930 | 1.1503
[5 empty] | 5935 | 26.1272 | 6438.68 | 10.7549 | 5935 | 1.1503
[5 empty] | 5940 | 26.1283 | 6444.32 | 10.7545 | 5940 | 1.1504
[5 empty] | 5945 | 26.1289 | 6449.96 | 10.7541 | 5945 | 1.1504
[5 empty] | 5950 | 26.1284 | 6455.6 | 10.7537 | 5950 | 1.1505
[5 empty] | 5955 | 26.1279 | 6461.24 | 10.7534 | 5955 | 1.1506
[5 empty] | 5960 | 26.1284 | 6466.88 | 10.7535 | 5960 | 1.1507
[5 empty] | 5965 | 26.1308 | 6472.52 | 10.7538 | 5965 | 1.1507
[5 empty] | 5970 | 26.1355 | 6478.16 | 10.7545 | 5970 | 1.1508
[5 empty] | 5975 | 26.1424 | 6483.8 | 10.7556 | 5975 | 1.1508
[5 empty] | 5980 | 26.1499 | 6489.44 | 10.7571 | 5980 | 1.1508
[5 empty] | 5985 | 26.1553 | 6495.08 | 10.759 | 5985 | 1.1507
[5 empty] | 5990 | 26.1568 | 6500.72 | 10.7614 | 5990 | 1.1506
[5 empty] | 5995 | 26.1533 | 6506.36 | 10.7643 | 5995 | 1.1505
[5 empty] | 6000 | 26.1463 | 6512 | 10.768 | 6000 | 1.1503
[5 empty] | 6005 | 26.1383 | 6517.64 | 10.7726 | 6005 | 1.1501
[5 empty] | 6010 | 26.1312 | 6523.28 | 10.7783 | 6010 | 1.1498
[5 empty] | 6015 | 26.1261 | 6528.92 | 10.7852 | 6015 | 1.1494
[5 empty] | 6020 | 26.1233 | 6534.56 | 10.7934 | 6020 | 1.1491
[5 empty] | 6025 | 26.1223 | 6540.2 | 10.803 | 6025 | 1.1487
[5 empty] | 6030 | 26.1224 | 6545.84 | 10.8138 | 6030 | 1.1482
[5 empty] | 6035 | 26.1232 | 6551.48 | 10.8256 | 6035 | 1.1478
[5 empty] | 6040 | 26.1254 | 6557.12 | 10.8379 | 6040 | 1.1474
[5 empty] | 6045 | 26.1297 | 6562.76 | 10.8504 | 6045 | 1.147
[5 empty] | 6050 | 26.1368 | 6568.4 | 10.8624 | 6050 | 1.1466
[5 empty] | 6055 | 26.1451 | 6574.04 | 10.8734 | 6055 | 1.1463
[5 empty] | 6060 | 26.1525 | 6579.68 | 10.883 | 6060 | 1.1461
[5 empty] | 6065 | 26.1566 | 6585.32 | 10.8904 | 6065 | 1.1459
[5 empty] | 6070 | 26.1562 | 6590.96 | 10.8952 | 6070 | 1.1458
[5 empty] | 6075 | 26.1513 | 6596.6 | 10.897 | 6075 | 1.1458
[5 empty] | 6080 | 26.1438 | 6602.24 | 10.8956 | 6080 | 1.1459
[5 empty] | 6085 | 26.1358 | 6607.88 | 10.8909 | 6085 | 1.146
[5 empty] | 6090 | 26.129 | 6613.52 | 10.8828 | 6090 | 1.1462
[5 empty] | 6095 | 26.1244 | 6619.16 | 10.8718 | 6095 | 1.1465
[5 empty] | 6100 | 26.1225 | 6624.8 | 10.8585 | 6100 | 1.1468
[5 empty] | 6105 | 26.1236 | 6630.44 | 10.8435 | 6105 | 1.1471
[5 empty] | 6110 | 26.127 | 6636.08 | 10.8279 | 6110 | 1.1474
[5 empty] | 6115 | 26.132 | 6641.72 | 10.8127 | 6115 | 1.1478
[5 empty] | 6120 | 26.1374 | 6647.36 | 10.7989 | 6120 | 1.1481
[5 empty] | 6125 | 26.1416 | 6653 | 10.7874 | 6125 | 1.1483
[5 empty] | 6130 | 26.1444 | 6658.64 | 10.7788 | 6130 | 1.1485
[5 empty] | 6135 | 26.1459 | 6664.28 | 10.7734 | 6135 | 1.1486
[5 empty] | 6140 | 26.1461 | 6669.92 | 10.7713 | 6140 | 1.1486
[5 empty] | 6145 | 26.1456 | 6675.56 | 10.772 | 6145 | 1.1486
[5 empty] | 6150 | 26.1446 | 6681.2 | 10.7748 | 6150 | 1.1485
[5 empty] | 6155 | 26.1432 | 6686.84 | 10.7789 | 6155 | 1.1483
[5 empty] | 6160 | 26.1416 | 6692.48 | 10.7832 | 6160 | 1.1482
[5 empty] | 6165 | 26.1398 | 6698.12 | 10.7869 | 6165 | 1.148
[5 empty] | 6170 | 26.1376 | 6703.76 | 10.7892 | 6170 | 1.1478
[5 empty] | 6175 | 26.1353 | 6709.4 | 10.7894 | 6175 | 1.1476
[5 empty] | 6180 | 26.1331 | 6715.04 | 10.7874 | 6180 | 1.1475
[5 empty] | 6185 | 26.1314 | 6720.68 | 10.7831 | 6185 | 1.1474
[5 empty] | 6190 | 26.1305 | 6726.32 | 10.7767 | 6190 | 1.1474
[5 empty] | 6195 | 26.1305 | 6731.96 | 10.7689 | 6195 | 1.1474
[5 empty] | 6200 | 26.131 | 6737.6 | 10.7601 | 6200 | 1.1474
[5 empty] | 6205 | 26.1318 | 6743.24 | 10.7513 | 6205 | 1.1475
[5 empty] | 6210 | 26.132 | 6748.88 | 10.743 | 6210 | 1.1476
[5 empty] | 6215 | 26.1314 | 6754.52 | 10.736 | 6215 | 1.1477
[5 empty] | 6220 | 26.1294 | 6760.16 | 10.7308 | 6220 | 1.1477
[5 empty] | 6225 | 26.1264 | 6765.8 | 10.7278 | 6225 | 1.1478
[5 empty] | 6230 | 26.123 | 6771.44 | 10.7271 | 6230 | 1.1479
[5 empty] | 6235 | 26.1197 | 6777.08 | 10.7285 | 6235 | 1.1479
[5 empty] | 6240 | 26.1172 | 6782.72 | 10.732 | 6240 | 1.148
[5 empty] | 6245 | 26.1157 | 6788.36 | 10.7372 | 6245 | 1.148
[5 empty] | 6250 | 26.1153 | 6794 | 10.7436 | 6250 | 1.1481
[5 empty] | 6255 | 26.1159 | 6799.64 | 10.7509 | 6255 | 1.1482
[5 empty] | 6260 | 26.117 | 6805.28 | 10.7587 | 6260 | 1.1483
[5 empty] | 6265 | 26.1185 | 6810.92 | 10.7668 | 6265 | 1.1484
[5 empty] | 6270 | 26.1199 | 6816.56 | 10.775 | 6270 | 1.1486
[5 empty] | 6275 | 26.1211 | 6822.2 | 10.783 | 6275 | 1.1488
[5 empty] | 6280 | 26.1222 | 6827.84 | 10.7907 | 6280 | 1.149
[5 empty] | 6285 | 26.1231 | 6833.48 | 10.7981 | 6285 | 1.1492
[5 empty] | 6290 | 26.124 | 6839.12 | 10.8051 | 6290 | 1.1494
[5 empty] | 6295 | 26.1251 | 6844.76 | 10.8114 | 6295 | 1.1496
[5 empty] | 6300 | 26.1264 | 6850.4 | 10.817 | 6300 | 1.1498
[5 empty] | 6305 | 26.1279 | 6856.04 | 10.8218 | 6305 | 1.15
[5 empty] | 6310 | 26.1297 | 6861.68 | 10.8257 | 6310 | 1.1501
[5 empty] | 6315 | 26.1315 | 6867.32 | 10.8287 | 6315 | 1.1503
[5 empty] | 6320 | 26.1331 | 6872.96 | 10.8308 | 6320 | 1.1504
[5 empty] | 6325 | 26.1346 | 6878.6 | 10.8321 | 6325 | 1.1505
[5 empty] | 6330 | 26.1356 | 6884.24 | 10.8326 | 6330 | 1.1505
[5 empty] | 6335 | 26.1363 | 6889.88 | 10.8323 | 6335 | 1.1505
[5 empty] | 6340 | 26.1366 | 6895.52 | 10.8313 | 6340 | 1.1505
[5 empty] | 6345 | 26.1366 | 6901.16 | 10.8294 | 6345 | 1.1505
[5 empty] | 6350 | 26.1363 | 6906.8 | 10.8266 | 6350 | 1.1505
[5 empty] | 6355 | 26.136 | 6912.44 | 10.8227 | 6355 | 1.1505
[5 empty] | 6360 | 26.1361 | 6918.08 | 10.8177 | 6360 | 1.1505
[5 empty] | 6365 | 26.1369 | 6923.72 | 10.8117 | 6365 | 1.1506
[5 empty] | 6370 | 26.1386 | 6929.36 | 10.8047 | 6370 | 1.1507
[5 empty] | 6375 | 26.1411 | 6935 | 10.7973 | 6375 | 1.1508
[5 empty] | 6380 | 26.1441 | 6940.64 | 10.7897 | 6380 | 1.151
[5 empty] | 6385 | 26.1471 | 6946.28 | 10.7827 | 6385 | 1.1513
[5 empty] | 6390 | 26.1489 | 6951.92 | 10.7766 | 6390 | 1.1517
[5 empty] | 6395 | 26.1491 | 6957.56 | 10.7721 | 6395 | 1.1521
[5 empty] | 6400 | 26.1466 | 6963.2 | 10.7695 | 6400 | 1.1526
[5 empty] | 6405 | 26.1416 | 6968.84 | 10.769 | 6405 | 1.1531
[5 empty] | 6410 | 26.1352 | 6974.48 | 10.7706 | 6410 | 1.1536
[5 empty] | 6415 | 26.1285 | 6980.12 | 10.774 | 6415 | 1.1542
[5 empty] | 6420 | 26.1235 | 6985.76 | 10.779 | 6420 | 1.1547
[5 empty] | 6425 | 26.1211 | 6991.4 | 10.7851 | 6425 | 1.1552
[5 empty] | 6430 | 26.1211 | 6997.04 | 10.7916 | 6430 | 1.1557
[5 empty] | 6435 | 26.123 | 7002.68 | 10.7981 | 6435 | 1.1561
[5 empty] | 6440 | 26.1251 | 7008.32 | 10.8043 | 6440 | 1.1564
[5 empty] | 6445 | 26.1272 | 7013.96 | 10.8097 | 6445 | 1.1567
[5 empty] | 6450 | 26.1289 | 7019.6 | 10.8142 | 6450 | 1.1568
[5 empty] | 6455 | 26.1302 | 7025.24 | 10.8178 | 6455 | 1.1569
[5 empty] | 6460 | 26.1311 | 7030.88 | 10.8206 | 6460 | 1.157
[5 empty] | 6465 | 26.1317 | 7036.52 | 10.8228 | 6465 | 1.157
[5 empty] | 6470 | 26.1319 | 7042.16 | 10.8246 | 6470 | 1.1571
[5 empty] | 6475 | 26.1318 | 7047.8 | 10.8262 | 6475 | 1.1571
[5 empty] | 6480 | 26.1317 | 7053.44 | 10.8277 | 6480 | 1.1571
[5 empty] | 6485 | 26.132 | 7059.08 | 10.8293 | 6485 | 1.157
[5 empty] | 6490 | 26.1339 | 7064.72 | 10.8309 | 6490 | 1.157
[5 empty] | 6495 | 26.1383 | 7070.36 | 10.8324 | 6495 | 1.157
[5 empty] | 6500 | 26.1452 | 7076 | 10.8335 | 6500 | 1.157
[5 empty] | 6505 | 26.154 | 7081.64 | 10.8339 | 6505 | 1.1569
[5 empty] | 6510 | 26.1631 | 7087.28 | 10.8333 | 6510 | 1.1568
[5 empty] | 6515 | 26.1717 | 7092.92 | 10.8314 | 6515 | 1.1567
[5 empty] | 6520 | 26.1789 | 7098.56 | 10.8279 | 6520 | 1.1566
[5 empty] | 6525 | 26.1843 | 7104.2 | 10.8227 | 6525 | 1.1565
[5 empty] | 6530 | 26.1879 | 7109.84 | 10.8158 | 6530 | 1.1563
[5 empty] | 6535 | 26.1899 | 7115.48 | 10.8076 | 6535 | 1.1562
[5 empty] | 6540 | 26.1911 | 7121.12 | 10.7984 | 6540 | 1.1561
[5 empty] | 6545 | 26.1925 | 7126.76 | 10.789 | 6545 | 1.1559
[5 empty] | 6550 | 26.1946 | 7132.4 | 10.7797 | 6550 | 1.1558
[5 empty] | 6555 | 26.1974 | 7138.04 | 10.7715 | 6555 | 1.1557
[5 empty] | 6560 | 26.2004 | 7143.68 | 10.7646 | 6560 | 1.1557
[5 empty] | 6565 | 26.2023 | 7149.32 | 10.7596 | 6565 | 1.1556
[5 empty] | 6570 | 26.2019 | 7154.96 | 10.7565 | 6570 | 1.1556
[5 empty] | 6575 | 26.1976 | 7160.6 | 10.7554 | 6575 | 1.1556
[5 empty] | 6580 | 26.1905 | 7166.24 | 10.756 | 6580 | 1.1556
[5 empty] | 6585 | 26.1813 | 7171.88 | 10.7581 | 6585 | 1.1557
[5 empty] | 6590 | 26.1734 | 7177.52 | 10.7611 | 6590 | 1.1557
[5 empty] | 6595 | 26.1675 | 7183.16 | 10.7648 | 6595 | 1.1557
[5 empty] | 6600 | 26.1644 | 7188.8 | 10.7688 | 6600 | 1.1558
[5 empty] | 6605 | 26.164 | 7194.44 | 10.7728 | 6605 | 1.1558
[5 empty] | 6610 | 26.1654 | 7200.08 | 10.7764 | 6610 | 1.1558
[5 empty] | 6615 | 26.1679 | 7205.72 | 10.7796 | 6615 | 1.1558
[5 empty] | 6620 | 26.1706 | 7211.36 | 10.7821 | 6620 | 1.1558
[5 empty] | 6625 | 26.1723 | 7217 | 10.7837 | 6625 | 1.1558
[5 empty] | 6630 | 26.1726 | 7222.64 | 10.784 | 6630 | 1.1559
[5 empty] | 6635 | 26.1714 | 7228.28 | 10.783 | 6635 | 1.156
[5 empty] | 6640 | 26.1698 | 7233.92 | 10.7805 | 6640 | 1.1561
[5 empty] | 6645 | 26.1682 | 7239.56 | 10.7764 | 6645 | 1.1562
[5 empty] | 6650 | 26.1673 | 7245.2 | 10.771 | 6650 | 1.1563
[5 empty] | 6655 | 26.1672 | 7250.84 | 10.7645 | 6655 | 1.1565
[5 empty] | 6660 | 26.1675 | 7256.48 | 10.7575 | 6660 | 1.1567
[5 empty] | 6665 | 26.168 | 7262.12 | 10.7505 | 6665 | 1.1569
[5 empty] | 6670 | 26.1679 | 7267.76 | 10.7444 | 6670 | 1.1571
[5 empty] | 6675 | 26.167 | 7273.4 | 10.7396 | 6675 | 1.1574
[5 empty] | 6680 | 26.1647 | 7279.04 | 10.7366 | 6680 | 1.1575
[5 empty] | 6685 | 26.1614 | 7284.68 | 10.7357 | 6685 | 1.1577
[5 empty] | 6690 | 26.1569 | 7290.32 | 10.7369 | 6690 | 1.1579
[5 empty] | 6695 | 26.1519 | 7295.96 | 10.7399 | 6695 | 1.158
[5 empty] | 6700 | 26.1471 | 7301.6 | 10.7442 | 6700 | 1.1581
[5 empty] | 6705 | 26.1433 | 7307.24 | 10.7494 | 6705 | 1.1581
[5 empty] | 6710 | 26.1411 | 7312.88 | 10.7549 | 6710 | 1.1581
[5 empty] | 6715 | 26.1417 | 7318.52 | 10.7601 | 6715 | 1.1581
[5 empty] | 6720 | 26.1448 | 7324.16 | 10.7648 | 6720 | 1.1581
[5 empty] | 6725 | 26.1499 | 7329.8 | 10.7687 | 6725 | 1.1581
[5 empty] | 6730 | 26.1554 | 7335.44 | 10.7719 | 6730 | 1.1581
[5 empty] | 6735 | 26.1607 | 7341.08 | 10.7745 | 6735 | 1.1581
[5 empty] | 6740 | 26.1648 | 7346.72 | 10.7768 | 6740 | 1.1581
[5 empty] | 6745 | 26.167 | 7352.36 | 10.7791 | 6745 | 1.1581
[5 empty] | 6750 | 26.1672 | 7358 | 10.7817 | 6750 | 1.1581
[5 empty] | 6755 | 26.1655 | 7363.64 | 10.7848 | 6755 | 1.1582
[5 empty] | 6760 | 26.1625 | 7369.28 | 10.7885 | 6760 | 1.1582
[5 empty] | 6765 | 26.1585 | 7374.92 | 10.7928 | 6765 | 1.1583
[5 empty] | 6770 | 26.155 | 7380.56 | 10.7975 | 6770 | 1.1583
[5 empty] | 6775 | 26.1525 | 7386.2 | 10.8026 | 6775 | 1.1584
[5 empty] | 6780 | 26.1517 | 7391.84 | 10.8076 | 6780 | 1.1585
[5 empty] | 6785 | 26.1525 | 7397.48 | 10.8124 | 6785 | 1.1585
[5 empty] | 6790 | 26.1545 | 7403.12 | 10.8167 | 6790 | 1.1586
[5 empty] | 6795 | 26.1568 | 7408.76 | 10.8204 | 6795 | 1.1587
[5 empty] | 6800 | 26.1586 | 7414.4 | 10.8233 | 6800 | 1.1587
[5 empty] | 6805 | 26.1593 | 7420.04 | 10.8254 | 6805 | 1.1588
[5 empty] | 6810 | 26.1588 | 7425.68 | 10.8267 | 6810 | 1.1588
[5 empty] | 6815 | 26.1577 | 7431.32 | 10.8273 | 6815 | 1.1588
[5 empty] | 6820 | 26.1565 | 7436.96 | 10.8275 | 6820 | 1.1588
[5 empty] | 6825 | 26.1556 | 7442.6 | 10.8273 | 6825 | 1.1588
[5 empty] | 6830 | 26.1553 | 7448.24 | 10.8271 | 6830 | 1.1588
[5 empty] | 6835 | 26.1558 | 7453.88 | 10.8268 | 6835 | 1.1588
[5 empty] | 6840 | 26.1571 | 7459.52 | 10.8267 | 6840 | 1.1588
[5 empty] | 6845 | 26.1585 | 7465.16 | 10.8269 | 6845 | 1.1588
[5 empty] | 6850 | 26.1598 | 7470.8 | 10.8273 | 6850 | 1.1588
[5 empty] | 6855 | 26.16 | 7476.44 | 10.8279 | 6855 | 1.1588
[5 empty] | 6860 | 26.1588 | 7482.08 | 10.8287 | 6860 | 1.1588
[5 empty] | 6865 | 26.1563 | 7487.72 | 10.8296 | 6865 | 1.1588
[5 empty] | 6870 | 26.1534 | 7493.36 | 10.8304 | 6870 | 1.1588
[5 empty] | 6875 | 26.1507 | 7499 | 10.8312 | 6875 | 1.1588
[5 empty] | 6880 | 26.149 | 7504.64 | 10.8318 | 6880 | 1.1588
[5 empty] | 6885 | 26.1487 | 7510.28 | 10.8321 | 6885 | 1.1588
[5 empty] | 6890 | 26.1495 | 7515.92 | 10.8323 | 6890 | 1.1587
[5 empty] | 6895 | 26.1516 | 7521.56 | 10.8322 | 6895 | 1.1587
[5 empty] | 6900 | 26.1537 | 7527.2 | 10.832 | 6900 | 1.1586
[5 empty] | 6905 | 26.1557 | 7532.84 | 10.8317 | 6905 | 1.1586
[5 empty] | 6910 | 26.1571 | 7538.48 | 10.8313 | 6910 | 1.1586
[5 empty] | 6915 | 26.1578 | 7544.12 | 10.831 | 6915 | 1.1585
[5 empty] | 6920 | 26.1579 | 7549.76 | 10.8308 | 6920 | 1.1585
[5 empty] | 6925 | 26.1575 | 7555.4 | 10.8308 | 6925 | 1.1585
[5 empty] | 6930 | 26.157 | 7561.04 | 10.8309 | 6930 | 1.1585
[5 empty] | 6935 | 26.1566 | 7566.68 | 10.8312 | 6935 | 1.1585
[5 empty] | 6940 | 26.1563 | 7572.32 | 10.8318 | 6940 | 1.1586
[5 empty] | 6945 | 26.1565 | 7577.96 | 10.8325 | 6945 | 1.1586
[5 empty] | 6950 | 26.1571 | 7583.6 | 10.8333 | 6950 | 1.1587
[5 empty] | 6955 | 26.1582 | 7589.24 | 10.8342 | 6955 | 1.1589
[5 empty] | 6960 | 26.1598 | 7594.88 | 10.8351 | 6960 | 1.159
[5 empty] | 6965 | 26.1615 | 7600.52 | 10.836 | 6965 | 1.1591
[5 empty] | 6970 | 26.1635 | 7606.16 | 10.8368 | 6970 | 1.1593
[5 empty] | 6975 | 26.1655 | 7611.8 | 10.8374 | 6975 | 1.1595
[5 empty] | 6980 | 26.1672 | 7617.44 | 10.8379 | 6980 | 1.1596
[5 empty] | 6985 | 26.1686 | 7623.08 | 10.8381 | 6985 | 1.1598
[5 empty] | 6990 | 26.1696 | 7628.72 | 10.8382 | 6990 | 1.1599
[5 empty] | 6995 | 26.1703 | 7634.36 | 10.838 | 6995 | 1.16
[5 empty] | 7000 | 26.1705 | 7640 | 10.8377 | 7000 | 1.1602
[5 empty] | 7005 | 26.1703 | 7645.64 | 10.8374 | 7005 | 1.1603
[5 empty] | 7010 | 26.1698 | 7651.28 | 10.8369 | 7010 | 1.1603
[5 empty] | 7015 | 26.1688 | 7656.92 | 10.8365 | 7015 | 1.1604
[5 empty] | 7020 | 26.1675 | 7662.56 | 10.8361 | 7020 | 1.1605
[5 empty] | 7025 | 26.1658 | 7668.2 | 10.8357 | 7025 | 1.1605
[5 empty] | 7030 | 26.1639 | 7673.84 | 10.8355 | 7030 | 1.1605
[5 empty] | 7035 | 26.1617 | 7679.48 | 10.8353 | 7035 | 1.1605
[5 empty] | 7040 | 26.1596 | 7685.12 | 10.8352 | 7040 | 1.1605
[5 empty] | 7045 | 26.1579 | 7690.76 | 10.8351 | 7045 | 1.1605
[5 empty] | 7050 | 26.1566 | 7696.4 | 10.8351 | 7050 | 1.1604
[5 empty] | 7055 | 26.156 | 7702.04 | 10.8351 | 7055 | 1.1603
[5 empty] | 7060 | 26.1559 | 7707.68 | 10.8352 | 7060 | 1.1601
[5 empty] | 7065 | 26.1564 | 7713.32 | 10.8354 | 7065 | 1.16
[5 empty] | 7070 | 26.1572 | 7718.96 | 10.8356 | 7070 | 1.1598
[5 empty] | 7075 | 26.1579 | 7724.6 | 10.836 | 7075 | 1.1595
[5 empty] | 7080 | 26.1583 | 7730.24 | 10.8364 | 7080 | 1.1593
[5 empty] | 7085 | 26.158 | 7735.88 | 10.837 | 7085 | 1.159
[5 empty] | 7090 | 26.1571 | 7741.52 | 10.8376 | 7090 | 1.1587
[5 empty] | 7095 | 26.1555 | 7747.16 | 10.8384 | 7095 | 1.1584
[5 empty] | 7100 | 26.1535 | 7752.8 | 10.8392 | 7100 | 1.1581
[5 empty] | 7105 | 26.1514 | 7758.44 | 10.8401 | 7105 | 1.1578
[5 empty] | 7110 | 26.1493 | 7764.08 | 10.8409 | 7110 | 1.1576
[5 empty] | 7115 | 26.1476 | 7769.72 | 10.8418 | 7115 | 1.1573
[5 empty] | 7120 | 26.1464 | 7775.36 | 10.8426 | 7120 | 1.1571
[5 empty] | 7125 | 26.1457 | 7781 | 10.8433 | 7125 | 1.1569
[5 empty] | 7130 | 26.1455 | 7786.64 | 10.8441 | 7130 | 1.1568
[5 empty] | 7135 | 26.1456 | 7792.28 | 10.8448 | 7135 | 1.1567
[5 empty] | 7140 | 26.1461 | 7797.92 | 10.8455 | 7140 | 1.1567
[5 empty] | 7145 | 26.1467 | 7803.56 | 10.8461 | 7145 | 1.1568
[5 empty] | 7150 | 26.1477 | 7809.2 | 10.8468 | 7150 | 1.1569
[5 empty] | 7155 | 26.149 | 7814.84 | 10.8475 | 7155 | 1.1571
[5 empty] | 7160 | 26.1506 | 7820.48 | 10.8482 | 7160 | 1.1574
[5 empty] | 7165 | 26.1526 | 7826.12 | 10.8489 | 7165 | 1.1577
[5 empty] | 7170 | 26.155 | 7831.76 | 10.8496 | 7170 | 1.158
[5 empty] | 7175 | 26.1575 | 7837.4 | 10.8502 | 7175 | 1.1584
[5 empty] | 7180 | 26.1599 | 7843.04 | 10.8507 | 7180 | 1.1588
[5 empty] | 7185 | 26.162 | 7848.68 | 10.8511 | 7185 | 1.1591
[5 empty] | 7190 | 26.1638 | 7854.32 | 10.8515 | 7190 | 1.1595
[5 empty] | 7195 | 26.1654 | 7859.96 | 10.8517 | 7195 | 1.1598
[5 empty] | 7200 | 26.1666 | 7865.6 | 10.852 | 7200 | 1.1601
[5 empty] | 7205 | 26.1677 | 7871.24 | 10.8522 | 7205 | 1.1603
[5 empty] | 7210 | 26.1688 | 7876.88 | 10.8524 | 7210 | 1.1605
[5 empty] | 7215 | 26.1699 | 7882.52 | 10.8526 | 7215 | 1.1606
[5 empty] | 7220 | 26.1712 | 7888.16 | 10.8529 | 7220 | 1.1607
[5 empty] | 7225 | 26.1727 | 7893.8 | 10.8532 | 7225 | 1.1607
[5 empty] | 7230 | 26.1742 | 7899.44 | 10.8535 | 7230 | 1.1606
[5 empty] | 7235 | 26.1753 | 7905.08 | 10.8539 | 7235 | 1.1606
[5 empty] | 7240 | 26.1761 | 7910.72 | 10.8543 | 7240 | 1.1604
[5 empty] | 7245 | 26.1764 | 7916.36 | 10.8546 | 7245 | 1.1603
[5 empty] | 7250 | 26.1762 | 7922 | 10.855 | 7250 | 1.1601
[5 empty] | 7255 | 26.1756 | 7927.64 | 10.8552 | 7255 | 1.1599
[5 empty] | 7260 | 26.1748 | 7933.28 | 10.8555 | 7260 | 1.1596
[5 empty] | 7265 | 26.174 | 7938.92 | 10.8556 | 7265 | 1.1594
[5 empty] | 7270 | 26.174 | 7944.56 | 10.8558 | 7270 | 1.1591
[5 empty] | 7275 | 26.1745 | 7950.2 | 10.8559 | 7275 | 1.1589
[5 empty] | 7280 | 26.1758 | 7955.84 | 10.8561 | 7280 | 1.1586
[5 empty] | 7285 | 26.1777 | 7961.48 | 10.8563 | 7285 | 1.1584
[5 empty] | 7290 | 26.18 | 7967.12 | 10.8566 | 7290 | 1.1581
[5 empty] | 7295 | 26.1826 | 7972.76 | 10.8569 | 7295 | 1.1579
[5 empty] | 7300 | 26.1849 | 7978.4 | 10.8573 | 7300 | 1.1576
[5 empty] | 7305 | 26.1868 | 7984.04 | 10.8578 | 7305 | 1.1573
[5 empty] | 7310 | 26.1881 | 7989.68 | 10.8583 | 7310 | 1.1571
[5 empty] | 7315 | 26.1887 | 7995.32 | 10.8588 | 7315 | 1.1568
[5 empty] | 7320 | 26.1888 | 8000.96 | 10.8594 | 7320 | 1.1566
[5 empty] | 7325 | 26.1885 | 8006.6 | 10.86 | 7325 | 1.1564
[5 empty] | 7330 | 26.188 | 8012.24 | 10.8606 | 7330 | 1.1561
[5 empty] | 7335 | 26.1874 | 8017.88 | 10.8611 | 7335 | 1.1559
[5 empty] | 7340 | 26.187 | 8023.52 | 10.8617 | 7340 | 1.1557
[5 empty] | 7345 | 26.1868 | 8029.16 | 10.8622 | 7345 | 1.1555
[5 empty] | 7350 | 26.1869 | 8034.8 | 10.8627 | 7350 | 1.1553
[5 empty] | 7355 | 26.1872 | 8040.44 | 10.8632 | 7355 | 1.1552
[5 empty] | 7360 | 26.1876 | 8046.08 | 10.8636 | 7360 | 1.1551
[5 empty] | 7365 | 26.188 | 8051.72 | 10.864 | 7365 | 1.155
[5 empty] | 7370 | 26.1885 | 8057.36 | 10.8644 | 7370 | 1.155
[5 empty] | 7375 | 26.1889 | 8063 | 10.8647 | 7375 | 1.155
[5 empty] | 7380 | 26.1892 | 8068.64 | 10.865 | 7380 | 1.155
[5 empty] | 7385 | 26.1897 | 8074.28 | 10.8652 | 7385 | 1.1551
[5 empty] | 7390 | 26.1901 | 8079.92 | 10.8655 | 7390 | 1.1552
[5 empty] | 7395 | 26.1906 | 8085.56 | 10.8657 | 7395 | 1.1554
[5 empty] | 7400 | 26.1913 | 8091.2 | 10.8659 | 7400 | 1.1555
[5 empty] | 7405 | 26.192 | 8096.84 | 10.8662 | 7405 | 1.1558
[5 empty] | 7410 | 26.1927 | 8102.48 | 10.8665 | 7410 | 1.156
[5 empty] | 7415 | 26.1934 | 8108.12 | 10.8668 | 7415 | 1.1563
[5 empty] | 7420 | 26.1939 | 8113.76 | 10.8673 | 7420 | 1.1566
[5 empty] | 7425 | 26.1943 | 8119.4 | 10.8678 | 7425 | 1.157
[5 empty] | 7430 | 26.1945 | 8125.04 | 10.8684 | 7430 | 1.1573
[5 empty] | 7435 | 26.1945 | 8130.68 | 10.8691 | 7435 | 1.1577
[5 empty] | 7440 | 26.1943 | 8136.32 | 10.8699 | 7440 | 1.1581
[5 empty] | 7445 | 26.194 | 8141.96 | 10.8707 | 7445 | 1.1585
[5 empty] | 7450 | 26.1938 | 8147.6 | 10.8715 | 7450 | 1.1589
[5 empty] | 7455 | 26.1936 | 8153.24 | 10.8722 | 7455 | 1.1593
[5 empty] | 7460 | 26.1935 | 8158.88 | 10.8729 | 7460 | 1.1597
[5 empty] | 7465 | 26.1936 | 8164.52 | 10.8734 | 7465 | 1.16
[5 empty] | 7470 | 26.1938 | 8170.16 | 10.8739 | 7470 | 1.1604
[5 empty] | 7475 | 26.1941 | 8175.8 | 10.8743 | 7475 | 1.1607
[5 empty] | 7480 | 26.1944 | 8181.44 | 10.8746 | 7480 | 1.1609
[5 empty] | 7485 | 26.1949 | 8187.08 | 10.8748 | 7485 | 1.1611
[5 empty] | 7490 | 26.1955 | 8192.72 | 10.8749 | 7490 | 1.1613
[5 empty] | 7495 | 26.1962 | 8198.36 | 10.8751 | 7495 | 1.1614
[5 empty] | 7500 | 26.1971 | 8204 | 10.8753 | 7500 | 1.1615
[5 empty] | 7505 | 26.1981 | 8209.64 | 10.8756 | 7505 | 1.1615
[5 empty] | 7510 | 26.1991 | 8215.28 | 10.876 | 7510 | 1.1614
[5 empty] | 7515 | 26.2001 | 8220.92 | 10.8766 | 7515 | 1.1614
[5 empty] | 7520 | 26.2011 | 8226.56 | 10.8774 | 7520 | 1.1613
[5 empty] | 7525 | 26.2019 | 8232.2 | 10.8783 | 7525 | 1.1611
[5 empty] | 7530 | 26.2024 | 8237.84 | 10.8795 | 7530 | 1.1609
[5 empty] | 7535 | 26.2026 | 8243.48 | 10.8808 | 7535 | 1.1607
[5 empty] | 7540 | 26.2024 | 8249.12 | 10.8822 | 7540 | 1.1605
[5 empty] | 7545 | 26.2018 | 8254.76 | 10.8837 | 7545 | 1.1603
[5 empty] | 7550 | 26.2011 | 8260.4 | 10.8853 | 7550 | 1.1601
[5 empty] | 7555 | 26.2003 | 8266.04 | 10.887 | 7555 | 1.1598
[5 empty] | 7560 | 26.1994 | 8271.68 | 10.8886 | 7560 | 1.1596
[5 empty] | 7565 | 26.1987 | 8277.32 | 10.8901 | 7565 | 1.1594
[5 empty] | 7570 | 26.1983 | 8282.96 | 10.8916 | 7570 | 1.1592
[5 empty] | 7575 | 26.1983 | 8288.6 | 10.8928 | 7575 | 1.159
[5 empty] | 7580 | 26.1987 | 8294.24 | 10.8939 | 7580 | 1.1589
[5 empty] | 7585 | 26.1995 | 8299.88 | 10.8947 | 7585 | 1.1587
[5 empty] | 7590 | 26.2006 | 8305.52 | 10.8953 | 7590 | 1.1586
[5 empty] | 7595 | 26.2018 | 8311.16 | 10.8956 | 7595 | 1.1586
[5 empty] | 7600 | 26.2031 | 8316.8 | 10.8956 | 7600 | 1.1585
[5 empty] | 7605 | 26.2042 | 8322.44 | 10.8955 | 7605 | 1.1585
[5 empty] | 7610 | 26.2053 | 8328.08 | 10.8953 | 7610 | 1.1585
[5 empty] | 7615 | 26.2061 | 8333.72 | 10.895 | 7615 | 1.1585
[5 empty] | 7620 | 26.2067 | 8339.36 | 10.8946 | 7620 | 1.1585
[5 empty] | 7625 | 26.207 | 8345 | 10.8944 | 7625 | 1.1586
[5 empty] | 7630 | 26.207 | 8350.64 | 10.8944 | 7630 | 1.1587
[5 empty] | 7635 | 26.2068 | 8356.28 | 10.8945 | 7635 | 1.1588
[5 empty] | 7640 | 26.2064 | 8361.92 | 10.895 | 7640 | 1.1588
[5 empty] | 7645 | 26.2057 | 8367.56 | 10.8958 | 7645 | 1.1589
[5 empty] | 7650 | 26.2048 | 8373.2 | 10.8969 | 7650 | 1.159
[5 empty] | 7655 | 26.2037 | 8378.84 | 10.8982 | 7655 | 1.1591
[5 empty] | 7660 | 26.2024 | 8384.48 | 10.8999 | 7660 | 1.1592
[5 empty] | 7665 | 26.2009 | 8390.12 | 10.9017 | 7665 | 1.1593
[5 empty] | 7670 | 26.1996 | 8395.76 | 10.9036 | 7670 | 1.1594
[5 empty] | 7675 | 26.1982 | 8401.4 | 10.9055 | 7675 | 1.1595
[5 empty] | 7680 | 26.1972 | 8407.04 | 10.9073 | 7680 | 1.1595
[5 empty] | 7685 | 26.1966 | 8412.68 | 10.909 | 7685 | 1.1596
[5 empty] | 7690 | 26.1963 | 8418.32 | 10.9105 | 7690 | 1.1596
[5 empty] | 7695 | 26.1966 | 8423.96 | 10.9118 | 7695 | 1.1596
[5 empty] | 7700 | 26.1974 | 8429.6 | 10.9128 | 7700 | 1.1596
[5 empty] | 7705 | 26.1986 | 8435.24 | 10.9136 | 7705 | 1.1595
[5 empty] | 7710 | 26.2002 | 8440.88 | 10.9142 | 7710 | 1.1595
[5 empty] | 7715 | 26.2021 | 8446.52 | 10.9146 | 7715 | 1.1594
[5 empty] | 7720 | 26.204 | 8452.16 | 10.9149 | 7720 | 1.1593
[5 empty] | 7725 | 26.2058 | 8457.8 | 10.9151 | 7725 | 1.1591
[5 empty] | 7730 | 26.2073 | 8463.44 | 10.9153 | 7730 | 1.159
[5 empty] | 7735 | 26.2086 | 8469.08 | 10.9156 | 7735 | 1.1588
[5 empty] | 7740 | 26.2094 | 8474.72 | 10.916 | 7740 | 1.1586
[5 empty] | 7745 | 26.2098 | 8480.36 | 10.9165 | 7745 | 1.1585
[5 empty] | 7750 | 26.21 | 8486 | 10.9172 | 7750 | 1.1583
[5 empty] | 7755 | 26.2098 | 8491.64 | 10.9181 | 7755 | 1.1581
[5 empty] | 7760 | 26.2093 | 8497.28 | 10.9191 | 7760 | 1.158
[5 empty] | 7765 | 26.2088 | 8502.92 | 10.9202 | 7765 | 1.1578
[5 empty] | 7770 | 26.2082 | 8508.56 | 10.9214 | 7770 | 1.1577
[5 empty] | 7775 | 26.2077 | 8514.2 | 10.9226 | 7775 | 1.1576
[5 empty] | 7780 | 26.2074 | 8519.84 | 10.9238 | 7780 | 1.1576
[5 empty] | 7785 | 26.2073 | 8525.48 | 10.9249 | 7785 | 1.1575
[5 empty] | 7790 | 26.2074 | 8531.12 | 10.926 | 7790 | 1.1575
[5 empty] | 7795 | 26.2077 | 8536.76 | 10.927 | 7795 | 1.1575
[5 empty] | 7800 | 26.2081 | 8542.4 | 10.9278 | 7800 | 1.1576
[5 empty] | 7805 | 26.2087 | 8548.04 | 10.9284 | 7805 | 1.1577
[5 empty] | 7810 | 26.2095 | 8553.68 | 10.929 | 7810 | 1.1578
[5 empty] | 7815 | 26.2104 | 8559.32 | 10.9295 | 7815 | 1.158
[5 empty] | 7820 | 26.2113 | 8564.96 | 10.9299 | 7820 | 1.1581
[5 empty] | 7825 | 26.2124 | 8570.6 | 10.9302 | 7825 | 1.1583
[5 empty] | 7830 | 26.2135 | 8576.24 | 10.9306 | 7830 | 1.1585
[5 empty] | 7835 | 26.2146 | 8581.88 | 10.931 | 7835 | 1.1587
[5 empty] | 7840 | 26.2158 | 8587.52 | 10.9314 | 7840 | 1.1589
[5 empty] | 7845 | 26.2171 | 8593.16 | 10.9319 | 7845 | 1.1591
[5 empty] | 7850 | 26.2184 | 8598.8 | 10.9325 | 7850 | 1.1594
[5 empty] | 7855 | 26.2196 | 8604.44 | 10.9331 | 7855 | 1.1596
[5 empty] | 7860 | 26.2209 | 8610.08 | 10.9338 | 7860 | 1.1598
[5 empty] | 7865 | 26.222 | 8615.72 | 10.9345 | 7865 | 1.16
[5 empty] | 7870 | 26.223 | 8621.36 | 10.9351 | 7870 | 1.1602
[5 empty] | 7875 | 26.2239 | 8627 | 10.9357 | 7875 | 1.1604
[5 empty] | 7880 | 26.2247 | 8632.64 | 10.9361 | 7880 | 1.1606
[5 empty] | 7885 | 26.2253 | 8638.28 | 10.9364 | 7885 | 1.1608
[5 empty] | 7890 | 26.2257 | 8643.92 | 10.9366 | 7890 | 1.161
[5 empty] | 7895 | 26.226 | 8649.56 | 10.9365 | 7895 | 1.1612
[5 empty] | 7900 | 26.2262 | 8655.2 | 10.9363 | 7900 | 1.1615
[5 empty] | 7905 | 26.2263 | 8660.84 | 10.9359 | 7905 | 1.1617
[5 empty] | 7910 | 26.2264 | 8666.48 | 10.9354 | 7910 | 1.1619
[5 empty] | 7915 | 26.2264 | 8672.12 | 10.9349 | 7915 | 1.1622
[5 empty] | 7920 | 26.2264 | 8677.76 | 10.9343 | 7920 | 1.1625
[5 empty] | 7925 | 26.2265 | 8683.4 | 10.9338 | 7925 | 1.1628
[5 empty] | 7930 | 26.2266 | 8689.04 | 10.9334 | 7930 | 1.1632
[5 empty] | 7935 | 26.2268 | 8694.68 | 10.9332 | 7935 | 1.1635
[5 empty] | 7940 | 26.2271 | 8700.32 | 10.9331 | 7940 | 1.1639
[5 empty] | 7945 | 26.2274 | 8705.96 | 10.9333 | 7945 | 1.1643
[5 empty] | 7950 | 26.2279 | 8711.6 | 10.9337 | 7950 | 1.1648
[5 empty] | 7955 | 26.2284 | 8717.24 | 10.9343 | 7955 | 1.1652
[5 empty] | 7960 | 26.2291 | 8722.88 | 10.935 | 7960 | 1.1656
[5 empty] | 7965 | 26.2299 | 8728.52 | 10.9359 | 7965 | 1.1661
[5 empty] | 7970 | 26.2308 | 8734.16 | 10.9369 | 7970 | 1.1666
[5 empty] | 7975 | 26.2317 | 8739.8 | 10.9379 | 7975 | 1.167
[5 empty] | 7980 | 26.2328 | 8745.44 | 10.9389 | 7980 | 1.1674
[5 empty] | 7985 | 26.2338 | 8751.08 | 10.9398 | 7985 | 1.1678
[5 empty] | 7990 | 26.2349 | 8756.72 | 10.9406 | 7990 | 1.1682
[5 empty] | 7995 | 26.236 | 8762.36 | 10.9414 | 7995 | 1.1686
[5 empty] | 8000 | 26.237 | 8768 | 10.9419 | 8000 | 1.1689
[5 empty] | 8005 | 26.2379 | 8773.64 | 10.9424 | 8005 | 1.1692
[5 empty] | 8010 | 26.2387 | 8779.28 | 10.9426 | 8010 | 1.1694
[5 empty] | 8015 | 26.2394 | 8784.92 | 10.9428 | 8015 | 1.1696
[5 empty] | 8020 | 26.24 | 8790.56 | 10.9427 | 8020 | 1.1698
[5 empty] | 8025 | 26.2404 | 8796.2 | 10.9426 | 8025 | 1.1699
[5 empty] | 8030 | 26.2408 | 8801.84 | 10.9423 | 8030 | 1.1699
[5 empty] | 8035 | 26.2411 | 8807.48 | 10.9419 | 8035 | 1.1699
[5 empty] | 8040 | 26.2414 | 8813.12 | 10.9414 | 8040 | 1.1699
[5 empty] | 8045 | 26.2417 | 8818.76 | 10.9409 | 8045 | 1.1699
[5 empty] | 8050 | 26.2422 | 8824.4 | 10.9403 | 8050 | 1.1698
[5 empty] | 8055 | 26.2428 | 8830.04 | 10.9397 | 8055 | 1.1697
[5 empty] | 8060 | 26.2435 | 8835.68 | 10.9391 | 8060 | 1.1695
[5 empty] | 8065 | 26.2444 | 8841.32 | 10.9386 | 8065 | 1.1694
[5 empty] | 8070 | 26.2453 | 8846.96 | 10.9382 | 8070 | 1.1692
[5 empty] | 8075 | 26.2464 | 8852.6 | 10.9379 | 8075 | 1.169
[5 empty] | 8080 | 26.2474 | 8858.24 | 10.9377 | 8080 | 1.1688
[5 empty] | 8085 | 26.2485 | 8863.88 | 10.9378 | 8085 | 1.1687
[5 empty] | 8090 | 26.2494 | 8869.52 | 10.9379 | 8090 | 1.1685
[5 empty] | 8095 | 26.2503 | 8875.16 | 10.9383 | 8095 | 1.1683
[5 empty] | 8100 | 26.251 | 8880.8 | 10.9388 | 8100 | 1.1682
[5 empty] | 8105 | 26.2515 | 8886.44 | 10.9394 | 8105 | 1.168
[5 empty] | 8110 | 26.252 | 8892.08 | 10.9402 | 8110 | 1.1679
[5 empty] | 8115 | 26.2523 | 8897.72 | 10.9409 | 8115 | 1.1678
[5 empty] | 8120 | 26.2528 | 8903.36 | 10.9417 | 8120 | 1.1677
[5 empty] | 8125 | 26.2532 | 8909 | 10.9425 | 8125 | 1.1676
[5 empty] | 8130 | 26.2538 | 8914.64 | 10.9432 | 8130 | 1.1675
[5 empty] | 8135 | 26.2546 | 8920.28 | 10.9438 | 8135 | 1.1675
[5 empty] | 8140 | 26.2555 | 8925.92 | 10.9442 | 8140 | 1.1674
[5 empty] | 8145 | 26.2566 | 8931.56 | 10.9445 | 8145 | 1.1674
[5 empty] | 8150 | 26.2579 | 8937.2 | 10.9447 | 8150 | 1.1674
[5 empty] | 8155 | 26.2592 | 8942.84 | 10.9447 | 8155 | 1.1673
[5 empty] | 8160 | 26.2608 | 8948.48 | 10.9446 | 8160 | 1.1673
[5 empty] | 8165 | 26.2624 | 8954.12 | 10.9445 | 8165 | 1.1673
[5 empty] | 8170 | 26.264 | 8959.76 | 10.9443 | 8170 | 1.1672
[5 empty] | 8175 | 26.2655 | 8965.4 | 10.9442 | 8175 | 1.1672
[5 empty] | 8180 | 26.2669 | 8971.04 | 10.9441 | 8180 | 1.1672
[5 empty] | 8185 | 26.2683 | 8976.68 | 10.9441 | 8185 | 1.1671
[5 empty] | 8190 | 26.2696 | 8982.32 | 10.9442 | 8190 | 1.1671
[5 empty] | 8195 | 26.2708 | 8987.96 | 10.9444 | 8195 | 1.1671
[5 empty] | 8200 | 26.2718 | 8993.6 | 10.9447 | 8200 | 1.1671
[5 empty] | 8205 | 26.2727 | 8999.24 | 10.9451 | 8205 | 1.1671
[5 empty] | 8210 | 26.2735 | 9004.88 | 10.9456 | 8210 | 1.1671
[5 empty] | 8215 | 26.2742 | 9010.52 | 10.9462 | 8215 | 1.1672
[5 empty] | 8220 | 26.2749 | 9016.16 | 10.9468 | 8220 | 1.1672
[5 empty] | 8225 | 26.2756 | 9021.8 | 10.9474 | 8225 | 1.1673
[5 empty] | 8230 | 26.2763 | 9027.44 | 10.948 | 8230 | 1.1674
[5 empty] | 8235 | 26.2769 | 9033.08 | 10.9486 | 8235 | 1.1675
[5 empty] | 8240 | 26.2777 | 9038.72 | 10.9492 | 8240 | 1.1676
[5 empty] | 8245 | 26.2785 | 9044.36 | 10.9497 | 8245 | 1.1677
[5 empty] | 8250 | 26.2792 | 9050 | 10.9501 | 8250 | 1.1679
[5 empty] | 8255 | 26.28 | 9055.64 | 10.9505 | 8255 | 1.1681
[5 empty] | 8260 | 26.2808 | 9061.28 | 10.9508 | 8260 | 1.1683
[5 empty] | 8265 | 26.2817 | 9066.92 | 10.9511 | 8265 | 1.1685
[5 empty] | 8270 | 26.2825 | 9072.56 | 10.9513 | 8270 | 1.1688
[5 empty] | 8275 | 26.2833 | 9078.2 | 10.9514 | 8275 | 1.1691
[5 empty] | 8280 | 26.2842 | 9083.84 | 10.9515 | 8280 | 1.1693
[5 empty] | 8285 | 26.285 | 9089.48 | 10.9515 | 8285 | 1.1696
[5 empty] | 8290 | 26.2859 | 9095.12 | 10.9515 | 8290 | 1.1699
[5 empty] | 8295 | 26.2868 | 9100.76 | 10.9515 | 8295 | 1.1702
[5 empty] | 8300 | 26.2877 | 9106.4 | 10.9514 | 8300 | 1.1705
[5 empty] | 8305 | 26.2886 | 9112.04 | 10.9513 | 8305 | 1.1708
[5 empty] | 8310 | 26.2895 | 9117.68 | 10.9511 | 8310 | 1.1711
[5 empty] | 8315 | 26.2903 | 9123.32 | 10.951 | 8315 | 1.1714
[5 empty] | 8320 | 26.2912 | 9128.96 | 10.9508 | 8320 | 1.1717
[5 empty] | 8325 | 26.292 | 9134.6 | 10.9506 | 8325 | 1.172
[5 empty] | 8330 | 26.2926 | 9140.24 | 10.9504 | 8330 | 1.1723
[5 empty] | 8335 | 26.2932 | 9145.88 | 10.9502 | 8335 | 1.1726
[5 empty] | 8340 | 26.2937 | 9151.52 | 10.95 | 8340 | 1.1728
[5 empty] | 8345 | 26.2941 | 9157.16 | 10.9498 | 8345 | 1.1731
[5 empty] | 8350 | 26.2944 | 9162.8 | 10.9496 | 8350 | 1.1733
[5 empty] | 8355 | 26.2946 | 9168.44 | 10.9494 | 8355 | 1.1736
[5 empty] | 8360 | 26.2947 | 9174.08 | 10.9492 | 8360 | 1.1738
[5 empty] | 8365 | 26.2949 | 9179.72 | 10.9491 | 8365 | 1.174
[5 empty] | 8370 | 26.2951 | 9185.36 | 10.949 | 8370 | 1.1742
[5 empty] | 8375 | 26.2953 | 9191 | 10.9489 | 8375 | 1.1744
[5 empty] | 8380 | 26.2957 | 9196.64 | 10.9488 | 8380 | 1.1746
[5 empty] | 8385 | 26.2962 | 9202.28 | 10.9488 | 8385 | 1.1748
[5 empty] | 8390 | 26.2968 | 9207.92 | 10.9488 | 8390 | 1.175
[5 empty] | 8395 | 26.2976 | 9213.56 | 10.9488 | 8395 | 1.1752
[5 empty] | 8400 | 26.2985 | 9219.2 | 10.9488 | 8400 | 1.1754
[5 empty] | 8405 | 26.2994 | 9224.84 | 10.9487 | 8405 | 1.1756
[5 empty] | 8410 | 26.3004 | 9230.48 | 10.9487 | 8410 | 1.1757
[5 empty] | 8415 | 26.3014 | 9236.12 | 10.9487 | 8415 | 1.1759
[5 empty] | 8420 | 26.3022 | 9241.76 | 10.9486 | 8420 | 1.176
[5 empty] | 8425 | 26.3031 | 9247.4 | 10.9485 | 8425 | 1.1761
[5 empty] | 8430 | 26.3038 | 9253.04 | 10.9484 | 8430 | 1.1762
[5 empty] | 8435 | 26.3044 | 9258.68 | 10.9482 | 8435 | 1.1763
[5 empty] | 8440 | 26.3049 | 9264.32 | 10.9481 | 8440 | 1.1764
[5 empty] | 8445 | 26.3054 | 9269.96 | 10.9479 | 8445 | 1.1764
[5 empty] | 8450 | 26.3058 | 9275.6 | 10.9477 | 8450 | 1.1765
[5 empty] | 8455 | 26.3062 | 9281.24 | 10.9476 | 8455 | 1.1765
[5 empty] | 8460 | 26.3066 | 9286.88 | 10.9474 | 8460 | 1.1764
[5 empty] | 8465 | 26.3071 | 9292.52 | 10.9473 | 8465 | 1.1764
[5 empty] | 8470 | 26.3075 | 9298.16 | 10.9472 | 8470 | 1.1764
[5 empty] | 8475 | 26.3081 | 9303.8 | 10.9472 | 8475 | 1.1763
[5 empty] | 8480 | 26.3086 | 9309.44 | 10.9472 | 8480 | 1.1762
[5 empty] | 8485 | 26.3091 | 9315.08 | 10.9472 | 8485 | 1.1761
[5 empty] | 8490 | 26.3094 | 9320.72 | 10.9471 | 8490 | 1.176
[5 empty] | 8495 | 26.3097 | 9326.36 | 10.9471 | 8495 | 1.1759
[5 empty] | 8500 | 26.3099 | 9332 | 10.947 | 8500 | 1.1757
[5 empty] | 8505 | 26.3099 | 9337.64 | 10.9469 | 8505 | 1.1756
[5 empty] | 8510 | 26.3097 | 9343.28 | 10.9467 | 8510 | 1.1754
[5 empty] | 8515 | 26.3093 | 9348.92 | 10.9464 | 8515 | 1.1753
[5 empty] | 8520 | 26.3088 | 9354.56 | 10.9461 | 8520 | 1.1751
[5 empty] | 8525 | 26.3082 | 9360.2 | 10.9457 | 8525 | 1.175
[5 empty] | 8530 | 26.3076 | 9365.84 | 10.9452 | 8530 | 1.1749
[5 empty] | 8535 | 26.3069 | 9371.48 | 10.9447 | 8535 | 1.1747
[5 empty] | 8540 | 26.3064 | 9377.12 | 10.9441 | 8540 | 1.1746
[5 empty] | 8545 | 26.306 | 9382.76 | 10.9436 | 8545 | 1.1745
[5 empty] | 8550 | 26.3059 | 9388.4 | 10.943 | 8550 | 1.1744
[5 empty] | 8555 | 26.3059 | 9394.04 | 10.9424 | 8555 | 1.1743
[5 empty] | 8560 | 26.3063 | 9399.68 | 10.9419 | 8560 | 1.1743
[5 empty] | 8565 | 26.307 | 9405.32 | 10.9413 | 8565 | 1.1742
[5 empty] | 8570 | 26.308 | 9410.96 | 10.9408 | 8570 | 1.1742
[5 empty] | 8575 | 26.3091 | 9416.6 | 10.9402 | 8575 | 1.1743
[5 empty] | 8580 | 26.3106 | 9422.24 | 10.9396 | 8580 | 1.1743
[5 empty] | 8585 | 26.3122 | 9427.88 | 10.939 | 8585 | 1.1744
[5 empty] | 8590 | 26.3141 | 9433.52 | 10.9384 | 8590 | 1.1745
[5 empty] | 8595 | 26.3159 | 9439.16 | 10.9377 | 8595 | 1.1747
[5 empty] | 8600 | 26.3178 | 9444.8 | 10.937 | 8600 | 1.1749
[5 empty] | 8605 | 26.3197 | 9450.44 | 10.9363 | 8605 | 1.1751
[5 empty] | 8610 | 26.3215 | 9456.08 | 10.9356 | 8610 | 1.1753
[5 empty] | 8615 | 26.3232 | 9461.72 | 10.9348 | 8615 | 1.1756
[5 empty] | 8620 | 26.3249 | 9467.36 | 10.9341 | 8620 | 1.1759
[5 empty] | 8625 | 26.3263 | 9473 | 10.9335 | 8625 | 1.1762
[5 empty] | 8630 | 26.3276 | 9478.64 | 10.9328 | 8630 | 1.1766
[5 empty] | 8635 | 26.3287 | 9484.28 | 10.9323 | 8635 | 1.1769
[5 empty] | 8640 | 26.3297 | 9489.92 | 10.9319 | 8640 | 1.1773
[5 empty] | 8645 | 26.3306 | 9495.56 | 10.9315 | 8645 | 1.1777
[5 empty] | 8650 | 26.3314 | 9501.2 | 10.9313 | 8650 | 1.1781
[5 empty] | 8655 | 26.332 | 9506.84 | 10.9312 | 8655 | 1.1785
[5 empty] | 8660 | 26.3326 | 9512.48 | 10.9312 | 8660 | 1.1789
[5 empty] | 8665 | 26.333 | 9518.12 | 10.9313 | 8665 | 1.1793
[5 empty] | 8670 | 26.3333 | 9523.76 | 10.9315 | 8670 | 1.1796
[5 empty] | 8675 | 26.3335 | 9529.4 | 10.9318 | 8675 | 1.18
[5 empty] | 8680 | 26.3337 | 9535.04 | 10.932 | 8680 | 1.1803
[5 empty] | 8685 | 26.3336 | 9540.68 | 10.9323 | 8685 | 1.1806
[5 empty] | 8690 | 26.3335 | 9546.32 | 10.9326 | 8690 | 1.1809
[5 empty] | 8695 | 26.3333 | 9551.96 | 10.9328 | 8695 | 1.1812
[5 empty] | 8700 | 26.333 | 9557.6 | 10.933 | 8700 | 1.1814
[5 empty] | 8705 | 26.3326 | 9563.24 | 10.9331 | 8705 | 1.1816
[5 empty] | 8710 | 26.332 | 9568.88 | 10.933 | 8710 | 1.1817
[5 empty] | 8715 | 26.3315 | 9574.52 | 10.9329 | 8715 | 1.1818
[5 empty] | 8720 | 26.3309 | 9580.16 | 10.9326 | 8720 | 1.1819
[5 empty] | 8725 | 26.3303 | 9585.8 | 10.9323 | 8725 | 1.1819
[5 empty] | 8730 | 26.3296 | 9591.44 | 10.9317 | 8730 | 1.1819
[5 empty] | 8735 | 26.329 | 9597.08 | 10.9311 | 8735 | 1.1819
[5 empty] | 8740 | 26.3285 | 9602.72 | 10.9304 | 8740 | 1.1818
[5 empty] | 8745 | 26.3281 | 9608.36 | 10.9296 | 8745 | 1.1817
[5 empty] | 8750 | 26.3278 | 9614 | 10.9287 | 8750 | 1.1816
[5 empty] | 8755 | 26.3277 | 9619.64 | 10.9277 | 8755 | 1.1815
[5 empty] | 8760 | 26.3279 | 9625.28 | 10.9267 | 8760 | 1.1813
[5 empty] | 8765 | 26.3283 | 9630.92 | 10.9257 | 8765 | 1.1812
[5 empty] | 8770 | 26.3288 | 9636.56 | 10.9248 | 8770 | 1.181
[5 empty] | 8775 | 26.3297 | 9642.2 | 10.9238 | 8775 | 1.1808
[5 empty] | 8780 | 26.3307 | 9647.84 | 10.923 | 8780 | 1.1806
[5 empty] | 8785 | 26.3318 | 9653.48 | 10.9222 | 8785 | 1.1805
[5 empty] | 8790 | 26.333 | 9659.12 | 10.9215 | 8790 | 1.1803
[5 empty] | 8795 | 26.3343 | 9664.76 | 10.921 | 8795 | 1.1801
[5 empty] | 8800 | 26.3356 | 9670.4 | 10.9205 | 8800 | 1.18
[5 empty] | 8805 | 26.3365 | 9676.04 | 10.9202 | 8805 | 1.1798
[5 empty] | 8810 | 26.3374 | 9681.68 | 10.92 | 8810 | 1.1797
[5 empty] | 8815 | 26.338 | 9687.32 | 10.9199 | 8815 | 1.1795
[5 empty] | 8820 | 26.3383 | 9692.96 | 10.9199 | 8820 | 1.1794
[5 empty] | 8825 | 26.3382 | 9698.6 | 10.92 | 8825 | 1.1793
[5 empty] | 8830 | 26.3379 | 9704.24 | 10.9202 | 8830 | 1.1792
[5 empty] | 8835 | 26.3373 | 9709.88 | 10.9205 | 8835 | 1.1791
[5 empty] | 8840 | 26.3364 | 9715.52 | 10.9208 | 8840 | 1.179
[5 empty] | 8845 | 26.3353 | 9721.16 | 10.9212 | 8845 | 1.179
[5 empty] | 8850 | 26.3341 | 9726.8 | 10.9216 | 8850 | 1.1789
[5 empty] | 8855 | 26.3328 | 9732.44 | 10.9221 | 8855 | 1.1788
[5 empty] | 8860 | 26.3314 | 9738.08 | 10.9226 | 8860 | 1.1788
[5 empty] | 8865 | 26.3301 | 9743.72 | 10.9231 | 8865 | 1.1788
[5 empty] | 8870 | 26.329 | 9749.36 | 10.9236 | 8870 | 1.1787
[5 empty] | 8875 | 26.3279 | 9755 | 10.9241 | 8875 | 1.1787
[5 empty] | 8880 | 26.3269 | 9760.64 | 10.9245 | 8880 | 1.1787
[5 empty] | 8885 | 26.3261 | 9766.28 | 10.9249 | 8885 | 1.1787
[5 empty] | 8890 | 26.3256 | 9771.92 | 10.9253 | 8890 | 1.1787
[5 empty] | 8895 | 26.3253 | 9777.56 | 10.9256 | 8895 | 1.1787
[5 empty] | 8900 | 26.3252 | 9783.2 | 10.9259 | 8900 | 1.1788
[5 empty] | 8905 | 26.3253 | 9788.84 | 10.9261 | 8905 | 1.1788
[5 empty] | 8910 | 26.3256 | 9794.48 | 10.9263 | 8910 | 1.1789
[5 empty] | 8915 | 26.3261 | 9800.12 | 10.9263 | 8915 | 1.1789
[5 empty] | 8920 | 26.3267 | 9805.76 | 10.9264 | 8920 | 1.179
[5 empty] | 8925 | 26.3275 | 9811.4 | 10.9264 | 8925 | 1.1791
[5 empty] | 8930 | 26.3284 | 9817.04 | 10.9263 | 8930 | 1.1793
[5 empty] | 8935 | 26.3295 | 9822.68 | 10.9262 | 8935 | 1.1794
[5 empty] | 8940 | 26.3306 | 9828.32 | 10.926 | 8940 | 1.1796
[5 empty] | 8945 | 26.3315 | 9833.96 | 10.9259 | 8945 | 1.1797
[5 empty] | 8950 | 26.3325 | 9839.6 | 10.9257 | 8950 | 1.1799
[5 empty] | 8955 | 26.3333 | 9845.24 | 10.9255 | 8955 | 1.1801
[5 empty] | 8960 | 26.3339 | 9850.88 | 10.9253 | 8960 | 1.1803
[5 empty] | 8965 | 26.3343 | 9856.52 | 10.9251 | 8965 | 1.1805
[5 empty] | 8970 | 26.3345 | 9862.16 | 10.9249 | 8970 | 1.1808
[5 empty] | 8975 | 26.3344 | 9867.8 | 10.9246 | 8975 | 1.181
[5 empty] | 8980 | 26.3341 | 9873.44 | 10.9244 | 8980 | 1.1812
[5 empty] | 8985 | 26.3335 | 9879.08 | 10.9242 | 8985 | 1.1815
[5 empty] | 8990 | 26.3329 | 9884.72 | 10.924 | 8990 | 1.1817
[5 empty] | 8995 | 26.3322 | 9890.36 | 10.9237 | 8995 | 1.182
[5 empty] | 9000 | 26.3315 | 9896 | 10.9234 | 9000 | 1.1822
[5 empty] | 9005 | 26.3308 | 9901.64 | 10.9231 | 9005 | 1.1825
[5 empty] | 9010 | 26.3303 | 9907.28 | 10.9228 | 9010 | 1.1827
[5 empty] | 9015 | 26.3301 | 9912.92 | 10.9225 | 9015 | 1.183
[5 empty] | 9020 | 26.3301 | 9918.56 | 10.9221 | 9020 | 1.1832
[5 empty] | 9025 | 26.3304 | 9924.2 | 10.9216 | 9025 | 1.1835
[5 empty] | 9030 | 26.331 | 9929.84 | 10.9212 | 9030 | 1.1837
[5 empty] | 9035 | 26.332 | 9935.48 | 10.9207 | 9035 | 1.1839
[5 empty] | 9040 | 26.3329 | 9941.12 | 10.9203 | 9040 | 1.1841
[5 empty] | 9045 | 26.3341 | 9946.76 | 10.9198 | 9045 | 1.1843
[5 empty] | 9050 | 26.3353 | 9952.4 | 10.9193 | 9050 | 1.1845
[5 empty] | 9055 | 26.3365 | 9958.04 | 10.9188 | 9055 | 1.1847
[5 empty] | 9060 | 26.3376 | 9963.68 | 10.9184 | 9060 | 1.1848
[5 empty] | 9065 | 26.3385 | 9969.32 | 10.918 | 9065 | 1.185
[5 empty] | 9070 | 26.3393 | 9974.96 | 10.9176 | 9070 | 1.1851
[5 empty] | 9075 | 26.3397 | 9980.6 | 10.9173 | 9075 | 1.1853
[5 empty] | 9080 | 26.3397 | 9986.24 | 10.917 | 9080 | 1.1854
[5 empty] | 9085 | 26.3395 | 9991.88 | 10.9168 | 9085 | 1.1855
[5 empty] | 9090 | 26.3391 | 9997.52 | 10.9167 | 9090 | 1.1855
[5 empty] | 9095 | 26.3385 | 10003.16 | 10.9166 | 9095 | 1.1856
[5 empty] | 9100 | 26.3376 | 10008.8 | 10.9165 | 9100 | 1.1857
[5 empty] | 9105 | 26.3366 | 10014.44 | 10.9165 | 9105 | 1.1857
[5 empty] | 9110 | 26.3357 | 10020.08 | 10.9165 | 9110 | 1.1857
[5 empty] | 9115 | 26.3347 | 10025.72 | 10.9165 | 9115 | 1.1857
[5 empty] | 9120 | 26.3336 | 10031.36 | 10.9165 | 9120 | 1.1857
[5 empty] | 9125 | 26.3328 | 10037 | 10.9166 | 9125 | 1.1856
[5 empty] | 9130 | 26.332 | 10042.64 | 10.9165 | 9130 | 1.1855
[5 empty] | 9135 | 26.3314 | 10048.28 | 10.9165 | 9135 | 1.1855
[5 empty] | 9140 | 26.3308 | 10053.92 | 10.9164 | 9140 | 1.1853
[5 empty] | 9145 | 26.3304 | 10059.56 | 10.9163 | 9145 | 1.1852
[5 empty] | 9150 | 26.33 | 10065.2 | 10.9161 | 9150 | 1.1851
[5 empty] | 9155 | 26.3298 | 10070.84 | 10.9158 | 9155 | 1.1849
[5 empty] | 9160 | 26.3295 | 10076.48 | 10.9156 | 9160 | 1.1848
[5 empty] | 9165 | 26.3293 | 10082.12 | 10.9152 | 9165 | 1.1846
[5 empty] | 9170 | 26.3291 | 10087.76 | 10.9149 | 9170 | 1.1844
[5 empty] | 9175 | 26.329 | 10093.4 | 10.9145 | 9175 | 1.1842
[5 empty] | 9180 | 26.3288 | 10099.04 | 10.9142 | 9180 | 1.184
[5 empty] | 9185 | 26.3287 | 10104.68 | 10.9138 | 9185 | 1.1838
[5 empty] | 9190 | 26.3286 | 10110.32 | 10.9135 | 9190 | 1.1836
[5 empty] | 9195 | 26.3286 | 10115.96 | 10.9132 | 9195 | 1.1834
[5 empty] | 9200 | 26.3286 | 10121.6 | 10.9129 | 9200 | 1.1832
[5 empty] | 9205 | 26.3289 | 10127.24 | 10.9127 | 9205 | 1.183
[5 empty] | 9210 | 26.3292 | 10132.88 | 10.9125 | 9210 | 1.1828
[5 empty] | 9215 | 26.3297 | 10138.52 | 10.9124 | 9215 | 1.1826
[5 empty] | 9220 | 26.3303 | 10144.16 | 10.9123 | 9220 | 1.1825
[5 empty] | 9225 | 26.3309 | 10149.8 | 10.9122 | 9225 | 1.1823
[5 empty] | 9230 | 26.3316 | 10155.44 | 10.9122 | 9230 | 1.1822
[5 empty] | 9235 | 26.3324 | 10161.08 | 10.9122 | 9235 | 1.182
[5 empty] | 9240 | 26.3332 | 10166.72 | 10.9122 | 9240 | 1.1819
[5 empty] | 9245 | 26.3338 | 10172.36 | 10.9123 | 9245 | 1.1818
[5 empty] | 9250 | 26.3344 | 10178 | 10.9123 | 9250 | 1.1817
[5 empty] | 9255 | 26.3349 | 10183.64 | 10.9124 | 9255 | 1.1816
[5 empty] | 9260 | 26.3351 | 10189.28 | 10.9125 | 9260 | 1.1816
[5 empty] | 9265 | 26.3351 | 10194.92 | 10.9126 | 9265 | 1.1815
[5 empty] | 9270 | 26.335 | 10200.56 | 10.9127 | 9270 | 1.1815
[5 empty] | 9275 | 26.3348 | 10206.2 | 10.9129 | 9275 | 1.1815
[5 empty] | 9280 | 26.3344 | 10211.84 | 10.913 | 9280 | 1.1815
[5 empty] | 9285 | 26.3338 | 10217.48 | 10.9132 | 9285 | 1.1815
[5 empty] | 9290 | 26.3332 | 10223.12 | 10.9133 | 9290 | 1.1815
[5 empty] | 9295 | 26.3327 | 10228.76 | 10.9135 | 9295 | 1.1815
[5 empty] | 9300 | 26.3321 | 10234.4 | 10.9138 | 9300 | 1.1815
[5 empty] | 9305 | 26.3317 | 10240.04 | 10.914 | 9305 | 1.1815
[5 empty] | 9310 | 26.3313 | 10245.68 | 10.9143 | 9310 | 1.1816
[5 empty] | 9315 | 26.3311 | 10251.32 | 10.9145 | 9315 | 1.1816
[5 empty] | 9320 | 26.3311 | 10256.96 | 10.9148 | 9320 | 1.1816
[5 empty] | 9325 | 26.3312 | 10262.6 | 10.9152 | 9325 | 1.1816
[5 empty] | 9330 | 26.3314 | 10268.24 | 10.9155 | 9330 | 1.1817
[5 empty] | 9335 | 26.3318 | 10273.88 | 10.9159 | 9335 | 1.1817
[5 empty] | 9340 | 26.3322 | 10279.52 | 10.9162 | 9340 | 1.1817
[5 empty] | 9345 | 26.3327 | 10285.16 | 10.9166 | 9345 | 1.1817
[5 empty] | 9350 | 26.3332 | 10290.8 | 10.9169 | 9350 | 1.1818
[5 empty] | 9355 | 26.3337 | 10296.44 | 10.9173 | 9355 | 1.1818
[5 empty] | 9360 | 26.3341 | 10302.08 | 10.9176 | 9360 | 1.1818
[5 empty] | 9365 | 26.3344 | 10307.72 | 10.9179 | 9365 | 1.1818
[5 empty] | 9370 | 26.3347 | 10313.36 | 10.9182 | 9370 | 1.1819
[5 empty] | 9375 | 26.3349 | 10319 | 10.9184 | 9375 | 1.1819
[5 empty] | 9380 | 26.335 | 10324.64 | 10.9186 | 9380 | 1.1819
[5 empty] | 9385 | 26.3349 | 10330.28 | 10.9186 | 9385 | 1.1819
[5 empty] | 9390 | 26.3347 | 10335.92 | 10.9186 | 9390 | 1.1819
[5 empty] | 9395 | 26.3344 | 10341.56 | 10.9185 | 9395 | 1.1819
[5 empty] | 9400 | 26.3342 | 10347.2 | 10.9183 | 9400 | 1.1819
[5 empty] | 9405 | 26.3338 | 10352.84 | 10.918 | 9405 | 1.1819
[5 empty] | 9410 | 26.3333 | 10358.48 | 10.9176 | 9410 | 1.1818
[5 empty] | 9415 | 26.3328 | 10364.12 | 10.9172 | 9415 | 1.1818
[5 empty] | 9420 | 26.3323 | 10369.76 | 10.9166 | 9420 | 1.1817
[5 empty] | 9425 | 26.3317 | 10375.4 | 10.916 | 9425 | 1.1817
[5 empty] | 9430 | 26.3312 | 10381.04 | 10.9153 | 9430 | 1.1816
[5 empty] | 9435 | 26.3305 | 10386.68 | 10.9145 | 9435 | 1.1815
[5 empty] | 9440 | 26.33 | 10392.32 | 10.9138 | 9440 | 1.1814
[5 empty] | 9445 | 26.3295 | 10397.96 | 10.913 | 9445 | 1.1813
[5 empty] | 9450 | 26.3289 | 10403.6 | 10.9122 | 9450 | 1.1811
[5 empty] | 9455 | 26.3284 | 10409.24 | 10.9115 | 9455 | 1.181
[5 empty] | 9460 | 26.328 | 10414.88 | 10.9107 | 9460 | 1.1808
[5 empty] | 9465 | 26.3276 | 10420.52 | 10.91 | 9465 | 1.1806
[5 empty] | 9470 | 26.3273 | 10426.16 | 10.9094 | 9470 | 1.1804
[5 empty] | 9475 | 26.3271 | 10431.8 | 10.9088 | 9475 | 1.1802
[5 empty] | 9480 | 26.3269 | 10437.44 | 10.9082 | 9480 | 1.1799
[5 empty] | 9485 | 26.3269 | 10443.08 | 10.9077 | 9485 | 1.1797
[5 empty] | 9490 | 26.327 | 10448.72 | 10.9073 | 9490 | 1.1794
[5 empty] | 9495 | 26.3272 | 10454.36 | 10.9069 | 9495 | 1.1791
[5 empty] | 9500 | 26.3275 | 10460 | 10.9065 | 9500 | 1.1788
[5 empty] | 9505 | 26.3278 | 10465.64 | 10.9062 | 9505 | 1.1785
[5 empty] | 9510 | 26.3282 | 10471.28 | 10.9059 | 9510 | 1.1781
[5 empty] | 9515 | 26.3287 | 10476.92 | 10.9056 | 9515 | 1.1778
[5 empty] | 9520 | 26.3294 | 10482.56 | 10.9054 | 9520 | 1.1775
[5 empty] | 9525 | 26.33 | 10488.2 | 10.9052 | 9525 | 1.1771
[5 empty] | 9530 | 26.3306 | 10493.84 | 10.905 | 9530 | 1.1768
[5 empty] | 9535 | 26.3313 | 10499.48 | 10.9049 | 9535 | 1.1764
[5 empty] | 9540 | 26.3319 | 10505.12 | 10.9048 | 9540 | 1.1761
[5 empty] | 9545 | 26.3326 | 10510.76 | 10.9047 | 9545 | 1.1757
[5 empty] | 9550 | 26.3331 | 10516.4 | 10.9047 | 9550 | 1.1754
[5 empty] | 9555 | 26.3337 | 10522.04 | 10.9047 | 9555 | 1.1751
[5 empty] | 9560 | 26.3342 | 10527.68 | 10.9047 | 9560 | 1.1748
[5 empty] | 9565 | 26.3346 | 10533.32 | 10.9047 | 9565 | 1.1745
[5 empty] | 9570 | 26.3349 | 10538.96 | 10.9048 | 9570 | 1.1742
[5 empty] | 9575 | 26.3352 | 10544.6 | 10.9049 | 9575 | 1.1739
[5 empty] | 9580 | 26.3352 | 10550.24 | 10.9051 | 9580 | 1.1737
[5 empty] | 9585 | 26.3351 | 10555.88 | 10.9053 | 9585 | 1.1734
[5 empty] | 9590 | 26.335 | 10561.52 | 10.9055 | 9590 | 1.1732
[5 empty] | 9595 | 26.3348 | 10567.16 | 10.9058 | 9595 | 1.173
[5 empty] | 9600 | 26.3344 | 10572.8 | 10.9061 | 9600 | 1.1729
[5 empty] | 9605 | 26.334 | 10578.44 | 10.9064 | 9605 | 1.1727
[5 empty] | 9610 | 26.3334 | 10584.08 | 10.9068 | 9610 | 1.1726
[5 empty] | 9615 | 26.3329 | 10589.72 | 10.9072 | 9615 | 1.1725
[5 empty] | 9620 | 26.3323 | 10595.36 | 10.9077 | 9620 | 1.1724
[5 empty] | 9625 | 26.3317 | 10601 | 10.9082 | 9625 | 1.1723
[5 empty] | 9630 | 26.3313 | 10606.64 | 10.9087 | 9630 | 1.1722
[5 empty] | 9635 | 26.3309 | 10612.28 | 10.9093 | 9635 | 1.1722
[5 empty] | 9640 | 26.3306 | 10617.92 | 10.9099 | 9640 | 1.1722
[5 empty] | 9645 | 26.3305 | 10623.56 | 10.9106 | 9645 | 1.1721
[5 empty] | 9650 | 26.3304 | 10629.2 | 10.9112 | 9650 | 1.1721
[5 empty] | 9655 | 26.3304 | 10634.84 | 10.9119 | 9655 | 1.1721
[5 empty] | 9660 | 26.3307 | 10640.48 | 10.9125 | 9660 | 1.1721
[5 empty] | 9665 | 26.3311 | 10646.12 | 10.9132 | 9665 | 1.1721
[5 empty] | 9670 | 26.3314 | 10651.76 | 10.9138 | 9670 | 1.1721
[5 empty] | 9675 | 26.3318 | 10657.4 | 10.9143 | 9675 | 1.1721
[5 empty] | 9680 | 26.3322 | 10663.04 | 10.9147 | 9680 | 1.1721
[5 empty] | 9685 | 26.3325 | 10668.68 | 10.9151 | 9685 | 1.1721
[5 empty] | 9690 | 26.3327 | 10674.32 | 10.9153 | 9690 | 1.1721
[5 empty] | 9695 | 26.3327 | 10679.96 | 10.9154 | 9695 | 1.1721
[5 empty] | 9700 | 26.3326 | 10685.6 | 10.9154 | 9700 | 1.1721
[5 empty] | 9705 | 26.3323 | 10691.24 | 10.9152 | 9705 | 1.1721
[5 empty] | 9710 | 26.3319 | 10696.88 | 10.9149 | 9710 | 1.1721
[5 empty] | 9715 | 26.3311 | 10702.52 | 10.9145 | 9715 | 1.172
[5 empty] | 9720 | 26.3302 | 10708.16 | 10.9139 | 9720 | 1.172
[5 empty] | 9725 | 26.3292 | 10713.8 | 10.9132 | 9725 | 1.172
[5 empty] | 9730 | 26.3282 | 10719.44 | 10.9124 | 9730 | 1.1719
[5 empty] | 9735 | 26.327 | 10725.08 | 10.9115 | 9735 | 1.1718
[5 empty] | 9740 | 26.3258 | 10730.72 | 10.9105 | 9740 | 1.1718
[5 empty] | 9745 | 26.3247 | 10736.36 | 10.9094 | 9745 | 1.1717
[5 empty] | 9750 | 26.3238 | 10742 | 10.9084 | 9750 | 1.1717
[5 empty] | 9755 | 26.3229 | 10747.64 | 10.9073 | 9755 | 1.1716
[5 empty] | 9760 | 26.3221 | 10753.28 | 10.9063 | 9760 | 1.1715
[5 empty] | 9765 | 26.3215 | 10758.92 | 10.9053 | 9765 | 1.1715
[5 empty] | 9770 | 26.3211 | 10764.56 | 10.9043 | 9770 | 1.1714
[5 empty] | 9775 | 26.3208 | 10770.2 | 10.9035 | 9775 | 1.1714
[5 empty] | 9780 | 26.3207 | 10775.84 | 10.9028 | 9780 | 1.1714
[5 empty] | 9785 | 26.3206 | 10781.48 | 10.9021 | 9785 | 1.1713
[5 empty] | 9790 | 26.3205 | 10787.12 | 10.9016 | 9790 | 1.1713
[5 empty] | 9795 | 26.3205 | 10792.76 | 10.9012 | 9795 | 1.1713
[5 empty] | 9800 | 26.3205 | 10798.4 | 10.901 | 9800 | 1.1714
[5 empty] | 9805 | 26.3203 | 10804.04 | 10.9008 | 9805 | 1.1714
[5 empty] | 9810 | 26.3201 | 10809.68 | 10.9007 | 9810 | 1.1715
[5 empty] | 9815 | 26.3198 | 10815.32 | 10.9008 | 9815 | 1.1716
[5 empty] | 9820 | 26.3194 | 10820.96 | 10.9009 | 9820 | 1.1717
[5 empty] | 9825 | 26.3189 | 10826.6 | 10.9011 | 9825 | 1.1718
[5 empty] | 9830 | 26.3182 | 10832.24 | 10.9013 | 9830 | 1.1719
[5 empty] | 9835 | 26.3175 | 10837.88 | 10.9016 | 9835 | 1.1721
[5 empty] | 9840 | 26.3168 | 10843.52 | 10.9019 | 9840 | 1.1722
[5 empty] | 9845 | 26.3162 | 10849.16 | 10.9022 | 9845 | 1.1724
[5 empty] | 9850 | 26.3155 | 10854.8 | 10.9025 | 9850 | 1.1726
[5 empty] | 9855 | 26.315 | 10860.44 | 10.9028 | 9855 | 1.1728
[5 empty] | 9860 | 26.3145 | 10866.08 | 10.903 | 9860 | 1.173
[5 empty] | 9865 | 26.3142 | 10871.72 | 10.9032 | 9865 | 1.1732
[5 empty] | 9870 | 26.3141 | 10877.36 | 10.9034 | 9870 | 1.1734
[5 empty] | 9875 | 26.314 | 10883 | 10.9035 | 9875 | 1.1736
[5 empty] | 9880 | 26.314 | 10888.64 | 10.9035 | 9880 | 1.1739
[5 empty] | 9885 | 26.3142 | 10894.28 | 10.9035 | 9885 | 1.1741
[5 empty] | 9890 | 26.3145 | 10899.92 | 10.9034 | 9890 | 1.1743
[5 empty] | 9895 | 26.3148 | 10905.56 | 10.9033 | 9895 | 1.1745
[5 empty] | 9900 | 26.315 | 10911.2 | 10.9031 | 9900 | 1.1746
[5 empty] | 9905 | 26.3152 | 10916.84 | 10.9029 | 9905 | 1.1748
[5 empty] | 9910 | 26.3154 | 10922.48 | 10.9026 | 9910 | 1.175
[5 empty] | 9915 | 26.3155 | 10928.12 | 10.9023 | 9915 | 1.1751
[5 empty] | 9920 | 26.3155 | 10933.76 | 10.902 | 9920 | 1.1752
[5 empty] | 9925 | 26.3154 | 10939.4 | 10.9017 | 9925 | 1.1753
[5 empty] | 9930 | 26.3151 | 10945.04 | 10.9013 | 9930 | 1.1754
[5 empty] | 9935 | 26.3147 | 10950.68 | 10.9009 | 9935 | 1.1755
[5 empty] | 9940 | 26.3144 | 10956.32 | 10.9005 | 9940 | 1.1755
[5 empty] | 9945 | 26.3139 | 10961.96 | 10.9001 | 9945 | 1.1756
[5 empty] | 9950 | 26.3134 | 10967.6 | 10.8997 | 9950 | 1.1756
[5 empty] | 9955 | 26.313 | 10973.24 | 10.8993 | 9955 | 1.1756
[5 empty] | 9960 | 26.3126 | 10978.88 | 10.8989 | 9960 | 1.1756
[5 empty] | 9965 | 26.3122 | 10984.52 | 10.8984 | 9965 | 1.1755
[5 empty] | 9970 | 26.312 | 10990.16 | 10.898 | 9970 | 1.1755
[5 empty] | 9975 | 26.3118 | 10995.8 | 10.8975 | 9975 | 1.1754
[5 empty] | 9980 | 26.3118 | 11001.44 | 10.8971 | 9980 | 1.1753
[5 empty] | 9985 | 26.312 | 11007.08 | 10.8966 | 9985 | 1.1752
[5 empty] | 9990 | 26.3122 | 11012.72 | 10.8962 | 9990 | 1.1751
[5 empty] | 9995 | 26.3125 | 11018.36 | 10.8957 | 9995 | 1.175
[5 empty] | 10000 | 26.3129 | 11024 | 10.8952 | 10000 | 1.1749
[5 empty] | 10005 | 26.3133 | 11029.64 | 10.8947 | 10005 | 1.1748
[5 empty] | 10010 | 26.3137 | 11035.28 | 10.8941 | 10010 | 1.1747
[5 empty] | 10015 | 26.3142 | 11040.92 | 10.8936 | 10015 | 1.1746
[5 empty] | 10020 | 26.3146 | 11046.56 | 10.8931 | 10020 | 1.1745
[5 empty] | 10025 | 26.3149 | 11052.2 | 10.8925 | 10025 | 1.1744
[5 empty] | 10030 | 26.3151 | 11057.84 | 10.892 | 10030 | 1.1743
[5 empty] | 10035 | 26.3153 | 11063.48 | 10.8915 | 10035 | 1.1743
[5 empty] | 10040 | 26.3155 | 11069.12 | 10.8909 | 10040 | 1.1742
[5 empty] | 10045 | 26.3155 | 11074.76 | 10.8904 | 10045 | 1.1742
[5 empty] | 10050 | 26.3155 | 11080.4 | 10.8899 | 10050 | 1.1741
[5 empty] | 10055 | 26.3153 | 11086.04 | 10.8894 | 10055 | 1.1741
[5 empty] | 10060 | 26.3152 | 11091.68 | 10.889 | 10060 | 1.1741
[5 empty] | 10065 | 26.3151 | 11097.32 | 10.8885 | 10065 | 1.1741
[5 empty] | 10070 | 26.3149 | 11102.96 | 10.8881 | 10070 | 1.1741
[5 empty] | 10075 | 26.3147 | 11108.6 | 10.8877 | 10075 | 1.1741
[5 empty] | 10080 | 26.3146 | 11114.24 | 10.8874 | 10080 | 1.1741
[5 empty] | 10085 | 26.3146 | 11119.88 | 10.8871 | 10085 | 1.1741
[5 empty] | 10090 | 26.3146 | 11125.52 | 10.8868 | 10090 | 1.1742
[5 empty] | 10095 | 26.3146 | 11131.16 | 10.8865 | 10095 | 1.1742
[5 empty] | 10100 | 26.3148 | 11136.8 | 10.8863 | 10100 | 1.1742
[5 empty] | 10105 | 26.315 | 11142.44 | 10.8861 | 10105 | 1.1743
[5 empty] | 10110 | 26.3152 | 11148.08 | 10.8859 | 10110 | 1.1743
[5 empty] | 10115 | 26.3155 | 11153.72 | 10.8857 | 10115 | 1.1744
[5 empty] | 10120 | 26.3158 | 11159.36 | 10.8856 | 10120 | 1.1744
[5 empty] | 10125 | 26.316 | 11165 | 10.8855 | 10125 | 1.1745
[5 empty] | 10130 | 26.3163 | 11170.64 | 10.8853 | 10130 | 1.1746
[5 empty] | 10135 | 26.3166 | 11176.28 | 10.8852 | 10135 | 1.1746
[5 empty] | 10140 | 26.3167 | 11181.92 | 10.8851 | 10140 | 1.1746
[5 empty] | 10145 | 26.3167 | 11187.56 | 10.885 | 10145 | 1.1747
[5 empty] | 10150 | 26.3167 | 11193.2 | 10.8849 | 10150 | 1.1747
[5 empty] | 10155 | 26.3165 | 11198.84 | 10.8847 | 10155 | 1.1748
[5 empty] | 10160 | 26.3161 | 11204.48 | 10.8846 | 10160 | 1.1748
[5 empty] | 10165 | 26.3156 | 11210.12 | 10.8845 | 10165 | 1.1748
[5 empty] | 10170 | 26.3148 | 11215.76 | 10.8844 | 10170 | 1.1749
[5 empty] | 10175 | 26.314 | 11221.4 | 10.8843 | 10175 | 1.1749
[5 empty] | 10180 | 26.313 | 11227.04 | 10.8842 | 10180 | 1.1749
[5 empty] | 10185 | 26.3119 | 11232.68 | 10.884 | 10185 | 1.1749
[5 empty] | 10190 | 26.3107 | 11238.32 | 10.8839 | 10190 | 1.1749
[5 empty] | 10195 | 26.3093 | 11243.96 | 10.8838 | 10195 | 1.1749
[5 empty] | 10200 | 26.308 | 11249.6 | 10.8837 | 10200 | 1.1749
[5 empty] | 10205 | 26.3066 | 11255.24 | 10.8836 | 10205 | 1.1748
[5 empty] | 10210 | 26.3052 | 11260.88 | 10.8834 | 10210 | 1.1748
[5 empty] | 10215 | 26.3039 | 11266.52 | 10.8833 | 10215 | 1.1748
[5 empty] | 10220 | 26.3026 | 11272.16 | 10.8832 | 10220 | 1.1747
[5 empty] | 10225 | 26.3015 | 11277.8 | 10.883 | 10225 | 1.1747
[5 empty] | 10230 | 26.3004 | 11283.44 | 10.8829 | 10230 | 1.1746
[5 empty] | 10235 | 26.2995 | 11289.08 | 10.8827 | 10235 | 1.1746
[5 empty] | 10240 | 26.2988 | 11294.72 | 10.8825 | 10240 | 1.1745
[5 empty] | 10245 | 26.2982 | 11300.36 | 10.8823 | 10245 | 1.1744
[5 empty] | 10250 | 26.2976 | 11306 | 10.8821 | 10250 | 1.1743
[5 empty] | 10255 | 26.2973 | 11311.64 | 10.8818 | 10255 | 1.1742
[5 empty] | 10260 | 26.2971 | 11317.28 | 10.8815 | 10260 | 1.1741
[5 empty] | 10265 | 26.2969 | 11322.92 | 10.8812 | 10265 | 1.174
[5 empty] | 10270 | 26.2968 | 11328.56 | 10.8808 | 10270 | 1.1738
[5 empty] | 10275 | 26.2967 | 11334.2 | 10.8804 | 10275 | 1.1737
[5 empty] | 10280 | 26.2967 | 11339.84 | 10.8799 | 10280 | 1.1735
[5 empty] | 10285 | 26.2967 | 11345.48 | 10.8795 | 10285 | 1.1734
[5 empty] | 10290 | 26.2966 | 11351.12 | 10.879 | 10290 | 1.1732
[5 empty] | 10295 | 26.2964 | 11356.76 | 10.8784 | 10295 | 1.173
[5 empty] | 10300 | 26.2962 | 11362.4 | 10.8778 | 10300 | 1.1728
[5 empty] | 10305 | 26.296 | 11368.04 | 10.8772 | 10305 | 1.1726
[5 empty] | 10310 | 26.2957 | 11373.68 | 10.8766 | 10310 | 1.1724
[5 empty] | 10315 | 26.2954 | 11379.32 | 10.8759 | 10315 | 1.1722
[5 empty] | 10320 | 26.295 | 11384.96 | 10.8752 | 10320 | 1.1719
[5 empty] | 10325 | 26.2945 | 11390.6 | 10.8745 | 10325 | 1.1717
[5 empty] | 10330 | 26.2941 | 11396.24 | 10.8738 | 10330 | 1.1714
[5 empty] | 10335 | 26.2936 | 11401.88 | 10.8731 | 10335 | 1.1711
[5 empty] | 10340 | 26.2931 | 11407.52 | 10.8723 | 10340 | 1.1709
[5 empty] | 10345 | 26.2926 | 11413.16 | 10.8716 | 10345 | 1.1706
[5 empty] | 10350 | 26.292 | 11418.8 | 10.8708 | 10350 | 1.1703
[5 empty] | 10355 | 26.2916 | 11424.44 | 10.87 | 10355 | 1.1699
[5 empty] | 10360 | 26.2911 | 11430.08 | 10.8692 | 10360 | 1.1696
[5 empty] | 10365 | 26.2906 | 11435.72 | 10.8683 | 10365 | 1.1693
[5 empty] | 10370 | 26.2901 | 11441.36 | 10.8675 | 10370 | 1.169
[5 empty] | 10375 | 26.2897 | 11447 | 10.8666 | 10375 | 1.1687
[5 empty] | 10380 | 26.2893 | 11452.64 | 10.8657 | 10380 | 1.1684
[5 empty] | 10385 | 26.2889 | 11458.28 | 10.8648 | 10385 | 1.168
[5 empty] | 10390 | 26.2885 | 11463.92 | 10.8639 | 10390 | 1.1677
[5 empty] | 10395 | 26.2882 | 11469.56 | 10.8629 | 10395 | 1.1674
[5 empty] | 10400 | 26.2879 | 11475.2 | 10.862 | 10400 | 1.1671
[5 empty] | 10405 | 26.2876 | 11480.84 | 10.861 | 10405 | 1.1668
[5 empty] | 10410 | 26.2873 | 11486.48 | 10.86 | 10410 | 1.1665
[5 empty] | 10415 | 26.287 | 11492.12 | 10.8589 | 10415 | 1.1662
[5 empty] | 10420 | 26.2867 | 11497.76 | 10.8579 | 10420 | 1.1659
[5 empty] | 10425 | 26.2865 | 11503.4 | 10.8568 | 10425 | 1.1656
[5 empty] | 10430 | 26.2863 | 11509.04 | 10.8557 | 10430 | 1.1654
[5 empty] | 10435 | 26.2861 | 11514.68 | 10.8545 | 10435 | 1.1651
[5 empty] | 10440 | 26.2859 | 11520.32 | 10.8534 | 10440 | 1.1649
[5 empty] | 10445 | 26.2858 | 11525.96 | 10.8521 | 10445 | 1.1646
[5 empty] | 10450 | 26.2857 | 11531.6 | 10.8509 | 10450 | 1.1644
[5 empty] | 10455 | 26.2856 | 11537.24 | 10.8496 | 10455 | 1.1642
[5 empty] | 10460 | 26.2854 | 11542.88 | 10.8482 | 10460 | 1.164
[5 empty] | 10465 | 26.2853 | 11548.52 | 10.8468 | 10465 | 1.1638
[5 empty] | 10470 | 26.2852 | 11554.16 | 10.8453 | 10470 | 1.1637
[5 empty] | 10475 | 26.285 | 11559.8 | 10.8438 | 10475 | 1.1635
[5 empty] | 10480 | 26.2847 | 11565.44 | 10.8422 | 10480 | 1.1634
[5 empty] | 10485 | 26.2844 | 11571.08 | 10.8405 | 10485 | 1.1633
[5 empty] | 10490 | 26.284 | 11576.72 | 10.8387 | 10490 | 1.1631
[5 empty] | 10495 | 26.2836 | 11582.36 | 10.8369 | 10495 | 1.1631
[5 empty] | 10500 | 26.283 | 11588 | 10.835 | 10500 | 1.163
[5 empty] | 10505 | 26.2825 | 11593.64 | 10.833 | 10505 | 1.1629
[5 empty] | 10510 | 26.2818 | 11599.28 | 10.831 | 10510 | 1.1629
[5 empty] | 10515 | 26.2811 | 11604.92 | 10.8288 | 10515 | 1.1628
[5 empty] | 10520 | 26.2803 | 11610.56 | 10.8266 | 10520 | 1.1628
[5 empty] | 10525 | 26.2793 | 11616.2 | 10.8243 | 10525 | 1.1628
[5 empty] | 10530 | 26.2783 | 11621.84 | 10.822 | 10530 | 1.1628
[5 empty] | 10535 | 26.2773 | 11627.48 | 10.8195 | 10535 | 1.1628
[5 empty] | 10540 | 26.2762 | 11633.12 | 10.817 | 10540 | 1.1628
[5 empty] | 10545 | 26.275 | 11638.76 | 10.8144 | 10545 | 1.1628
[5 empty] | 10550 | 26.2737 | 11644.4 | 10.8118 | 10550 | 1.1629
[5 empty] | 10555 | 26.2724 | 11650.04 | 10.8091 | 10555 | 1.1629
[5 empty] | 10560 | 26.271 | 11655.68 | 10.8063 | 10560 | 1.163
[5 empty] | 10565 | 26.2696 | 11661.32 | 10.8035 | 10565 | 1.1631
[5 empty] | 10570 | 26.2681 | 11666.96 | 10.8006 | 10570 | 1.1632
[5 empty] | 10575 | 26.2666 | 11672.6 | 10.7976 | 10575 | 1.1633
[5 empty] | 10580 | 26.2651 | 11678.24 | 10.7946 | 10580 | 1.1634
[5 empty] | 10585 | 26.2636 | 11683.88 | 10.7916 | 10585 | 1.1635
[5 empty] | 10590 | 26.262 | 11689.52 | 10.7885 | 10590 | 1.1636
[5 empty] | 10595 | 26.2605 | 11695.16 | 10.7854 | 10595 | 1.1638
[5 empty] | 10600 | 26.2591 | 11700.8 | 10.7822 | 10600 | 1.1639
[5 empty] | 10605 | 26.2576 | 11706.44 | 10.7791 | 10605 | 1.1641
[5 empty] | 10610 | 26.2562 | 11712.08 | 10.7759 | 10610 | 1.1643
[5 empty] | 10615 | 26.2548 | 11717.72 | 10.7727 | 10615 | 1.1645
[5 empty] | 10620 | 26.2536 | 11723.36 | 10.7695 | 10620 | 1.1647
[5 empty] | 10625 | 26.2524 | 11729 | 10.7663 | 10625 | 1.1648
[5 empty] | 10630 | 26.2514 | 11734.64 | 10.7631 | 10630 | 1.165
[5 empty] | 10635 | 26.2504 | 11740.28 | 10.7599 | 10635 | 1.1653
[5 empty] | 10640 | 26.2496 | 11745.92 | 10.7568 | 10640 | 1.1655
[5 empty] | 10645 | 26.2489 | 11751.56 | 10.7537 | 10645 | 1.1657
[5 empty] | 10650 | 26.2484 | 11757.2 | 10.7507 | 10650 | 1.166
[5 empty] | 10655 | 26.2481 | 11762.84 | 10.7477 | 10655 | 1.1662
[5 empty] | 10660 | 26.2478 | 11768.48 | 10.7448 | 10660 | 1.1665
[5 empty] | 10665 | 26.2476 | 11774.12 | 10.742 | 10665 | 1.1668
[5 empty] | 10670 | 26.2477 | 11779.76 | 10.7393 | 10670 | 1.1671
[5 empty] | 10675 | 26.2478 | 11785.4 | 10.7367 | 10675 | 1.1674
[5 empty] | 10680 | 26.248 | 11791.04 | 10.7342 | 10680 | 1.1676
[5 empty] | 10685 | 26.2483 | 11796.68 | 10.7318 | 10685 | 1.168
[5 empty] | 10690 | 26.2487 | 11802.32 | 10.7295 | 10690 | 1.1683
[5 empty] | 10695 | 26.2492 | 11807.96 | 10.7274 | 10695 | 1.1686
[5 empty] | 10700 | 26.2497 | 11813.6 | 10.7254 | 10700 | 1.1689
[5 empty] | 10705 | 26.2503 | 11819.24 | 10.7236 | 10705 | 1.1692
[5 empty] | 10710 | 26.2508 | 11824.88 | 10.722 | 10710 | 1.1695
[5 empty] | 10715 | 26.2513 | 11830.52 | 10.7205 | 10715 | 1.1699
[5 empty] | 10720 | 26.2519 | 11836.16 | 10.7191 | 10720 | 1.1702
[5 empty] | 10725 | 26.2524 | 11841.8 | 10.7179 | 10725 | 1.1705
[5 empty] | 10730 | 26.2529 | 11847.44 | 10.7169 | 10730 | 1.1709
[5 empty] | 10735 | 26.2532 | 11853.08 | 10.716 | 10735 | 1.1712
[5 empty] | 10740 | 26.2534 | 11858.72 | 10.7153 | 10740 | 1.1716
[5 empty] | 10745 | 26.2536 | 11864.36 | 10.7147 | 10745 | 1.1719
[5 empty] | 10750 | 26.2537 | 11870 | 10.7142 | 10750 | 1.1723
[5 empty] | 10755 | 26.2536 | 11875.64 | 10.7139 | 10755 | 1.1726
[5 empty] | 10760 | 26.2534 | 11881.28 | 10.7137 | 10760 | 1.1729
[5 empty] | 10765 | 26.2531 | 11886.92 | 10.7136 | 10765 | 1.1733
[5 empty] | 10770 | 26.2527 | 11892.56 | 10.7136 | 10770 | 1.1736
[5 empty] | 10775 | 26.2521 | 11898.2 | 10.7137 | 10775 | 1.1739
[5 empty] | 10780 | 26.2515 | 11903.84 | 10.7139 | 10780 | 1.1743
[5 empty] | 10785 | 26.2507 | 11909.48 | 10.7141 | 10785 | 1.1746
[5 empty] | 10790 | 26.2498 | 11915.12 | 10.7145 | 10790 | 1.1749
[5 empty] | 10795 | 26.2489 | 11920.76 | 10.7149 | 10795 | 1.1752
[5 empty] | 10800 | 26.248 | 11926.4 | 10.7153 | 10800 | 1.1756
[5 empty] | 10805 | 26.247 | 11932.04 | 10.7159 | 10805 | 1.1759
[5 empty] | 10810 | 26.246 | 11937.68 | 10.7165 | 10810 | 1.1762
[5 empty] | 10815 | 26.2449 | 11943.32 | 10.7172 | 10815 | 1.1764
[5 empty] | 10820 | 26.244 | 11948.96 | 10.7179 | 10820 | 1.1767
[5 empty] | 10825 | 26.2432 | 11954.6 | 10.7187 | 10825 | 1.177
[5 empty] | 10830 | 26.2425 | 11960.24 | 10.7196 | 10830 | 1.1773
[5 empty] | 10835 | 26.2417 | 11965.88 | 10.7205 | 10835 | 1.1775
[5 empty] | 10840 | 26.241 | 11971.52 | 10.7215 | 10840 | 1.1778
[5 empty] | 10845 | 26.2406 | 11977.16 | 10.7227 | 10845 | 1.178
[5 empty] | 10850 | 26.2401 | 11982.8 | 10.7238 | 10850 | 1.1783
[5 empty] | 10855 | 26.2399 | 11988.44 | 10.7251 | 10855 | 1.1785
[5 empty] | 10860 | 26.2397 | 11994.08 | 10.7265 | 10860 | 1.1787
[5 empty] | 10865 | 26.2396 | 11999.72 | 10.7279 | 10865 | 1.179
[5 empty] | 10870 | 26.2395 | 12005.36 | 10.7295 | 10870 | 1.1792
[5 empty] | 10875 | 26.2395 | 12011 | 10.7311 | 10875 | 1.1794
[5 empty] | 10880 | 26.2396 | 12016.64 | 10.7327 | 10880 | 1.1795
[5 empty] | 10885 | 26.2397 | 12022.28 | 10.7345 | 10885 | 1.1797
[5 empty] | 10890 | 26.2399 | 12027.92 | 10.7363 | 10890 | 1.1799
[5 empty] | 10895 | 26.24 | 12033.56 | 10.7381 | 10895 | 1.18
[5 empty] | 10900 | 26.2402 | 12039.2 | 10.74 | 10900 | 1.1802
[5 empty] | 10905 | 26.2404 | 12044.84 | 10.7419 | 10905 | 1.1803
[5 empty] | 10910 | 26.2406 | 12050.48 | 10.7437 | 10910 | 1.1804
[5 empty] | 10915 | 26.2408 | 12056.12 | 10.7456 | 10915 | 1.1805
[5 empty] | 10920 | 26.2409 | 12061.76 | 10.7475 | 10920 | 1.1806
[5 empty] | 10925 | 26.2411 | 12067.4 | 10.7493 | 10925 | 1.1807
[5 empty] | 10930 | 26.2412 | 12073.04 | 10.751 | 10930 | 1.1808
[5 empty] | 10935 | 26.2414 | 12078.68 | 10.7527 | 10935 | 1.1809
[5 empty] | 10940 | 26.2415 | 12084.32 | 10.7544 | 10940 | 1.181
[5 empty] | 10945 | 26.2417 | 12089.96 | 10.7559 | 10945 | 1.181
[5 empty] | 10950 | 26.2419 | 12095.6 | 10.7574 | 10950 | 1.1811
[5 empty] | 10955 | 26.2421 | 12101.24 | 10.7588 | 10955 | 1.1811
[5 empty] | 10960 | 26.2424 | 12106.88 | 10.7601 | 10960 | 1.1812
[5 empty] | 10965 | 26.2426 | 12112.52 | 10.7613 | 10965 | 1.1812
[5 empty] | 10970 | 26.2428 | 12118.16 | 10.7625 | 10970 | 1.1812
[5 empty] | 10975 | 26.2431 | 12123.8 | 10.7636 | 10975 | 1.1812
[5 empty] | 10980 | 26.2434 | 12129.44 | 10.7646 | 10980 | 1.1813
[5 empty] | 10985 | 26.2436 | 12135.08 | 10.7656 | 10985 | 1.1813
[5 empty] | 10990 | 26.2438 | 12140.72 | 10.7665 | 10990 | 1.1812
[5 empty] | 10995 | 26.244 | 12146.36 | 10.7674 | 10995 | 1.1812
[5 empty] | 11000 | 26.2441 | 12152 | 10.7683 | 11000 | 1.1812
[5 empty] | 11005 | 26.2442 | 12157.64 | 10.7692 | 11005 | 1.1812
[5 empty] | 11010 | 26.2442 | 12163.28 | 10.7701 | 11010 | 1.1812
[5 empty] | 11015 | 26.2441 | 12168.92 | 10.7709 | 11015 | 1.1811
[5 empty] | 11020 | 26.2438 | 12174.56 | 10.7718 | 11020 | 1.1811
[5 empty] | 11025 | 26.2435 | 12180.2 | 10.7728 | 11025 | 1.181
[5 empty] | 11030 | 26.243 | 12185.84 | 10.7737 | 11030 | 1.181
[5 empty] | 11035 | 26.2424 | 12191.48 | 10.7747 | 11035 | 1.1809
[5 empty] | 11040 | 26.2416 | 12197.12 | 10.7757 | 11040 | 1.1808
[5 empty] | 11045 | 26.2407 | 12202.76 | 10.7767 | 11045 | 1.1807
[5 empty] | 11050 | 26.2397 | 12208.4 | 10.7777 | 11050 | 1.1806
[5 empty] | 11055 | 26.2385 | 12214.04 | 10.7787 | 11055 | 1.1805
[5 empty] | 11060 | 26.2373 | 12219.68 | 10.7797 | 11060 | 1.1804
[5 empty] | 11065 | 26.2359 | 12225.32 | 10.7807 | 11065 | 1.1803
[5 empty] | 11070 | 26.2345 | 12230.96 | 10.7817 | 11070 | 1.1802
[5 empty] | 11075 | 26.2329 | 12236.6 | 10.7827 | 11075 | 1.1801
[5 empty] | 11080 | 26.2312 | 12242.24 | 10.7836 | 11080 | 1.1799
[5 empty] | 11085 | 26.2296 | 12247.88 | 10.7844 | 11085 | 1.1798
[5 empty] | 11090 | 26.228 | 12253.52 | 10.7852 | 11090 | 1.1796
[5 empty] | 11095 | 26.2264 | 12259.16 | 10.7859 | 11095 | 1.1794
[5 empty] | 11100 | 26.2248 | 12264.8 | 10.7865 | 11100 | 1.1793
[5 empty] | 11105 | 26.2232 | 12270.44 | 10.787 | 11105 | 1.1791
[5 empty] | 11110 | 26.2217 | 12276.08 | 10.7874 | 11110 | 1.1789
[5 empty] | 11115 | 26.2205 | 12281.72 | 10.7877 | 11115 | 1.1787
[5 empty] | 11120 | 26.2192 | 12287.36 | 10.7879 | 11120 | 1.1785
[5 empty] | 11125 | 26.218 | 12293 | 10.7879 | 11125 | 1.1784
[5 empty] | 11130 | 26.217 | 12298.64 | 10.7879 | 11130 | 1.1782
[5 empty] | 11135 | 26.2162 | 12304.28 | 10.7877 | 11135 | 1.178
[5 empty] | 11140 | 26.2155 | 12309.92 | 10.7874 | 11140 | 1.1778
[5 empty] | 11145 | 26.215 | 12315.56 | 10.787 | 11145 | 1.1776
[5 empty] | 11150 | 26.2147 | 12321.2 | 10.7866 | 11150 | 1.1773
[5 empty] | 11155 | 26.2144 | 12326.84 | 10.786 | 11155 | 1.1771
[5 empty] | 11160 | 26.2143 | 12332.48 | 10.7854 | 11160 | 1.1769
[5 empty] | 11165 | 26.2145 | 12338.12 | 10.7847 | 11165 | 1.1767
[5 empty] | 11170 | 26.2147 | 12343.76 | 10.784 | 11170 | 1.1765
[5 empty] | 11175 | 26.215 | 12349.4 | 10.7833 | 11175 | 1.1763
[5 empty] | 11180 | 26.2155 | 12355.04 | 10.7825 | 11180 | 1.1761
[5 empty] | 11185 | 26.216 | 12360.68 | 10.7818 | 11185 | 1.1759
[5 empty] | 11190 | 26.2166 | 12366.32 | 10.781 | 11190 | 1.1758
[5 empty] | 11195 | 26.2172 | 12371.96 | 10.7803 | 11195 | 1.1756
[5 empty] | 11200 | 26.2179 | 12377.6 | 10.7797 | 11200 | 1.1754
[5 empty] | 11205 | 26.2187 | 12383.24 | 10.779 | 11205 | 1.1753
[5 empty] | 11210 | 26.2194 | 12388.88 | 10.7785 | 11210 | 1.1751
[5 empty] | 11215 | 26.22 | 12394.52 | 10.778 | 11215 | 1.175
[5 empty] | 11220 | 26.2207 | 12400.16 | 10.7776 | 11220 | 1.1749
[5 empty] | 11225 | 26.2213 | 12405.8 | 10.7772 | 11225 | 1.1748
[5 empty] | 11230 | 26.2218 | 12411.44 | 10.7769 | 11230 | 1.1747
[5 empty] | 11235 | 26.2223 | 12417.08 | 10.7768 | 11235 | 1.1746
[5 empty] | 11240 | 26.2225 | 12422.72 | 10.7766 | 11240 | 1.1745
[5 empty] | 11245 | 26.2228 | 12428.36 | 10.7766 | 11245 | 1.1745
[5 empty] | 11250 | 26.2229 | 12434 | 10.7766 | 11250 | 1.1744
[5 empty] | 11255 | 26.2229 | 12439.64 | 10.7767 | 11255 | 1.1744
[5 empty] | 11260 | 26.2228 | 12445.28 | 10.7768 | 11260 | 1.1744
[5 empty] | 11265 | 26.2224 | 12450.92 | 10.777 | 11265 | 1.1743
[5 empty] | 11270 | 26.2219 | 12456.56 | 10.7772 | 11270 | 1.1743
[5 empty] | 11275 | 26.2214 | 12462.2 | 10.7775 | 11275 | 1.1744
[5 empty] | 11280 | 26.2207 | 12467.84 | 10.7778 | 11280 | 1.1744
[5 empty] | 11285 | 26.2199 | 12473.48 | 10.778 | 11285 | 1.1744
[5 empty] | 11290 | 26.2189 | 12479.12 | 10.7784 | 11290 | 1.1744
[5 empty] | 11295 | 26.2178 | 12484.76 | 10.7787 | 11295 | 1.1745
[5 empty] | 11300 | 26.2166 | 12490.4 | 10.779 | 11300 | 1.1745
[5 empty] | 11305 | 26.2153 | 12496.04 | 10.7793 | 11305 | 1.1746
[5 empty] | 11310 | 26.2139 | 12501.68 | 10.7797 | 11310 | 1.1747
[5 empty] | 11315 | 26.2124 | 12507.32 | 10.78 | 11315 | 1.1748
[5 empty] | 11320 | 26.211 | 12512.96 | 10.7804 | 11320 | 1.1748
[5 empty] | 11325 | 26.2094 | 12518.6 | 10.7808 | 11325 | 1.1749
[5 empty] | 11330 | 26.2078 | 12524.24 | 10.7812 | 11330 | 1.175
[5 empty] | 11335 | 26.2062 | 12529.88 | 10.7816 | 11335 | 1.1751
[5 empty] | 11340 | 26.2046 | 12535.52 | 10.7821 | 11340 | 1.1752
[5 empty] | 11345 | 26.203 | 12541.16 | 10.7825 | 11345 | 1.1752
[5 empty] | 11350 | 26.2014 | 12546.8 | 10.783 | 11350 | 1.1753
[5 empty] | 11355 | 26.1998 | 12552.44 | 10.7835 | 11355 | 1.1754
[5 empty] | 11360 | 26.1982 | 12558.08 | 10.7841 | 11360 | 1.1754
[5 empty] | 11365 | 26.1966 | 12563.72 | 10.7846 | 11365 | 1.1755
[5 empty] | 11370 | 26.1949 | 12569.36 | 10.7852 | 11370 | 1.1755
[5 empty] | 11375 | 26.1933 | 12575 | 10.7858 | 11375 | 1.1756
[5 empty] | 11380 | 26.1917 | 12580.64 | 10.7864 | 11380 | 1.1756
[5 empty] | 11385 | 26.19 | 12586.28 | 10.7871 | 11385 | 1.1757
[5 empty] | 11390 | 26.1883 | 12591.92 | 10.7877 | 11390 | 1.1757
[5 empty] | 11395 | 26.1864 | 12597.56 | 10.7884 | 11395 | 1.1757
[5 empty] | 11400 | 26.1845 | 12603.2 | 10.789 | 11400 | 1.1757
[5 empty] | 11405 | 26.1825 | 12608.84 | 10.7896 | 11405 | 1.1757
[5 empty] | 11410 | 26.1804 | 12614.48 | 10.7902 | 11410 | 1.1757
[5 empty] | 11415 | 26.1781 | 12620.12 | 10.7907 | 11415 | 1.1757
[5 empty] | 11420 | 26.1756 | 12625.76 | 10.7913 | 11420 | 1.1756
[5 empty] | 11425 | 26.1731 | 12631.4 | 10.7918 | 11425 | 1.1756
[5 empty] | 11430 | 26.1703 | 12637.04 | 10.7922 | 11430 | 1.1755
[5 empty] | 11435 | 26.1675 | 12642.68 | 10.7926 | 11435 | 1.1754
[5 empty] | 11440 | 26.1644 | 12648.32 | 10.7929 | 11440 | 1.1754
[5 empty] | 11445 | 26.1612 | 12653.96 | 10.7932 | 11445 | 1.1753
[5 empty] | 11450 | 26.1578 | 12659.6 | 10.7935 | 11450 | 1.1752
[5 empty] | 11455 | 26.1542 | 12665.24 | 10.7937 | 11455 | 1.1751
[5 empty] | 11460 | 26.1505 | 12670.88 | 10.7938 | 11460 | 1.175
[5 empty] | 11465 | 26.1467 | 12676.52 | 10.7939 | 11465 | 1.1748
[5 empty] | 11470 | 26.1429 | 12682.16 | 10.794 | 11470 | 1.1747
[5 empty] | 11475 | 26.1388 | 12687.8 | 10.794 | 11475 | 1.1746
[5 empty] | 11480 | 26.1346 | 12693.44 | 10.794 | 11480 | 1.1744
[5 empty] | 11485 | 26.1303 | 12699.08 | 10.7939 | 11485 | 1.1742
[5 empty] | 11490 | 26.1259 | 12704.72 | 10.7939 | 11490 | 1.1741
[5 empty] | 11495 | 26.1215 | 12710.36 | 10.7938 | 11495 | 1.1739
[5 empty] | 11500 | 26.117 | 12716 | 10.7936 | 11500 | 1.1737
[5 empty] | 11505 | 26.1124 | 12721.64 | 10.7935 | 11505 | 1.1735
[5 empty] | 11510 | 26.1078 | 12727.28 | 10.7933 | 11510 | 1.1733
[5 empty] | 11515 | 26.1031 | 12732.92 | 10.7931 | 11515 | 1.1731
[5 empty] | 11520 | 26.0984 | 12738.56 | 10.7928 | 11520 | 1.1728
[5 empty] | 11525 | 26.0935 | 12744.2 | 10.7926 | 11525 | 1.1726
[5 empty] | 11530 | 26.0886 | 12749.84 | 10.7923 | 11530 | 1.1724
[5 empty] | 11535 | 26.0836 | 12755.48 | 10.792 | 11535 | 1.1722
[5 empty] | 11540 | 26.0785 | 12761.12 | 10.7917 | 11540 | 1.1719
[5 empty] | 11545 | 26.0734 | 12766.76 | 10.7913 | 11545 | 1.1717
[5 empty] | 11550 | 26.0682 | 12772.4 | 10.7909 | 11550 | 1.1714
[5 empty] | 11555 | 26.0628 | 12778.04 | 10.7905 | 11555 | 1.1712
[5 empty] | 11560 | 26.0574 | 12783.68 | 10.79 | 11560 | 1.1709
[5 empty] | 11565 | 26.0519 | 12789.32 | 10.7896 | 11565 | 1.1706
[5 empty] | 11570 | 26.0462 | 12794.96 | 10.7891 | 11570 | 1.1704
[5 empty] | 11575 | 26.0405 | 12800.6 | 10.7885 | 11575 | 1.1701
[5 empty] | 11580 | 26.0347 | 12806.24 | 10.788 | 11580 | 1.1698
[5 empty] | 11585 | 26.0287 | 12811.88 | 10.7874 | 11585 | 1.1695
[5 empty] | 11590 | 26.0227 | 12817.52 | 10.7868 | 11590 | 1.1693
[5 empty] | 11595 | 26.0166 | 12823.16 | 10.7862 | 11595 | 1.169
[5 empty] | 11600 | 26.0104 | 12828.8 | 10.7856 | 11600 | 1.1687
[5 empty] | 11605 | 26.0042 | 12834.44 | 10.785 | 11605 | 1.1685
[5 empty] | 11610 | 25.9979 | 12840.08 | 10.7844 | 11610 | 1.1682
[5 empty] | 11615 | 25.9916 | 12845.72 | 10.7838 | 11615 | 1.1679
[5 empty] | 11620 | 25.9851 | 12851.36 | 10.7833 | 11620 | 1.1676
[5 empty] | 11625 | 25.9787 | 12857 | 10.7827 | 11625 | 1.1673
[5 empty] | 11630 | 25.9722 | 12862.64 | 10.7822 | 11630 | 1.167
[5 empty] | 11635 | 25.9656 | 12868.28 | 10.7817 | 11635 | 1.1667
[5 empty] | 11640 | 25.9591 | 12873.92 | 10.7812 | 11640 | 1.1665
[5 empty] | 11645 | 25.9525 | 12879.56 | 10.7808 | 11645 | 1.1662
[5 empty] | 11650 | 25.9458 | 12885.2 | 10.7803 | 11650 | 1.1659
[5 empty] | 11655 | 25.939 | 12890.84 | 10.7799 | 11655 | 1.1656
[5 empty] | 11660 | 25.9322 | 12896.48 | 10.7796 | 11660 | 1.1653
[5 empty] | 11665 | 25.9252 | 12902.12 | 10.7792 | 11665 | 1.1651
[5 empty] | 11670 | 25.9183 | 12907.76 | 10.7789 | 11670 | 1.1648
[5 empty] | 11675 | 25.9113 | 12913.4 | 10.7786 | 11675 | 1.1645
[5 empty] | 11680 | 25.9041 | 12919.04 | 10.7783 | 11680 | 1.1642
[5 empty] | 11685 | 25.8967 | 12924.68 | 10.778 | 11685 | 1.1639
[5 empty] | 11690 | 25.8893 | 12930.32 | 10.7777 | 11690 | 1.1636
[5 empty] | 11695 | 25.8818 | 12935.96 | 10.7774 | 11695 | 1.1633
[5 empty] | 11700 | 25.8742 | 12941.6 | 10.7771 | 11700 | 1.1631
[5 empty] | 11705 | 25.8664 | 12947.24 | 10.7767 | 11705 | 1.1628
[5 empty] | 11710 | 25.8585 | 12952.88 | 10.7763 | 11710 | 1.1625
[5 empty] | 11715 | 25.8505 | 12958.52 | 10.7759 | 11715 | 1.1622
[5 empty] | 11720 | 25.8423 | 12964.16 | 10.7754 | 11720 | 1.1619
[5 empty] | 11725 | 25.834 | 12969.8 | 10.7749 | 11725 | 1.1616
[5 empty] | 11730 | 25.8257 | 12975.44 | 10.7743 | 11730 | 1.1613
[5 empty] | 11735 | 25.8174 | 12981.08 | 10.7737 | 11735 | 1.161
[5 empty] | 11740 | 25.809 | 12986.72 | 10.773 | 11740 | 1.1607
[5 empty] | 11745 | 25.8005 | 12992.36 | 10.7722 | 11745 | 1.1603
[5 empty] | 11750 | 25.7921 | 12998 | 10.7713 | 11750 | 1.16
[5 empty] | 11755 | 25.7836 | 13003.64 | 10.7704 | 11755 | 1.1597
[5 empty] | 11760 | 25.7753 | 13009.28 | 10.7694 | 11760 | 1.1593
[5 empty] | 11765 | 25.7672 | 13014.92 | 10.7683 | 11765 | 1.159
[5 empty] | 11770 | 25.7592 | 13020.56 | 10.7671 | 11770 | 1.1586
[5 empty] | 11775 | 25.7514 | 13026.2 | 10.7659 | 11775 | 1.1583
[5 empty] | 11780 | 25.7436 | 13031.84 | 10.7646 | 11780 | 1.1579
[5 empty] | 11785 | 25.7361 | 13037.48 | 10.7632 | 11785 | 1.1575
[5 empty] | 11790 | 25.7289 | 13043.12 | 10.7617 | 11790 | 1.1572
[5 empty] | 11795 | 25.7221 | 13048.76 | 10.7602 | 11795 | 1.1568
[5 empty] | 11800 | 25.7156 | 13054.4 | 10.7586 | 11800 | 1.1564
[5 empty] | 11805 | 25.7093 | 13060.04 | 10.7569 | 11805 | 1.156
[5 empty] | 11810 | 25.7033 | 13065.68 | 10.7552 | 11810 | 1.1556
[5 empty] | 11815 | 25.6978 | 13071.32 | 10.7535 | 11815 | 1.1552
[5 empty] | 11820 | 25.6926 | 13076.96 | 10.7517 | 11820 | 1.1548
[5 empty] | 11825 | 25.6878 | 13082.6 | 10.7499 | 11825 | 1.1544
[5 empty] | 11830 | 25.6834 | 13088.24 | 10.7481 | 11830 | 1.1539
[5 empty] | 11835 | 25.6794 | 13093.88 | 10.7462 | 11835 | 1.1535
[5 empty] | 11840 | 25.6758 | 13099.52 | 10.7444 | 11840 | 1.1531
[5 empty] | 11845 | 25.6727 | 13105.16 | 10.7426 | 11845 | 1.1526
[5 empty] | 11850 | 25.6699 | 13110.8 | 10.7407 | 11850 | 1.1522
[5 empty] | 11855 | 25.6675 | 13116.44 | 10.7389 | 11855 | 1.1518
[5 empty] | 11860 | 25.6655 | 13122.08 | 10.7372 | 11860 | 1.1514
[5 empty] | 11865 | 25.6638 | 13127.72 | 10.7354 | 11865 | 1.1509
[5 empty] | 11870 | 25.6625 | 13133.36 | 10.7338 | 11870 | 1.1505
[5 empty] | 11875 | 25.6615 | 13139 | 10.7321 | 11875 | 1.1501
[5 empty] | 11880 | 25.6608 | 13144.64 | 10.7306 | 11880 | 1.1497
[5 empty] | 11885 | 25.6605 | 13150.28 | 10.7291 | 11885 | 1.1493
[5 empty] | 11890 | 25.6606 | 13155.92 | 10.7277 | 11890 | 1.1489
[5 empty] | 11895 | 25.6609 | 13161.56 | 10.7263 | 11895 | 1.1485
[5 empty] | 11900 | 25.6616 | 13167.2 | 10.7251 | 11900 | 1.1482
[5 empty] | 11905 | 25.6624 | 13172.84 | 10.7239 | 11905 | 1.1478
[5 empty] | 11910 | 25.6634 | 13178.48 | 10.7228 | 11910 | 1.1475
[5 empty] | 11915 | 25.6648 | 13184.12 | 10.7218 | 11915 | 1.1472
[5 empty] | 11920 | 25.6664 | 13189.76 | 10.7209 | 11920 | 1.1469
[5 empty] | 11925 | 25.6682 | 13195.4 | 10.7201 | 11925 | 1.1466
[5 empty] | 11930 | 25.6704 | 13201.04 | 10.7193 | 11930 | 1.1463
[5 empty] | 11935 | 25.6726 | 13206.68 | 10.7186 | 11935 | 1.1461
[5 empty] | 11940 | 25.6751 | 13212.32 | 10.7179 | 11940 | 1.1459
[5 empty] | 11945 | 25.6778 | 13217.96 | 10.7174 | 11945 | 1.1457
[5 empty] | 11950 | 25.6807 | 13223.6 | 10.7169 | 11950 | 1.1455
[5 empty] | 11955 | 25.6838 | 13229.24 | 10.7164 | 11955 | 1.1454
[5 empty] | 11960 | 25.6871 | 13234.88 | 10.716 | 11960 | 1.1453
[5 empty] | 11965 | 25.6905 | 13240.52 | 10.7156 | 11965 | 1.1452
[5 empty] | 11970 | 25.6942 | 13246.16 | 10.7152 | 11970 | 1.1451
[5 empty] | 11975 | 25.6981 | 13251.8 | 10.7149 | 11975 | 1.1451
[5 empty] | 11980 | 25.7021 | 13257.44 | 10.7146 | 11980 | 1.1451
[5 empty] | 11985 | 25.7062 | 13263.08 | 10.7143 | 11985 | 1.1451
[5 empty] | 11990 | 25.7105 | 13268.72 | 10.714 | 11990 | 1.1451
[5 empty] | 11995 | 25.715 | 13274.36 | 10.7137 | 11995 | 1.1452
[5 empty] | 12000 | 25.7196 | 13280 | 10.7134 | 12000 | 1.1453
[5 empty] | 12005 | 25.7244 | 13285.64 | 10.7132 | 12005 | 1.1454
[5 empty] | 12010 | 25.7293 | 13291.28 | 10.7129 | 12010 | 1.1456
[5 empty] | 12015 | 25.7342 | 13296.92 | 10.7127 | 12015 | 1.1457
[5 empty] | 12020 | 25.7391 | 13302.56 | 10.7124 | 12020 | 1.1459
[5 empty] | 12025 | 25.7441 | 13308.2 | 10.7122 | 12025 | 1.1462
[5 empty] | 12030 | 25.7492 | 13313.84 | 10.712 | 12030 | 1.1464
[5 empty] | 12035 | 25.7544 | 13319.48 | 10.7118 | 12035 | 1.1467
[5 empty] | 12040 | 25.7595 | 13325.12 | 10.7116 | 12040 | 1.147
[5 empty] | 12045 | 25.7646 | 13330.76 | 10.7114 | 12045 | 1.1473
[5 empty] | 12050 | 25.7695 | 13336.4 | 10.7112 | 12050 | 1.1476
[5 empty] | 12055 | 25.7744 | 13342.04 | 10.7111 | 12055 | 1.1479
[5 empty] | 12060 | 25.7793 | 13347.68 | 10.711 | 12060 | 1.1482
[5 empty] | 12065 | 25.7841 | 13353.32 | 10.7109 | 12065 | 1.1486
[5 empty] | 12070 | 25.7887 | 13358.96 | 10.7109 | 12070 | 1.1489
[5 empty] | 12075 | 25.7932 | 13364.6 | 10.7109 | 12075 | 1.1493
[5 empty] | 12080 | 25.7975 | 13370.24 | 10.7109 | 12080 | 1.1496
[5 empty] | 12085 | 25.8016 | 13375.88 | 10.7109 | 12085 | 1.15
[5 empty] | 12090 | 25.8055 | 13381.52 | 10.711 | 12090 | 1.1504
[5 empty] | 12095 | 25.8093 | 13387.16 | 10.7111 | 12095 | 1.1507
[5 empty] | 12100 | 25.8129 | 13392.8 | 10.7113 | 12100 | 1.1511
[5 empty] | 12105 | 25.8162 | 13398.44 | 10.7115 | 12105 | 1.1514
[5 empty] | 12110 | 25.8193 | 13404.08 | 10.7117 | 12110 | 1.1518
[5 empty] | 12115 | 25.8221 | 13409.72 | 10.712 | 12115 | 1.1521
[5 empty] | 12120 | 25.8248 | 13415.36 | 10.7122 | 12120 | 1.1525
[5 empty] | 12125 | 25.8272 | 13421 | 10.7126 | 12125 | 1.1528
[5 empty] | 12130 | 25.8294 | 13426.64 | 10.7129 | 12130 | 1.1531
[5 empty] | 12135 | 25.8314 | 13432.28 | 10.7133 | 12135 | 1.1534
[5 empty] | 12140 | 25.8332 | 13437.92 | 10.7137 | 12140 | 1.1537
[5 empty] | 12145 | 25.8347 | 13443.56 | 10.7142 | 12145 | 1.1539
[5 empty] | 12150 | 25.8361 | 13449.2 | 10.7146 | 12150 | 1.1542
[5 empty] | 12155 | 25.8372 | 13454.84 | 10.7152 | 12155 | 1.1544
[5 empty] | 12160 | 25.8383 | 13460.48 | 10.7157 | 12160 | 1.1546
[5 empty] | 12165 | 25.8391 | 13466.12 | 10.7163 | 12165 | 1.1547
[5 empty] | 12170 | 25.8399 | 13471.76 | 10.7169 | 12170 | 1.1549
[5 empty] | 12175 | 25.8404 | 13477.4 | 10.7175 | 12175 | 1.155
[5 empty] | 12180 | 25.8409 | 13483.04 | 10.7182 | 12180 | 1.1552
[5 empty] | 12185 | 25.8412 | 13488.68 | 10.7189 | 12185 | 1.1552
[5 empty] | 12190 | 25.8414 | 13494.32 | 10.7197 | 12190 | 1.1553
[5 empty] | 12195 | 25.8416 | 13499.96 | 10.7205 | 12195 | 1.1554
[5 empty] | 12200 | 25.8417 | 13505.6 | 10.7213 | 12200 | 1.1554
[5 empty] | 12205 | 25.8418 | 13511.24 | 10.7222 | 12205 | 1.1554
[5 empty] | 12210 | 25.8417 | 13516.88 | 10.7232 | 12210 | 1.1554
[5 empty] | 12215 | 25.8417 | 13522.52 | 10.7242 | 12215 | 1.1554
[5 empty] | 12220 | 25.8417 | 13528.16 | 10.7252 | 12220 | 1.1554
[5 empty] | 12225 | 25.8417 | 13533.8 | 10.7263 | 12225 | 1.1553
[5 empty] | 12230 | 25.8417 | 13539.44 | 10.7275 | 12230 | 1.1552
[5 empty] | 12235 | 25.8417 | 13545.08 | 10.7287 | 12235 | 1.1551
[5 empty] | 12240 | 25.8418 | 13550.72 | 10.73 | 12240 | 1.155
[5 empty] | 12245 | 25.8419 | 13556.36 | 10.7313 | 12245 | 1.1549
[5 empty] | 12250 | 25.8421 | 13562 | 10.7327 | 12250 | 1.1548
[5 empty] | 12255 | 25.8423 | 13567.64 | 10.7341 | 12255 | 1.1546
[5 empty] | 12260 | 25.8426 | 13573.28 | 10.7356 | 12260 | 1.1545
[5 empty] | 12265 | 25.8431 | 13578.92 | 10.7371 | 12265 | 1.1543
[5 empty] | 12270 | 25.8435 | 13584.56 | 10.7386 | 12270 | 1.1541
[5 empty] | 12275 | 25.8441 | 13590.2 | 10.7402 | 12275 | 1.1539
[5 empty] | 12280 | 25.8446 | 13595.84 | 10.7418 | 12280 | 1.1538
[5 empty] | 12285 | 25.8454 | 13601.48 | 10.7434 | 12285 | 1.1536
[5 empty] | 12290 | 25.8462 | 13607.12 | 10.745 | 12290 | 1.1534
[5 empty] | 12295 | 25.8471 | 13612.76 | 10.7466 | 12295 | 1.1532
[5 empty] | 12300 | 25.8479 | 13618.4 | 10.7482 | 12300 | 1.153
[5 empty] | 12305 | 25.849 | 13624.04 | 10.7498 | 12305 | 1.1528
[5 empty] | 12310 | 25.85 | 13629.68 | 10.7514 | 12310 | 1.1526
[5 empty] | 12315 | 25.8511 | 13635.32 | 10.7529 | 12315 | 1.1524
[5 empty] | 12320 | 25.8521 | 13640.96 | 10.7545 | 12320 | 1.1522
[5 empty] | 12325 | 25.8533 | 13646.6 | 10.756 | 12325 | 1.1521
[5 empty] | 12330 | 25.8545 | 13652.24 | 10.7575 | 12330 | 1.1519
[5 empty] | 12335 | 25.8557 | 13657.88 | 10.7589 | 12335 | 1.1518
[5 empty] | 12340 | 25.8568 | 13663.52 | 10.7603 | 12340 | 1.1516
[5 empty] | 12345 | 25.858 | 13669.16 | 10.7616 | 12345 | 1.1515
[5 empty] | 12350 | 25.8591 | 13674.8 | 10.7629 | 12350 | 1.1514
[5 empty] | 12355 | 25.8602 | 13680.44 | 10.7642 | 12355 | 1.1513
[5 empty] | 12360 | 25.8613 | 13686.08 | 10.7653 | 12360 | 1.1512
[5 empty] | 12365 | 25.8623 | 13691.72 | 10.7665 | 12365 | 1.1512
[5 empty] | 12370 | 25.8633 | 13697.36 | 10.7676 | 12370 | 1.1511
[5 empty] | 12375 | 25.8643 | 13703 | 10.7686 | 12375 | 1.1511
[5 empty] | 12380 | 25.8652 | 13708.64 | 10.7696 | 12380 | 1.1511
[5 empty] | 12385 | 25.866 | 13714.28 | 10.7705 | 12385 | 1.1511
[5 empty] | 12390 | 25.8667 | 13719.92 | 10.7714 | 12390 | 1.1511
[5 empty] | 12395 | 25.8674 | 13725.56 | 10.7723 | 12395 | 1.1512
[5 empty] | 12400 | 25.8679 | 13731.2 | 10.7731 | 12400 | 1.1512
[5 empty] | 12405 | 25.8685 | 13736.84 | 10.7738 | 12405 | 1.1513
[5 empty] | 12410 | 25.8689 | 13742.48 | 10.7745 | 12410 | 1.1514
[5 empty] | 12415 | 25.8693 | 13748.12 | 10.7752 | 12415 | 1.1515
[5 empty] | 12420 | 25.8695 | 13753.76 | 10.7758 | 12420 | 1.1517
[5 empty] | 12425 | 25.8697 | 13759.4 | 10.7764 | 12425 | 1.1518
[5 empty] | 12430 | 25.8699 | 13765.04 | 10.777 | 12430 | 1.152
[5 empty] | 12435 | 25.8699 | 13770.68 | 10.7775 | 12435 | 1.1522
[5 empty] | 12440 | 25.8699 | 13776.32 | 10.778 | 12440 | 1.1525
[5 empty] | 12445 | 25.8698 | 13781.96 | 10.7785 | 12445 | 1.1527
[5 empty] | 12450 | 25.8697 | 13787.6 | 10.779 | 12450 | 1.153
[5 empty] | 12455 | 25.8695 | 13793.24 | 10.7794 | 12455 | 1.1532
[5 empty] | 12460 | 25.8693 | 13798.88 | 10.7798 | 12460 | 1.1535
[5 empty] | 12465 | 25.8691 | 13804.52 | 10.7802 | 12465 | 1.1538
[5 empty] | 12470 | 25.8688 | 13810.16 | 10.7806 | 12470 | 1.1542
[5 empty] | 12475 | 25.8685 | 13815.8 | 10.7809 | 12475 | 1.1545
[5 empty] | 12480 | 25.8683 | 13821.44 | 10.7813 | 12480 | 1.1549
[5 empty] | 12485 | 25.8679 | 13827.08 | 10.7816 | 12485 | 1.1552
[5 empty] | 12490 | 25.8676 | 13832.72 | 10.7819 | 12490 | 1.1556
[5 empty] | 12495 | 25.8674 | 13838.36 | 10.7822 | 12495 | 1.156
[5 empty] | 12500 | 25.8673 | 13844 | 10.7825 | 12500 | 1.1564
[5 empty] | 12505 | 25.8672 | 13849.64 | 10.7828 | 12505 | 1.1568
[5 empty] | 12510 | 25.8672 | 13855.28 | 10.7831 | 12510 | 1.1572
[5 empty] | 12515 | 25.8672 | 13860.92 | 10.7834 | 12515 | 1.1576
[5 empty] | 12520 | 25.8673 | 13866.56 | 10.7837 | 12520 | 1.158
[5 empty] | 12525 | 25.8675 | 13872.2 | 10.784 | 12525 | 1.1584
[5 empty] | 12530 | 25.8679 | 13877.84 | 10.7843 | 12530 | 1.1588
[5 empty] | 12535 | 25.8684 | 13883.48 | 10.7846 | 12535 | 1.1592
[5 empty] | 12540 | 25.8691 | 13889.12 | 10.785 | 12540 | 1.1596
[5 empty] | 12545 | 25.8698 | 13894.76 | 10.7853 | 12545 | 1.16
[5 empty] | 12550 | 25.8706 | 13900.4 | 10.7857 | 12550 | 1.1604
[5 empty] | 12555 | 25.8715 | 13906.04 | 10.7861 | 12555 | 1.1607
[5 empty] | 12560 | 25.8727 | 13911.68 | 10.7865 | 12560 | 1.1611
[5 empty] | 12565 | 25.8739 | 13917.32 | 10.787 | 12565 | 1.1615
[5 empty] | 12570 | 25.8753 | 13922.96 | 10.7874 | 12570 | 1.1618
[5 empty] | 12575 | 25.8767 | 13928.6 | 10.7879 | 12575 | 1.1622
[5 empty] | 12580 | 25.8783 | 13934.24 | 10.7884 | 12580 | 1.1625
[5 empty] | 12585 | 25.88 | 13939.88 | 10.789 | 12585 | 1.1628
[5 empty] | 12590 | 25.8816 | 13945.52 | 10.7896 | 12590 | 1.1631
[5 empty] | 12595 | 25.8834 | 13951.16 | 10.7902 | 12595 | 1.1634
[5 empty] | 12600 | 25.8852 | 13956.8 | 10.7908 | 12600 | 1.1637
[5 empty] | 12605 | 25.887 | 13962.44 | 10.7915 | 12605 | 1.1639
[5 empty] | 12610 | 25.889 | 13968.08 | 10.7922 | 12610 | 1.1641
[5 empty] | 12615 | 25.8909 | 13973.72 | 10.7929 | 12615 | 1.1643
[5 empty] | 12620 | 25.8928 | 13979.36 | 10.7936 | 12620 | 1.1645
[5 empty] | 12625 | 25.8946 | 13985 | 10.7944 | 12625 | 1.1647
[5 empty] | 12630 | 25.8965 | 13990.64 | 10.7951 | 12630 | 1.1649
[5 empty] | 12635 | 25.8983 | 13996.28 | 10.7959 | 12635 | 1.165
[5 empty] | 12640 | 25.9 | 14001.92 | 10.7967 | 12640 | 1.1651
[5 empty] | 12645 | 25.9017 | 14007.56 | 10.7974 | 12645 | 1.1652
[5 empty] | 12650 | 25.9033 | 14013.2 | 10.7982 | 12650 | 1.1653
[5 empty] | 12655 | 25.9049 | 14018.84 | 10.7989 | 12655 | 1.1653
[5 empty] | 12660 | 25.9064 | 14024.48 | 10.7996 | 12660 | 1.1653
[5 empty] | 12665 | 25.9078 | 14030.12 | 10.8003 | 12665 | 1.1653
[5 empty] | 12670 | 25.9091 | 14035.76 | 10.801 | 12670 | 1.1653
[5 empty] | 12675 | 25.9103 | 14041.4 | 10.8016 | 12675 | 1.1653
[5 empty] | 12680 | 25.9115 | 14047.04 | 10.8022 | 12680 | 1.1652
[5 empty] | 12685 | 25.9126 | 14052.68 | 10.8027 | 12685 | 1.1652
[5 empty] | 12690 | 25.9135 | 14058.32 | 10.8032 | 12690 | 1.1651
[5 empty] | 12695 | 25.9145 | 14063.96 | 10.8036 | 12695 | 1.165
[5 empty] | 12700 | 25.9154 | 14069.6 | 10.8041 | 12700 | 1.1649
[5 empty] | 12705 | 25.9162 | 14075.24 | 10.8044 | 12705 | 1.1647
[5 empty] | 12710 | 25.917 | 14080.88 | 10.8048 | 12710 | 1.1646
[5 empty] | 12715 | 25.9176 | 14086.52 | 10.8053 | 12715 | 1.1644
[5 empty] | 12720 | 25.9182 |  |  | 12720 | 1.1642
[5 empty] | 12725 | 25.9188 |  |  | 12725 | 1.164
[5 empty] | 12730 | 25.9193 |  |  | 12730 | 1.1638
[5 empty] | 12735 | 25.9198 |  |  | 12735 | 1.1636
[5 empty] | 12740 | 25.9202 |  |  | 12740 | 1.1634
[5 empty] | 12745 | 25.9206 |  |  | 12745 | 1.1631
[5 empty] | 12750 | 25.9208 |  |  | 12750 | 1.1629
[5 empty] | 12755 | 25.921 |  |  | 12755 | 1.1626
[5 empty] | 12760 | 25.9212 |  |  | 12760 | 1.1623
[5 empty] | 12765 | 25.9212 |  |  | 12765 | 1.162
[5 empty] | 12770 | 25.9213 |  |  | 12770 | 1.1618
[5 empty] | 12775 | 25.9212 |  |  | 12775 | 1.1615
[5 empty] | 12780 | 25.9211 |  |  | 12780 | 1.1612
[5 empty] | 12785 | 25.9208 |  |  | 12785 | 1.1609
[5 empty] | 12790 | 25.9203 |  |  | 12790 | 1.1605
[5 empty] | 12795 | 25.9198 |  |  | 12795 | 1.1602
[5 empty] | 12800 | 25.9191 |  |  | 12800 | 1.1599
[5 empty] | 12805 | 25.9184 |  |  | 12805 | 1.1596
[5 empty] | 12810 | 25.9176 |  |  | 12810 | 1.1593
[5 empty] | 12815 | 25.9165 |  |  | 12815 | 1.159
[5 empty] | 12820 | 25.9154 |  |  | 12820 | 1.1587
[5 empty] | 12825 | 25.9141 |  |  | 12825 | 1.1584
[5 empty] | 12830 | 25.9127 |  |  | 12830 | 1.1581
[5 empty] | 12835 | 25.9111 |  |  | 12835 | 1.1577
[5 empty] | 12840 | 25.9094 |  |  | 12840 | 1.1574
[5 empty] | 12845 | 25.9076 |  |  | 12845 | 1.1571
[5 empty] | 12850 | 25.9057 |  |  | 12850 | 1.1568
[5 empty] | 12855 | 25.9038 |  |  | 12855 | 1.1565
[5 empty] | 12860 | 25.9017 |  |  | 12860 | 1.1562
[5 empty] | 12865 | 25.8995 |  |  | 12865 | 1.156
[5 empty] | 12870 | 25.8972 |  |  | 12870 | 1.1557
[5 empty] | 12875 | 25.8949 |  |  | 12875 | 1.1554
[5 empty] | 12880 | 25.8925 |  |  | 12880 | 1.1552
[5 empty] | 12885 | 25.8901 |  |  | 12885 | 1.1549
[5 empty] | 12890 | 25.8876 |  |  | 12890 | 1.1547
[5 empty] | 12895 | 25.8852 |  |  | 12895 | 1.1544
[5 empty] | 12900 | 25.8828 |  |  | 12900 | 1.1542
[5 empty] | 12905 | 25.8803 |  |  | 12905 | 1.154
[5 empty] | 12910 | 25.8778 |  |  | 12910 | 1.1538
[5 empty] | 12915 | 25.8755 |  |  | 12915 | 1.1536
[5 empty] | 12920 | 25.8731 |  |  | 12920 | 1.1534
[5 empty] | 12925 | 25.8708 |  |  | 12925 | 1.1532
[5 empty] | 12930 | 25.8686 |  |  | 12930 | 1.153
[5 empty] | 12935 | 25.8664 |  |  | 12935 | 1.1528
[5 empty] | 12940 | 25.8642 |  |  | 12940 | 1.1527
[5 empty] | 12945 | 25.8621 |  |  | 12945 | 1.1525
[5 empty] | 12950 | 25.86 |  |  | 12950 | 1.1524
[5 empty] | 12955 | 25.858 |  |  | 12955 | 1.1523
[5 empty] | 12960 | 25.856 |  |  | 12960 | 1.1522
[5 empty] | 12965 | 25.8541 |  |  | 12965 | 1.1521
[5 empty] | 12970 | 25.8522 |  |  | 12970 | 1.152
[5 empty] | 12975 | 25.8504 |  |  | 12975 | 1.152
[5 empty] | 12980 | 25.8486 |  |  | 12980 | 1.1519
[5 empty] | 12985 | 25.8467 |  |  | 12985 | 1.1519
[5 empty] | 12990 | 25.8449 |  |  | 12990 | 1.1519
[5 empty] | 12995 | 25.8431 |  |  | 12995 | 1.1519
[5 empty] | 13000 | 25.8413 |  |  | 13000 | 1.1519
[5 empty] | 13005 | 25.8395 |  |  | 13005 | 1.1519
[5 empty] | 13010 | 25.8377 |  |  | 13010 | 1.152
[5 empty] | 13015 | 25.8359 |  |  | 13015 | 1.152
[5 empty] | 13020 | 25.834 |  |  | 13020 | 1.1521
[5 empty] | 13025 | 25.8321 |  |  | 13025 | 1.1522
[5 empty] | 13030 | 25.8303 |  |  | 13030 | 1.1523
[5 empty] | 13035 | 25.8283 |  |  | 13035 | 1.1524
[5 empty] | 13040 | 25.8264 |  |  | 13040 | 1.1525
[5 empty] | 13045 | 25.8243 |  |  | 13045 | 1.1526
[5 empty] | 13050 | 25.8224 |  |  | 13050 | 1.1528
[5 empty] | 13055 | 25.8203 |  |  | 13055 | 1.1529
[5 empty] | 13060 | 25.8182 |  |  | 13060 | 1.1531
[5 empty] | 13065 | 25.816 |  |  | 13065 | 1.1533
[5 empty] | 13070 | 25.8138 |  |  | 13070 | 1.1535
[5 empty] | 13075 | 25.8116 |  |  | 13075 | 1.1537
[5 empty] | 13080 | 25.8093 |  |  | 13080 | 1.1539
[5 empty] | 13085 | 25.807 |  |  | 13085 | 1.1541
[5 empty] | 13090 | 25.8047 |  |  | 13090 | 1.1543
[5 empty] | 13095 | 25.8024 |  |  | 13095 | 1.1546
[5 empty] | 13100 | 25.8 |  |  | 13100 | 1.1548
[5 empty] | 13105 | 25.7976 |  |  | 13105 | 1.155
[5 empty] | 13110 | 25.7951 |  |  | 13110 | 1.1553
[5 empty] | 13115 | 25.7926 |  |  | 13115 | 1.1555
[5 empty] | 13120 | 25.7901 |  |  | 13120 | 1.1558
[5 empty] | 13125 | 25.7875 |  |  | 13125 | 1.1561
[5 empty] | 13130 | 25.7849 |  |  | 13130 | 1.1563
[5 empty] | 13135 | 25.7822 |  |  | 13135 | 1.1566
[5 empty] | 13140 | 25.7796 |  |  | 13140 | 1.1569
[5 empty] | 13145 | 25.7769 |  |  | 13145 | 1.1571
[5 empty] | 13150 | 25.7741 |  |  | 13150 | 1.1574
[5 empty] | 13155 | 25.7714 |  |  | 13155 | 1.1577
[5 empty] | 13160 | 25.7686 |  |  | 13160 | 1.1579
[5 empty] | 13165 | 25.7658 |  |  | 13165 | 1.1582
[5 empty] | 13170 | 25.7629 |  |  | 13170 | 1.1584
[5 empty] | 13175 | 25.76 |  |  | 13175 | 1.1587
[5 empty] | 13180 | 25.7571 |  |  | 13180 | 1.1589
[5 empty] | 13185 | 25.7542 |  |  | 13185 | 1.1591
[5 empty] | 13190 | 25.7513 |  |  | 13190 | 1.1594
[5 empty] | 13195 | 25.7484 |  |  | 13195 | 1.1596
[5 empty] | 13200 | 25.7455 |  |  | 13200 | 1.1598
[5 empty] | 13205 | 25.7426 |  |  | 13205 | 1.16
[5 empty] | 13210 | 25.7397 |  |  | 13210 | 1.1601
[5 empty] | 13215 | 25.7368 |  |  | 13215 | 1.1603
[5 empty] | 13220 | 25.734 |  |  | 13220 | 1.1604
[5 empty] | 13225 | 25.7313 |  |  | 13225 | 1.1606
[5 empty] | 13230 | 25.7287 |  |  | 13230 | 1.1607
[5 empty] | 13235 | 25.7261 |  |  | 13235 | 1.1608
[5 empty] | 13240 | 25.7235 |  |  | 13240 | 1.1609
[5 empty] | 13245 | 25.721 |  |  | 13245 | 1.161
[5 empty] | 13250 | 25.7187 |  |  | 13250 | 1.161
[5 empty] | 13255 | 25.7164 |  |  | 13255 | 1.1611
[5 empty] | 13260 | 25.7142 |  |  | 13260 | 1.1611
[5 empty] | 13265 | 25.7122 |  |  | 13265 | 1.1611
[5 empty] | 13270 | 25.7103 |  |  | 13270 | 1.1611
[5 empty] | 13275 | 25.7085 |  |  | 13275 | 1.1611
[5 empty] | 13280 | 25.7068 |  |  | 13280 | 1.1611
[5 empty] | 13285 | 25.7053 |  |  | 13285 | 1.1611
[5 empty] | 13290 | 25.7038 |  |  | 13290 | 1.161
[5 empty] | 13295 | 25.7024 |  |  | 13295 | 1.1609
[5 empty] | 13300 | 25.7012 |  |  | 13300 | 1.1608
[5 empty] | 13305 | 25.7001 |  |  | 13305 | 1.1607
[5 empty] | 13310 | 25.6991 |  |  | 13310 | 1.1606
[5 empty] | 13315 | 25.6983 |  |  | 13315 | 1.1605
[5 empty] | 13320 | 25.6975 |  |  | 13320 | 1.1604
[5 empty] | 13325 | 25.6968 |  |  | 13325 | 1.1602
[5 empty] | 13330 | 25.6961 |  |  | 13330 | 1.16
[5 empty] | 13335 | 25.6956 |  |  | 13335 | 1.1598
[5 empty] | 13340 | 25.695 |  |  | 13340 | 1.1597
[5 empty] | 13345 | 25.6946 |  |  | 13345 | 1.1595
[5 empty] | 13350 | 25.6942 |  |  | 13350 | 1.1593
[5 empty] | 13355 | 25.6939 |  |  | 13355 | 1.159
[5 empty] | 13360 | 25.6936 |  |  | 13360 | 1.1588
[5 empty] | 13365 | 25.6933 |  |  | 13365 | 1.1586
[5 empty] | 13370 | 25.6931 |  |  | 13370 | 1.1583
[5 empty] | 13375 | 25.6929 |  |  | 13375 | 1.1581
[5 empty] | 13380 | 25.6926 |  |  | 13380 | 1.1578
[5 empty] | 13385 | 25.6925 |  |  | 13385 | 1.1576
[5 empty] | 13390 | 25.6924 |  |  | 13390 | 1.1573
[5 empty] | 13395 | 25.6923 |  |  | 13395 | 1.157
[5 empty] | 13400 | 25.6923 |  |  | 13400 | 1.1568
[5 empty] | 13405 | 25.6923 |  |  | 13405 | 1.1565
[5 empty] | 13410 | 25.6924 |  |  | 13410 | 1.1562
[5 empty] | 13415 | 25.6925 |  |  | 13415 | 1.1559
[5 empty] | 13420 | 25.6926 |  |  | 13420 | 1.1556
[5 empty] | 13425 | 25.6928 |  |  | 13425 | 1.1553
[5 empty] | 13430 | 25.6931 |  |  | 13430 | 1.155
[5 empty] | 13435 | 25.6935 |  |  | 13435 | 1.1547
[5 empty] | 13440 | 25.6941 |  |  | 13440 | 1.1544
[5 empty] | 13445 | 25.6948 |  |  | 13445 | 1.1542
[5 empty] | 13450 | 25.6955 |  |  | 13450 | 1.1539
[5 empty] | 13455 | 25.6964 |  |  | 13455 | 1.1536
[5 empty] | 13460 | 25.6974 |  |  | 13460 | 1.1533
[5 empty] | 13465 | 25.6984 |  |  | 13465 | 1.153
[5 empty] | 13470 | 25.6997 |  |  | 13470 | 1.1527
[5 empty] | 13475 | 25.7011 |  |  | 13475 | 1.1525
[5 empty] | 13480 | 25.7027 |  |  | 13480 | 1.1522
[5 empty] | 13485 | 25.7045 |  |  | 13485 | 1.1519
[5 empty] | 13490 | 25.7064 |  |  | 13490 | 1.1516
[5 empty] | 13495 | 25.7084 |  |  | 13495 | 1.1514
[5 empty] | 13500 | 25.7106 |  |  | 13500 | 1.1511
[5 empty] | 13505 | 25.7129 |  |  | 13505 | 1.1509
[5 empty] | 13510 | 25.7153 |  |  | 13510 | 1.1506
[5 empty] | 13515 | 25.7179 |  |  | 13515 | 1.1504
[5 empty] | 13520 | 25.7206 |  |  | 13520 | 1.1501
[5 empty] | 13525 | 25.7235 |  |  | 13525 | 1.1499
[5 empty] | 13530 | 25.7265 |  |  | 13530 | 1.1497
[5 empty] | 13535 | 25.7297 |  |  | 13535 | 1.1494
[5 empty] | 13540 | 25.7329 |  |  | 13540 | 1.1492
[5 empty] | 13545 | 25.7361 |  |  | 13545 | 1.149
[5 empty] | 13550 | 25.7394 |  |  | 13550 | 1.1488
[5 empty] | 13555 | 25.7427 |  |  | 13555 | 1.1486
[5 empty] | 13560 | 25.7462 |  |  | 13560 | 1.1484
[5 empty] | 13565 | 25.7496 |  |  | 13565 | 1.1482
[5 empty] | 13570 | 25.7532 |  |  | 13570 | 1.148
[5 empty] | 13575 | 25.7568 |  |  | 13575 | 1.1478
[5 empty] | 13580 | 25.7603 |  |  | 13580 | 1.1477
[5 empty] | 13585 | 25.7638 |  |  | 13585 | 1.1475
[5 empty] | 13590 | 25.7673 |  |  | 13590 | 1.1473
[5 empty] | 13595 | 25.7708 |  |  | 13595 | 1.1472
[5 empty] | 13600 | 25.7742 |  |  | 13600 | 1.147
[5 empty] | 13605 | 25.7777 |  |  | 13605 | 1.1469
[5 empty] | 13610 | 25.7812 |  |  | 13610 | 1.1468
[5 empty] | 13615 | 25.7845 |  |  | 13615 | 1.1466
[5 empty] | 13620 | 25.7878 |  |  | 13620 | 1.1465
[5 empty] | 13625 | 25.791 |  |  | 13625 | 1.1464
[5 empty] | 13630 | 25.7942 |  |  | 13630 | 1.1463
[5 empty] | 13635 | 25.7973 |  |  | 13635 | 1.1463
[5 empty] | 13640 | 25.8004 |  |  | 13640 | 1.1462
[5 empty] | 13645 | 25.8034 |  |  | 13645 | 1.1461
[5 empty] | 13650 | 25.8063 |  |  | 13650 | 1.146
[5 empty] | 13655 | 25.8092 |  |  | 13655 | 1.146
[5 empty] | 13660 | 25.812 |  |  | 13660 | 1.1459
[5 empty] | 13665 | 25.8147 |  |  | 13665 | 1.1459
[5 empty] | 13670 | 25.8174 |  |  | 13670 | 1.1458
[5 empty] | 13675 | 25.82 |  |  | 13675 | 1.1458
[5 empty] | 13680 | 25.8225 |  |  | 13680 | 1.1458
[5 empty] | 13685 | 25.825 |  |  | 13685 | 1.1458
[5 empty] | 13690 | 25.8274 |  |  | 13690 | 1.1458
[5 empty] | 13695 | 25.8298 |  |  | 13695 | 1.1458
[5 empty] | 13700 | 25.832 |  |  | 13700 | 1.1458
[5 empty] | 13705 | 25.8343 |  |  | 13705 | 1.1458
[5 empty] | 13710 | 25.8365 |  |  | 13710 | 1.1459
[5 empty] | 13715 | 25.8386 |  |  | 13715 | 1.1459
[5 empty] | 13720 | 25.8406 |  |  | 13720 | 1.146
[5 empty] | 13725 | 25.8427 |  |  | 13725 | 1.146
[5 empty] | 13730 | 25.8446 |  |  | 13730 | 1.1461
[5 empty] | 13735 | 25.8465 |  |  | 13735 | 1.1462
[5 empty] | 13740 | 25.8483 |  |  | 13740 | 1.1463
[5 empty] | 13745 | 25.8502 |  |  | 13745 | 1.1464
[5 empty] | 13750 | 25.8519 |  |  | 13750 | 1.1465
[5 empty] | 13755 | 25.8536 |  |  | 13755 | 1.1466
[5 empty] | 13760 | 25.8553 |  |  | 13760 | 1.1468
[5 empty] | 13765 | 25.8569 |  |  | 13765 | 1.1469
[5 empty] | 13770 | 25.8586 |  |  | 13770 | 1.147
[5 empty] | 13775 | 25.8602 |  |  | 13775 | 1.1472
[5 empty] | 13780 | 25.8617 |  |  | 13780 | 1.1473
[5 empty] | 13785 | 25.8633 |  |  | 13785 | 1.1475
[5 empty] | 13790 | 25.8648 |  |  | 13790 | 1.1477
[5 empty] | 13795 | 25.8662 |  |  | 13795 | 1.1479
[5 empty] | 13800 | 25.8676 |  |  | 13800 | 1.1481
[5 empty] | 13805 | 25.8691 |  |  | 13805 | 1.1483
[5 empty] | 13810 | 25.8706 |  |  | 13810 | 1.1485
[5 empty] | 13815 | 25.872 |  |  | 13815 | 1.1487
[5 empty] | 13820 | 25.8734 |  |  | 13820 | 1.1489
[5 empty] | 13825 | 25.8748 |  |  | 13825 | 1.1491
[5 empty] | 13830 | 25.8763 |  |  | 13830 | 1.1494
[5 empty] | 13835 | 25.8777 |  |  | 13835 | 1.1496
[5 empty] | 13840 | 25.8791 |  |  | 13840 | 1.1498
[5 empty] | 13845 | 25.8805 |  |  | 13845 | 1.1501
[5 empty] | 13850 | 25.882 |  |  | 13850 | 1.1503
[5 empty] | 13855 | 25.8834 |  |  | 13855 | 1.1506
[5 empty] | 13860 | 25.8848 |  |  | 13860 | 1.1508
[5 empty] | 13865 | 25.8863 |  |  | 13865 | 1.1511
[5 empty] | 13870 | 25.8878 |  |  | 13870 | 1.1513
[5 empty] | 13875 | 25.8893 |  |  | 13875 | 1.1516
[5 empty] | 13880 | 25.8908 |  |  | 13880 | 1.1519
[5 empty] | 13885 | 25.8923 |  |  | 13885 | 1.1521
[5 empty] | 13890 | 25.8938 |  |  | 13890 | 1.1524
[5 empty] | 13895 | 25.8953 |  |  | 13895 | 1.1527
[5 empty] | 13900 | 25.8968 |  |  | 13900 | 1.1529
[5 empty] | 13905 | 25.8984 |  |  | 13905 | 1.1532
[5 empty] | 13910 | 25.8999 |  |  | 13910 | 1.1535
[5 empty] | 13915 | 25.9014 |  |  | 13915 | 1.1537
[5 empty] | 13920 | 25.903 |  |  | 13920 | 1.154
[5 empty] | 13925 | 25.9044 |  |  | 13925 | 1.1543
[5 empty] | 13930 | 25.9059 |  |  | 13930 | 1.1546
[5 empty] | 13935 | 25.9074 |  |  | 13935 | 1.1549
[5 empty] | 13940 | 25.9089 |  |  | 13940 | 1.1551
[5 empty] | 13945 | 25.9104 |  |  | 13945 | 1.1554
[5 empty] | 13950 | 25.9119 |  |  | 13950 | 1.1557
[5 empty] | 13955 | 25.9134 |  |  | 13955 | 1.1559
[5 empty] | 13960 | 25.9148 |  |  | 13960 | 1.1562
[5 empty] | 13965 | 25.9161 |  |  | 13965 | 1.1564
[5 empty] | 13970 | 25.9175 |  |  | 13970 | 1.1567
[5 empty] | 13975 | 25.9189 |  |  | 13975 | 1.1569
[5 empty] | 13980 | 25.9202 |  |  | 13980 | 1.1572
[5 empty] | 13985 | 25.9215 |  |  | 13985 | 1.1574
[5 empty] | 13990 | 25.9227 |  |  | 13990 | 1.1577
[5 empty] | 13995 | 25.924 |  |  | 13995 | 1.158
[5 empty] | 14000 | 25.9252 |  |  | 14000 | 1.1582

## Response Curves (hidden)
LCM4 |  | LCM4-NE
Wavelength | Response % | Wavelength | Response %
400 | 32.4054 | 400 | 35.8774
405 | 40.4216 | 405 | 39.6637
410 | 41.4621 | 410 | 37.6928
415 | 44.5271 | 415 | 42.7694
420 | 46.0875 | 420 | 40.9667
425 | 47.3473 | 425 | 44.4681
430 | 54.0259 | 430 | 46.5637
435 | 49.4824 | 435 | 48.6143
440 | 51.6475 | 440 | 51.0815
445 | 55.3994 | 445 | 48.8498
450 | 58.4612 | 450 | 47.4997
455 | 59.3794 | 455 | 51.0887
460 | 60.0824 | 460 | 52.025
465 | 59.4404 | 465 | 57.5405
470 | 56.6602 | 470 | 54.6515
475 | 61.6423 | 475 | 55.2979
480 | 60.9335 | 480 | 59.4581
485 | 60.9593 | 485 | 58.3028
490 | 57.262 | 490 | 56.1991
495 | 56.9867 | 495 | 59.7064
500 | 61.4079 | 500 | 63.8231
505 | 61.7169 | 505 | 60.33
510 | 62.6422 | 510 | 60.9806
515 | 60.971 | 515 | 62.2479
520 | 62.1591 | 520 | 58.2231
525 | 64.7004 | 525 | 56.7729
530 | 60.5157 | 530 | 59.3748
535 | 60.7254 | 535 | 59.3484
540 | 63.974 | 540 | 62.1525
545 | 62.3857 | 545 | 60.4913
550 | 63.673 | 550 | 58.3599
555 | 61.688 | 555 | 60.4681
560 | 59.2577 | 560 | 63.298
565 | 60.0315 | 565 | 60.5341
570 | 61.8935 | 570 | 59.0613
575 | 58.9168 | 575 | 60.0996
580 | 55.8765 | 580 | 61.2325
585 | 56.9772 | 585 | 63.8598
590 | 58.141 | 590 | 63.2619
595 | 60.5741 | 595 | 57.0149
600 | 61.4238 | 600 | 55.7296
605 | 58.4145 | 605 | 61.0682
610 | 55.7011 | 610 | 64.5976
615 | 56.7253 | 615 | 60.8972
620 | 60.6507 | 620 | 57.1657
625 | 60.5435 | 625 | 57.377
630 | 55.1537 | 630 | 58.793
635 | 52.6742 | 635 | 60.1857
640 | 53.4327 | 640 | 63.7535
645 | 54.5021 | 645 | 65.0509
650 | 53.5796 | 650 | 61.1943
655 | 53.2318 | 655 | 56.4765
660 | 55.023 | 660 | 55.5736
665 | 54.0687 | 665 | 59.4357
670 | 49.8468 | 670 | 61.8606
675 | 47.3568 | 675 | 58.568
680 | 48.6328 | 680 | 53.9744
685 | 51.6799 | 685 | 50.2978
690 | 52.0223 | 690 | 50.7079
695 | 49.5848 | 695 | 49.6419
700 | 46.2683 | 700 | 51.0868
705 | 43.162 | 705 | 52.3967
710 | 42.4083 | 710 | 55.7134
715 | 43.2478 | 715 | 57.8586
720 | 42.1857 | 720 | 54.2109
725 | 41.801 | 725 | 50.4169
730 | 41.9641 | 730 | 47.959
735 | 41.8769 | 735 | 47.9635
740 | 41.6014 | 740 | 52.2562
745 | 38.2234 | 745 | 54.1569
750 | 34.9791 | 750 | 53.7786
755 | 33.7577 | 755 | 51.9035
760 | 33.7707 | 760 | 48.3072
765 | 34.3821 | 765 | 45.4431
770 | 35.8639 | 770 | 45.5676
775 | 33.8886 | 775 | 45.6893
780 | 31.1042 | 780 | 44.6455
785 | 29.088 | 785 | 43.9836
790 | 26.6204 | 790 | 44.2588
795 | 24.3069 | 795 | 45.1698
800 | 23.8878 | 800 | 45.5856
805 | 24.6185 | 805 | 43.7586
810 | 25.3428 | 810 | 38.3096
815 | 24.142 | 815 | 35.8962
820 | 24.1216 | 820 | 33.5358
825 | 24.4774 | 825 | 33.4596
830 | 24.2763 | 830 | 33.721
835 | 23.5873 | 835 | 34.1944
840 | 21.2077 | 840 | 35.705
845 | 18.5279 | 845 | 35.773
850 | 16.8565 | 850 | 34.6651
855 | 17.8562 | 855 | 31.2919
860 | 17.9781 | 860 | 28.2023
865 | 18.7157 | 865 | 26.6403
870 | 18.6907 | 870 | 25.0753
875 | 17.9585 | 875 | 25.2004
880 | 16.8631 | 880 | 25.1673
885 | 15.8945 | 885 | 23.8729
890 | 15.6367 | 890 | 23.6994
895 | 12.8187 | 895 | 21.6841
900 | 12.0464 | 900 | 21.5073
905 | 11.3911 | 905 | 22.4362
910 | 11.1154 | 910 | 22.3443
915 | 10.9139 | 915 | 22.1628
920 | 8.5596 | 920 | 19.9815
925 | 8.4431 | 925 | 17.6193
930 | 8.8104 | 930 | 15.0154
935 | 8.0504 | 935 | 13.8349
940 | 7.5418 | 940 | 11.4053
945 | 7.416 | 945 | 11.4659
950 | 7.44 | 950 | 10.5356
955 | 6.7411 | 955 | 10.5234
960 | 6.3177 | 960 | 9.9351
965 | 5.1835 | 965 | 8.7235
970 | 3.7933 | 970 | 8.8678
975 | 3.2978 | 975 | 9.3437
980 | 3.158 | 980 | 8.1065
985 | 3.1957 | 985 | 7.0602
990 | 2.1232 | 990 | 5.8018
995 | 1.8478 | 995 | 5.3955
1000 | 1.0579 | 1000 | 3.764
1290 | 1.2527e-06
1340 | 1.2527e-07
1550 | 0.00013088